## Slide 1: Jefferson Lab 12 GeV Upgrade Accelerator Readiness Review Phase 1B
October 22-24, 2013

## Slide 2: Jefferson Lab 12 GeV Upgrade Accelerator Readiness Review Phase 1B
Roger Erickson
SLAC National Accelerator Laboratory
ARR Team Chairman
October 22-24, 2013

## Slide 3: The ARR Team
Roger Erickson, SLAC – Chair
John Anderson, FNAL
Roger Carlini, JLab Ian Evans, SLAC Harry Fanning, JLab
Jim Floyd, LBNL
Dennis Parzyck, Consultant Karen White, ORNL Facilitator: Robert May, JLab
DOE Observers:
Salma El-Safwany, LBNL Site Office
Several from JLab DOE Site Office and from DOE Oak Ridge Office.

## Slide 4: The ARR Plan
Two Phases: Phase 1A (August 19-23) and Phase 1B (now).
No new activity was expected to be authorized in response to the week’s Phase 1A Review, regardless of outcome.
The decision to commence beam commissioning to be based partly on the recommendations from the Phase 1B Review.
The Phase 1B Review focused on the findings of the Phase 1A Review.
It was agreed that pre-beam “Hot Check-Out” would proceed, under control of JLab line management and subject to established safety processes, while the ARR process unfolded.
The ARR team agreed with this approach.

## Slide 5: F2. Close out upgrade activities
Phase 1A Finding:
Close out completed upgrade related activities needed for initial commissioning operation with beam.
Observations:
All safety-credited systems are ready.
The remaining incomplete hardware and tasks are well defined and are being tracked to completion.
Status: Pre-start finding:
Use Hot Checkout, ATLis, and associated processes to manage remaining work needed for initial commissioning.

## Slide 6: F6. Beam Authorization
Phase 1A Finding:
Formalize how controls needed to ensure operations within the ASE are incorporated into the beam operations authorization process.
Observations:
A new software tool has been developed to capture credited controls and track readiness for authorizing beam operations.
Status: Complete.

## Slide 7: F8. CEBAF Element Database
Phase 1A Finding:
Populate the CEBAF Element Database with all elements needed for initial commissioning.
Observations:
The CED is a powerful tool for ensuring consistency among optical design, hardware implementation, and controls development.
Status: Complete for initial commissioning.

## Slide 8: F9. Cyber Security of Conventional Facilities
Phase 1A Finding:
The cyber security and SQA of Conventional Facilities (CF) control systems should be reviewed and steps taken to manage remote access.
Observations:
A formal review covering these issues was conducted on August 23, shortly after the ARR Phase 1A review.
Findings from that review are being tracked in the CATS system.
Status: Complete.

## Slide 9: F10. Conventional Facilities Controls
Phase 1A Finding:
Changes to controls software of CF systems that directly serve the accelerator should be planned, coordinated, and communicated with attention to cyber security and SQA similar to that of other accelerator systems.
Observations:
These concerns were addressed in the August 23 review.
Findings from that review are being tracked in the CATS system.
Status: Complete.

## Slide 10: 12GeV Accelerator Readiness Review, Phase 1B
Ian Evans, SLAC
October, 2013

## Slide 11: Finding
F.1 Populate Hot Checkout (HCO) Checklist data base with 100% of required elements needed for Phase1 start up & commissioning.
Observation:
Reviewed Hot Checkout Database and remaining open items.
  Checklists are written by the subject matter experts
  Approved and published by the group leader
  Independently reviewed by the HCO team
Status: Pre-Start Finding
Continue to populate and implement checklists

## Slide 12: Finding
F.3 Assess whether there are any other instrument or equipment specific commissioning plans, in addition to the radiological controls and surveys, that need to be integrated into the 12 GeV Commissioning Plan. Finalize and approve the plan
Observation:
Reviewed 12 GeV Commissioning Plan
Radiological hold points have been added
Details added on beam based calibrations and commissioning
Status:
Closed
  Plan approved and issued

## Slide 13: Finding
F.4 Update and approve all procedures for Phase 1 start up and commissioning.
Observation:
28 procedures identified as “required” for beam operations.
All have been revised for 12 GeV operations
All have been issued/released by Operations Document Coordinator
Status:
Closed

## Slide 14: Finding
F.5 Include guidance in the 12 GeV Beam Commissioning Plan on how to proceed during the execution phase if a step is found to be unclear or if steps need to be changed to continue commissioning.
Observation:
Accelerator Operations Directives (AOD) and the AOD Supplement for 12GeV Commissioning provides necessary guidance on program creation, execution and modification.
The Beam Commissioning Plan has been updated to include references to these over-arching documents.
Required training for identified staff includes course modules on the AOD and AOD Supplement.
Status:
Closed

## Slide 15: Finding
F.7 Define the minimum required training elements for commissioning operations and ensure operations staff (operators, accelerator physicists etc.) are fully trained
Observation:
Minimum training elements have been defined for both Accelerator Operators and Scientific support staff.
Staff required for operating the machine have been identified.
Training has been completed for all identified individuals.
Status:
Closed

## Slide 16: Conclusion

## Slide 17: Radiation Control
Jim Floyd

## Slide 18: F11. Complete/Integrate RadCon Plans
1A Finding:
Complete the plan for radiological control and survey activities needed for commissioning and integrate it into the Commissioning Plan.
Observations:
Monitoring plan has been completed
  Good integration with Physics and Operations
Plan has been integrated into Commissioning Plan
ATLis work items are created
Status: Closed

## Slide 19: F12. Develop Shielding Process
1A Finding:
JLab should ensure that all parts of the shielding analysis, review, approval, validation, installation, and configuration control are integrated for each of the processes (experiment, temporary, and permanent).
Observations:
Procedure has been drafted
  Scope includes all shielding (experiment, temporary, permanent)
  Integrates all elements (analysis, review, construction, validation)
  Addresses concern from 1A review
Status: Continue development (Post-Start)

## Slide 20: Resolution of ARR 1A Finding
Dennis Parzyk

## Slide 21: ARR 1A VOIP Risk Assessment
ARR 1A Finding:
The Emergency Management Team should perform a risk assessment to evaluate emergency communications capabilities in light of the move to VOIP telephones.

## Slide 22: VOIP Risk Assessment, cont’d.
ARR 1B Review:
Jefferson Lab has performed a risk assessment of emergency communications, a mitigation plan developed and a training program initiated.
The risk assessment concluded that providing multi cellular repeaters was the preferred alternative with signals going straight to a cell tower.

## Slide 23: VOIP Risk Assessment, cont’d.
ARR 1B Review, cont’d:
Efforts continue to replace tunnel wiring, deploy cell phone repeaters, and investigate the possibility of a separate router.
Jefferson lab will continue with plans to provide VOIP service on Smart Phones.
IT and ESH&Q organizations are tasked with developing training regarding the communication options available for emergency communications.

## Slide 24: VOIP Risk Assessment, cont’d.
ARR 1B Conclusion:
The Finding on the VOIP Risk Assessment is considered closed with the completion of the aforementioned risk assessment and initiation of the associated mitigation steps.

## Slide 25: Charge to the Committee: Six Issues
1. Commissioning and routine operational procedures necessary for the safe operation of the activity have been developed, reviewed, and approved, and an appropriate process for the development, review, and approval of new and revised procedures is in place.
August: Not yet. All the routine operational procedures necessary for the safe commissioning are not yet ready; however, a program is in place to identify and manage the preparation of all procedures needed to begin safe commissioning.
October: Yes!

## Slide 26: Charge to the Committee: Six Issues
2. Procedures to deal with abnormal and emergency situations have been prepared, reviewed, and approved for use.
August: Yes. Procedures for dealing with general emergency situations are ready; however, reliability of critical communications infrastructure should be evaluated.
October: Yes!

## Slide 27: Charge to the Committee
3. Records important for pre-operational, operational, and post-operational activities are adequately managed and controlled.
August: Yes! Processes for managing and controlling critical records are in place.

## Slide 28: Charge to the Committee
4. Equipment and systems having safety significance meet criteria established in the SAD and are fully operational and are managed as part of the laboratory's configuration management process.
August: Not yet. Many systems are ready, but some are not ready yet.
Configuration management of safety systems is effective.
October: Yes!

## Slide 29: Charge to the Committee
5. Training and qualification programs necessary for safe accelerator operations are in place in compliance with the ASO and JLab Integrated Safety Management (ISM) program.
August: On going. Systems for managing training activities are ready, and training activities have started; however, most operations staff have not yet received minimum refresher training.
October: Yes!

## Slide 30: Charge to the Committee
6. Staffing requirements specified in the ASE are met.
August: Yes! The ASE requires at least one operator on duty in MCC when accelerator is operating.

## Slide 31: Conclusions of Phase 1
The 12 GeV Upgrade Project is nearly ready for beam, pending close out of a few remaining items.
The processes for managing this work are well developed and on track for successful completion within a few days.
The team has prepared a letter to Laboratory Director Dr. Hugh Montgomery recommending that Jlab and the DOE Site Office partner to approve initial commissioning:
Full duty cycle cw beam in the injector.
Tune mode beam to R2 dumplette.

## Slide 32: Conclusions of Phase 1
Preparations for multi-pass operation are progressing well.
Credited safety systems for multi-pass operation are ready.
The team recommends that JLab management and the DOE Site Office partner to approve multi-pass operation with tune-mode beam as soon as remaining installation work is complete.

## Slide 33: Closing Thoughts
The ARR Team appreciates the efforts by the JLab staff to prepare for this review.
The presentations were clear and comprehensive.
The supporting material was well organized and made available to the team members.
Thanks to all the JLab staff who helped make this visit an enjoyable and educational experience.
Special thanks to Bob May for making this review process enjoyable and productive.