
	

	[image: image1.png]Je son Lab

jomas Jefferson National Accelerator Facility




	Laser Operational Safety Procedure (LOSP)     THA Guidance
(See ES&H Manual Chapter 6410 Appendix T1 
Laser Operational Safety Procedure)

	


1.0 Purpose
To provide the Laser System Supervisor (LSS) with a guide to help develop the Task Hazard Analysis (THA) and the associated Laser Operational Safety Procedure (LOSP). 

2.0 Scope
This guidance document is intended for use with Class 3B and Class 4 laser systems, but can also be used for lower powered laser systems. It is non-mandatory.  

3.0 Hazard Identification 
Development of the LOSP starts with creating the THA. Although each laser system is unique, there are hazards common to Class 3B or Class 4 lasers. The following THA (pages 2-6) captures many of the hazards and can be used to develop your THA. It is not intended to be cut and pasted into your THA, but should be tailored to the specific laser system. 

4.0 Work Planning
After the completion of the THA, the LOSP can be developed. The first page identifies the location and describes the project. List the lasers to be used in the system, record the wavelength power level, and laser class on the form. Referring to the information on the THA complete the LOSP form. Use the explanatory information found in the marked up LOSP starting on page 7 of this document. 
	
	

	[image: image2.png]Wgﬁ:ﬁ)on Lab 5
lomas Jefferson National Accelerator Facility




	Task Hazard Analysis (THA) Worksheet

(See ES&H Manual Chapter 3210 Appendix T1 
Work Planning, Control, and Authorization Procedure)
	

	
	


	THA must be accompanied by a LOSP

	Author:
	Marc McMullen
	Date: 10/20/2022

	Task Title:
	Operation of the DIRC 3B laser
	Task Location:
	EEL 108 DIRC laser lab

	Engineering Control requirements:
	Administrative requirements:

	Laser enclosure with interlocks for tooled housing.
Interlocked Laser Controlled Area (LCA) for room with multiple laser hazards not inside a Class 1 enclosure.

	Training: 

· SAF114O Laser Safety Orientation

· MED02 Laser eye exam

· SAFXXX Laser Specific Training

· SAF104 Lock, Tag, Try for service & maintenance

· SAFXXX other training necessary for identified hazards (i.e. Nanomaterial, ODH,  Chemical use)
	Postings/Labels/Warning systems:

· LCA warning signs at entrance

· Laser Class and aperture labels

· Laser on indicator (yellow flashing light)




	Laser Process
	Potential Hazards

Select hazard applicable to your LCA (delete non-applicable or add other hazard if not listed)
	Consequence Level
	Probability Level
	Risk Code (before mitigation)
	Proposed Mitigation
(Required for Risk Code >2)

Select control applicable to your LCA (delete non-applicable or add other mitigation if not listed)
	Safety Procedures/ Practices/Controls/Training

Select control applicable to your LCA (delete non-applicable or add other mitigation if not listed)
	Risk Code
(after mitigation)

	OPERATIONS
	Temporary or permanent eye damage
	M
	M
	3
	Standard operation (quartz crystal measurement) is done from outside LCA

For alignment of components done inside LCA, the interlock system requires a filter to be positioned which reduces the laser power to less than MPE. 
	Engineered controls: Interlock system

Authorized admittance

Required laser training

Laser specific protective eyewear
	1

	OPERATIONS
	Specular reflection hazard
	M
	M
	3
	The LCA is a closed room with interlocked controls. 

Standard operations are done from outside the LCA. No exposure is possible due to the interlock system. 
	Engineered controls: interlock system

No jewelry, watches, portable electronics, or reflective clothing
	1

	
	Exposure to high voltages (i.e. pockel cells)
	M
	L
	2
	Electrically insulated connections
	Installation of electrical components will be done by QEW
	1

	
	Chemical exposure (skin /eyes)
	L
	L
	1
	Safety Data Sheet review by Industrial Hygiene (IH) 
	Controls per IH based on SDS review
	1

	
	Human Error
	M
	M
	3
	A robust interlock system is designed prevent the human error factor

Apply the hierarchy of controls

Identify and manage error precursors
	Prevention through design should be applied up front to eliminate or reduce hazards

Incorporate one or more of the HPI tools (such as Three-way Communication, Independent Verification, Pre-task Brief, Procedure Use & Adherence, Job Site review, etc.)
	1

	
	Temporary or permanent eye damage
	M
	M
	3
	Optical beamline layout at or below waist level

The controls prevent beam-on exposure without a power reducing filter that lowers the laser power to = Class 2 

Terminate the beam at the end of its useful path.
	Engineered controls: interlock system with low power filter

Laser specific protective eyewear

Locate any specular reflections of the beam and block them as close to the source as possible, before moving to the next optical component or section
Use Beam blocks behind optics (mirrors) to terminate beam in case of misalignment

All beams and reflections are properly terminated prior to high power operation
	1

	ALIGNMENT
	Specular reflection hazard
	M
	M
	3
	The controls prevent beam-on exposure without a power reducing filter that lowers the laser power to = Class 2 

Optical beamline layout, beam stops
	Engineered controls: interlock system with low power filter

Laser specific protective eyewear


	1

	
	Manual measurement: Specular reflection hazard, human error
	M
	M
	3
	The controls are set to expert mode. Only personnel designated by the LSS and approved by the LSO may configure the set-up for manual measurements

Expert mode measurements are only to be done after the alignment procedure (in low power configuration)


	Laser specific protective eyewear

Additional training provided by the LSS

Manual measurement procedure adherence 
	2

	
	Trip/fall due to low light or limited vision due to laser PPE
	M
	L
	2
	Clear walkways and effective housekeeping

Walkway lighting
	LSS will ensure walkways are clear


	1

	
	Cleaning of laser components and optics
	M
	M
	3
	All manipulations of optical table equipment will be done with the laser off or in Bypass mode which reduces the laser output below MPE
	Engineered controls: interlock system with low power filter

DIRC laser alignment procedure
	2

	
	Replacement of components of parts
	M
	M
	3
	All manipulations of optical table equipment will be done with the laser off or in Bypass mode which reduces the laser output below MPE
	Engineered controls: interlock system with low power filter

DIRC laser alignment procedure
	2

	
	
	
	Highest Risk Code before Mitigation:
	3
	
	Highest Risk Code after Mitigation:
	2


	

	[image: image3.png]Je son Lab

jomas Jefferson National Accelerator Facility




	Laser Operational Safety Procedure (LOSP) Form
(See ES&H Manual Chapter 6410 Appendix T1 
Laser Operational Safety Procedure)

	


	Serial Number:
	
	

	
	(Assigned by ESH&Q Document Control x7277)
	

	
	
	

	*Attach the Task Hazard Analysis (THA) related to this procedure

	Issue Date:
	
	Expiration Date:
	

	Title:
	Operation of the DIRC 3B laser

	Location where the laser will be used:
	EEL 108 DIRC laser room

	Description of Project
	The DIRC laser will be used to evaluate synthetic fused silica bars received from SLAC for their mechanical and optical quality

	Document Owner(s):

Usually the LSS, but could be someone else
	Benedikt Zihlmann, Marc McMullen
	Date:
	10/20/2022

	
	
	
	
	


	Laser Inventory

	Laser Serial #
	Laser Class
	Wavelength(s)
	Maximum Power/Energy

	1. 
	S5935
	
	3B
	
	325nm/442nm
	
	14mW/44mW

	2. 
	
	
	
	
	
	
	

	3. 
	
	
	
	
	
	
	

	4. 
	
	
	
	
	
	
	

	


	Approval Signatures:
	Print




Signature
	Date:

	Laser System Supervisor:
	Benedikt Zihlmann
	
	
	

	Laser Safety Officer:
	Jennifer Williams
	
	
	

	Division Safety Officer
	
	
	
	

	Department or Group Head:
	
	
	
	

	Other Approval(s):
	Class 4 will require Fire Protection SME
	
	
	

	
	Where applicable, SMEs for LTT, IH, High Voltage, etc.
	
	
	

	
	
	
	
	

	
	
	
	
	


	Document History:

	Revision:
	Reason for revision or update:
	Serial number of superseded document

	
	
	


Distribution: ESH&Q Document Control (x7277, MS6B); affected area(s); Document Owner; Division Safety Officer

	Introduction: In areas containing more than one laser, define operational sequence or parameters.


· The DIRC laser is located in EEL 108
· The operational use of the laser is to evaluate synthetic fused silica bars received from SLAC for their mechanical and optical quality
· There is a single Kimmon model IK5351R-D class 3b laser located in the LCA
· Include a diagram or a floor plan for the LCA 
	Personnel

Only those authorized by the LSS are permitted to enter the location noted on the cover sheet of this document.
	List:

· Training and qualification requirements (including refresher training).

· Medical requirements. 

· Spectator protection requirements. 
	


· Indicate the Laser System Supervisor (LSS) and contact information. 

· Laser Operators must complete the following training: 

· SAF1140 – Laser Safety Orientation

· MED02 – Laser Eye Exam

· SAFXXX – Laser Specific Training for the DIRC laser
· GERT
· Will spectators be allowed? Yes, they must adhere to the following stipulations:  
· LSS approval is required
· Spectator shall be escorted at all times by a qualified laser operator
· Spectator shall be briefed on the hazards within the LCA and the controls in place to mitigate the hazards
	Laser

	Define:

· Laser system specifications 
· Laser system components
· Copy of laser operating manuals or reference the location of the manual(s). 


· Provide a detailed description of what the essential system consists of: 

The laser set up is on a 5’ wide x 10’ long optical table which is ~3’ tall. The beam path will stay level at a height of about 3 to 5 inches above the surface of the optical table. Blocking panels ensure the beam does not extent past the limits of the table top. The panels are interlocked to prevent laser output when the panels are down. 
[image: image4.emf]
Figure 1. Optical table rendering with components. The beam path is in violet.

The schematic of the setup is shown in the figure above. All components are mounted on an optical table.

The He-Cd laser operates at two wavelengths, 325nm (“UV”) and 442nm (“blue”), which are selected with a manual switch at the front of the laser. The laser housing is mounted on a stage to adjust the height and assure the beam is parallel to the table surface. A pipe takes the hot air from the laser cooling fan away from the setup. At the laser exit, the UV line has an output power of up to 14mW, a beam diameter below 1mm, and a beam divergence of 0.5mrad. The blue line has an output power of up to 44mW, a beam diameter below 1mm, and a beam divergence of 0.5mrad.

UV-enhanced aluminum mirrors, coated with MgF2, guide the laser beam in the optical plane, which is parallel to the table surface. The Lambda/2 plate changes the orientation of the polarization plane of the beam.

The laser beam will be 3 to 6 inches above the surface of the table (36”). All segments of the beam will terminate at either an absorber or photo diode sensor. 

All walls of the LCA will have at least one emergency stop to shut off the laser output. 

· List each laser in use in your LCA, include the following information (examples below):

	Laser Description
	Kimmon IK5351R-D 

	Laser Type 
	Helium Cadmium (He-Cd)

	Laser Class
	Class 3B/IIIb

	Laser Manufacturer
	Kimmon

	Laser Model Number(s) 
	IK5351R-D

	Laser Serial Number(s)
	S5935

	Wavelength Range
	(325nm/442nm)

	Power Range
	14mW/44mW

	Mode (i.e., time structure)
	CW

	Beam Diameter (collimated, typical)
	0.9/1.0 mm

	Divergence (uncollimated, typical)
	< 0.5 mrad*2 +/- 10%

	Laser Eyewear O.D.
	+5


	Hazards and Mitigation

	Define:

· Laser system hazards
· Occupational exposure hazards beyond laser light (e.g., fumes, noise, etc.)
· Credible non-beam hazards (e.g., environmental hazards)
· Describe all required personal protective equipment (include clothing requirements, e.g. no reflective jewelry, etc.). Refer to ES&H Manual Chapter 6410 Appendix T2 Laser Personal Protective Equipment (PPE)
· Control of Hazardous Energy (includes beam and non-beam hazards such as electrical) 


Hazards:

· Specular reflection – mitigated by an interlocked door, interlock system controls, Jlab-required training, and laser eye protection (7+)

· Temporary/permanent eye damage – mitigated by an interlocked door, interlock system controls, Jlab-required training, and laser eye protection (7+)

· HV exposure – mitigated the use of NRTL-approved equipment installed by Jlab QEW and PPE

· Chemical exposure – mitigated through the use of Jlab IH-approved cleaning chemicals, disposal practices, and PPE

· Trips/Fall – mitigated by appropriate lighting of walkways, proper housekeeping, and oversight by LSS and area safety warden

· Human error - mitigated by an interlocked door, interlock system controls, Jlab-required training, and laser eye protection (7+)

Experts will determine any modifications or repairs needed to any equipment. Electrical repairs will be completed by qualified electrical workers. Any work needed by external vendors shall be reviewed by the LSS prior to assignment
	Laser Environment 
System designs, including interlocks, require hazard evaluation review by SME.
	Define:

· Laser system hazards
· Layout of the laser controlled area and/or table. Show beam location in relation to user – waist height preferable.
· Interlock schematic (or similar), including smoke detector interlocks. 
· Room lighting conditions during laser use and alignment procedure(s).

· Targets.

· Primary and all likely beam paths (open or enclosed). 


Figure 3. Map of DIRC laser lab area
[image: image5.jpg]m O 3 M

- 0 0

Drawing not to scale

Laser lab

A :F
\ 8%
\ I
\ &g
~ 3
~

Outside
laser enable
key switch

Waste basket

PPE
storage

sjonuod
¥oopaau|

E-stop

N
e
e
inside laser E-stop
enable key
switch

Table shield

Optical table

E-stop





Figure 4. Optical table layout

[image: image6.jpg]10’

5!

Low

a)

b)

a)

power filter
\ Lasers 266nm Absorber
405 nm
/] ;& 442 nm
532 nm
635 nm
)
) d) <) 9)

c)?
-

€)

Schematic drawing of the optics setup measuring the reflection coefficient. a)
Mirrors, b) polarizing cube (for 325 nm and 442 nm additional /2 plate), c)
diaphragms, d) beam splitter, e) Brewster mirror, f) radiator, g) reference diode,

h) value diode.





	Written procedure for use and alignment

	Provide:

· All process steps – including unattended operation controls.

· All process steps for detailed alignment – include manufacturer’s protocols for alignment. 

· Maintenance and service

· Off-normal and emergency procedures (e.g., beam loss, fire).  


Alignment procedure:
The greatest probability of a hazard is due to specular reflection during alignment, this risk is reduced through the use of a low power filter which reduces the laser power less than or equal to the power of a class 3R laser. This alignment will dictate the laser path and interaction with components of the optical table. 

1. Alignment of the laser requires:

. personnel to visually observe the laser while it is on, inside the room

a. This requires bypass of the outside room interlock loop
a. Entry into the laser lab will trigger the interlock system to disable the laser output

a. Manual repositioning or cleaning of components on the optical table is done while the laser is disabled
a. To operate the laser, configure the laser to low power mode by positioning the filter directly in front of the beam

a. The filter will reduce laser power to less than or equal to the power of a 3R laser 

a. The table side shields should be up/in position to stop the beam from going past the extent of the table

a. The final step to enable the laser is to insert the in-room key and rotate it to “ON”
a. The laser will be enabled

b. Outside-room key control will be disabled
a. Remote control of stages can be done with laser on (in low power configuration)

Manual measurement (expert mode):
Fine tuning of the laser measurements require manual measurements at full power (without the low power filter in-line). These measurement require the laser to be on with an operator present in the room. These measurements are only authorized by designated personnel who have experience and training to perform the manual measurement. The LSS shall designate the specific personnel, and the LSO shall approved the designated personnel for manual measurement operations. The LSS shall provide specific training for manual measurements and determine if adjustments to the procedure are necessary. The “expert” shall not deviate from the manual measurement procedure.

1. The alignment procedure should already be completed
2. Rotate internal key switch to “expert” to reconfigure setup to bypass the low power filter

3. Use controls to reposition quartz bar and/or photo diode sensor to complete manual measurements

1. DAq will record and log the measurements

4. If any adjustments to the optical table equipment is required, rotate key to off prior and verify status of the laser prior to any hands-on operation

Automated measurements (standard operations):
Standard operation measurements are done outside the LCA with no personnel in the laser room. The laser can be armed from the “outside room” key switch location.

1. The operator shall enter the room to ensure it is clear of all personnel
2. Clear all interlocks
3. Insert key into the outside room key switch and rotate to “on”

4. Start automated measurements program
All alignment and manual measurements shall be done with laser safety eyewear donned. 
	Laser Controls

	Based on the information developed in the THA:
· Describe all controls (administrative and engineering) required to be in place during laser operations. 
· If a different control is recommended, describe the rationale for not using a typical/recommended control.  


An engineered control based interlock system has been developed with the following elements:

· Key control switches

· magnetically locked door with contacts

· Emergency stop switches
· Interlock shields for the table sides prevent the laser beam from extending beyond the table

· A filter to reduce the laser power during alignment

The system is designed to disable the laser output if any element triggers the interlock. The system is designed to latch in a way that forces the operator to affirm all interlocks are cleared before resuming laser output. The system is designed to only allow laser power below MPE with the presence of humans in the LCA
Figure 5. Laser interlocks diagram
[image: image7.jpg]All emergency

stops

Outside room loop N.C.

Insided room loop

N.O.
®

N.C.

Inside-room
key control

Table

POWEL shielding

filter





	Required Calculations 

	· Maximum permissible exposure.
· Optical density.
· Nominal hazard zone.  


	Labeling/Posting

See ES&H Manual Chapter 6410 Appendix T6 Laser Labeling/Posting Requirements

	· Equipment/area labeling/posting requirements
· Area signs  


· Signage will be posted for a Class 3B/IIIb laser at the entry way of the area and doorway into the LCA
	Authorized/Trained Individuals

	Print Name/Signature
	Date

	
	

	
	

	
	

	
	

	
	

	
	


Page 6 of 14

