## CLEOII-Chassis Layout
PLC I/O Usage Table
Magnets | Quench Detector / PSU | Temperature sensors | Temperature sensors | CL mass Flow Controller | JT motor relays | Reset and Pwr enabled | Load Sensors | Load Sensors | 8 Channels
Module | IV16D | IF16 | IF16 | OF4 | OW16I | OW16I | IF16 | IF16 | IF16
Configuration |  | Single_Ended voltage | Single_Ended Voltage | 0-5 volts output |  |  | Single_Ended_Mode Voltage | Single_Ended_Mode Voltage/Current | Differential_Mode Voltage
Chasis Slot # | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
Channel
0 | QD_U1 | TS_1 | TS_27_CL_Neg | Setpnt_m_flow_Neg | MOT_JTV1_Close | Danf_QD_Reset | Radial - A | Axial - T | Current Lead A
1 | QD_L1 | TS_2 | TS_28_CL_POS | Setpnt_m_flow_Pos | MOT_JTV1_Open | Vacuum_power | Radial - B | Axial - V | Current Lead B
2 | QD_U2 | TS_3 | TS_29 (LN2) |  | MOT_JTV2_Close | Keep_Alive_Reset | Radial - C | Axial - W | Coil 2
3 | QD_L2 | TS_4 | TS_30 (LN2) |  | MOT_JTV2_Open | Keep_Alive_Control | Radial - D | Axial - X | Coil 1
4 | QD_U3 | TS_5 | TS_31 (LN2) | N/A | MOT_JTV3_Close | Power_liquid_level_meter | Radial - E | LL_ He | Magnet
5 | QD_L3 | TS_6 | TS_32 (LN2) | N/A | MOT_JTV3_Open | PLC_Fast_Relay | Radial - F | LL_N2 | Current Loop 1
6 | QD_U4 | TS_21 | TS_12 (LN2) | N/A | MOT_JTV4_Close | PLC_Slow_Relay | Radial - G | Mass_Flow_CL_Neg | Current Loop 2
7 | QD_L4 | TS_22 | TS_13 (LN2) | N/A | MOT_JTV4_Open | MPS_CPU_Reset | Radial - H | Mass_Flow_CL_Pos
8 | QD_Sum | TS_23 | TS_9 (LN2_Supply) | N/A | MOT_JTV5_Close |  | Radial - J | LVDT_JTV1 | N/A
9 |  | TS_24 | TS_10 (LN2_Return) | N/A | MOT_JTV5_Open |  | Radial - K | LVDT_JTV2 | N/A
10 |  | TS_25 | TS_11 (LN2_Center) | N/A | MOT_JTV6-Close |  | Radial - L | LVDT_JTV3 | N/A
11 | JVT Mtrs KS remote | TS_26 |  | N/A | MOT_JTV6_Open |  | Radial - M | LVDT_JTV4 | N/A
12 | PSU_Control_pwr_on | TS_33 |  | N/A | Warm_return_Close |  | Radial - N | LVDT_JTV5 | N/A
13 | PSU_Main_pwr_on | TS_34 | Vacuum | N/A | Warm_return_Open |  | Radial - P | LVDT_JTV6 | N/A
14 | PSU_ready | TS_7_He_Supply | Pressure He | N/A |  |  | Radial - R | Warm_Return_Position | N/A
15 | PSU_Sum_Interlocks | TS_8_He_Return | Pressure LN2 | N/A |  |  | Radial - S | I_coarse | N/A
Dataforth Signal Conditioner Table SCM5B modules
Channel -> board # |  | 1 | 3 |  |  |  | 5 | 7 | 8
1 |  | 30-01D [+/- 10mV +/- 10V] | 30-01D [+/- 10mV +/- 10V] |  | JT1 LN2 Cooldown |  | 38-31D [+/-10mV - +/-3.33V] | 38-31D [+/-10mV - +/-3.33V] | 31-09D [+/-40V - +/-10V]
2 |  | 30-01D [+/- 10mV +/- 10V] | 30-01D [+/- 10mV +/- 10V] |  | JT2 LN2 Top Fill |  | 38-31D [+/-10mV - +/-3.33V] | 38-31D [+/-10mV - +/-3.33V] | 31-09D [+/-40V - +/-10V]
3 |  | 30-01D [+/- 10mV +/- 10V] | 30-01D [+/- 10mV +/- 10V] |  | JT3 He Cooldown |  | 38-31D [+/-10mV - +/-3.33V] | 38-31D [+/-10mV - +/-3.33V] | 31-09D [+/-40V - +/-10V]
4 |  | 30-01D [+/- 10mV +/- 10V] | 30-01D [+/- 10mV +/- 10V] |  | JT4 He Bottom Fill |  | 38-31D [+/-10mV - +/-3.33V] | 38-31D [+/-10mV - +/-3.33V] | 31-09D [+/-40V - +/-10V]
5 |  | 30-01D [+/- 10mV +/- 10V] | 30-01D [+/- 10mV +/- 10V] |  | JT5 He Top Fill |  | 38-31D [+/-10mV - +/-3.33V] |  | 31-09D [+/-40V - +/-10V]
6 |  | 30-01D [+/- 10mV +/- 10V] | 30-01D [+/- 10mV +/- 10V] |  | JT6 He Cold Return |  | 38-31D [+/-10mV - +/-3.33V] |  | 31-09D [+/-40V - +/-10V]
7 |  | 30-01D [+/- 10mV +/- 10V] | 30-01D [+/- 10mV +/- 10V] |  | Warm return Valve |  | 38-31D [+/-10mV - +/-3.33V] |  | 31-09D [+/-40V - +/-10V]
8 |  | 30-01D [+/- 10mV +/- 10V] | 30-01D [+/- 10mV +/- 10V] |  |  |  | 38-31D [+/-10mV - +/-3.33V] |  | Used
Channel -> board # |  | 2 | 4 |  |  |  | 6 | LVDT Modules
1 |  | 30-01D [+/- 10mV +/- 10V] | 30-01D [+/- 10mV +/- 10V] |  |  |  | 38-31D [+/-10mV - +/-3.33V] | LVC-2500
2 |  | 30-01D [+/- 10mV +/- 10V] | 30-01D [+/- 10mV +/- 10V] |  |  |  | 38-31D [+/-10mV - +/-3.33V] | LVC-2500
3 |  | 30-01D [+/- 10mV +/- 10V] | 30-01D [+/- 10mV +/- 10V] |  |  |  | 38-31D [+/-10mV - +/-3.33V] | LVC-2500
4 |  | 30-01D [+/- 10mV +/- 10V] |  |  |  |  | 38-31D [+/-10mV - +/-3.33V] | LVC-2500
5 |  | 30-01D [+/- 10mV +/- 10V] |  |  |  |  | 38-31D [+/-10mV - +/-3.33V] | LVC-2500
6 |  | 30-01D [+/- 10mV +/- 10V] | 30-01D [+/- 10mV +/- 10V] |  |  |  | 38-31D [+/-10mV - +/-3.33V] | LVC-2500
7 |  | 30-01D [+/- 10mV +/- 10V] | 30-01D [+/- 10mV +/- 10V] |  |  |  | 38-31D [+/-10mV - +/-3.33V] | LVC-2500
8 |  | 30-01D [+/- 10mV +/- 10V] | 30-01D [+/- 10mV +/- 10V] |  |  |  | 38-31D [+/-10mV - +/-3.33V]

## Temp Sensors
Signal description | Range |  | CCL sensor # | Description |  | Internal Cable # | Wire color | Internal connector / pin | Measured room temperature resistance |  |  |  | Internal Cable # | Wire Color | Vacuum Feed Thru Connector / Pin | Wire color | Terminal Strip | Dataforth SCM5B30-01D | Dataforth feed out | Constant Current Source | Wire Color | Terminal Strip | 1756 IF16 | Temperature | Resistance | Δ T [K] | Voltage from a 100uA | Power | Coefficients
[8 empty] | 25 pin D connectors | Fluke 87 III S# 82800752 with remote term blk db 25 pin patch panel (102FA) | [5 empty] | 10 pin vacuum feedthrough |  |  | +/- 10mV in +/- 10V out |  | 100uA | [6 empty] | mV | uW | Rodium Fe sensors
Temperature |  |  |  | Temperature sensors |  |  |  | 5 | Pin to Pin | Residual terminal block resistance |  | Ohms | [5 empty] | pin # | Board # / Pin # | 8 Channels per card |  |  |  | K | R |  | 0.0001 |  | Table 7.1 Page 75
Temperature Sensor He Coil Shell | 300 - 4 K | [7 empty] | Pin to Pin | Voltmeter lead resistance null out | Total resistance | Probe resistance |  |  |  |  | TS-02 |  |  | J1 |  | TS-PLC-02 | slot #/ Channel # / pin # | [5 empty] | x^2 | x | b
Rhodium/Fe 4-wire PH20002
Coil Outer shell |  | 1 | 1552 | TS1 | V + | 13 | B | H1 - 1 | 1-3 | 2.34 | 32.54 | 30.2 | 50 | B | D1 - A |  | 0-1 | 3 | 1 / 1 |  |  | 0 - 1 | 2:I.ch0+ / 2 | m = 10 b = 1 |  |  |  |  | -0.21783 | 16.9703 | -26.2339
[5 empty] | I + | 13 | BR | H1 - 2 | 2-4 | 2.43 | 32.59 | 30.16 | 50 | BR | D1 - B |  | 1-1 |  |  | CH 1+ |  |  |  | 4.22 | 1.8379 | -0.0115 | 0.18379 | 0.018379
Train gauge |  |  |  |  | V - | 13 | B | H1 - 3 | 1-4 | 2.35 | 32.6 | 30.25 | 50 | B | D1 - C |  | 0-2 | 2 | 1 / 3 |  |  | 0 - 2 | RTN / 9 | 77.35 | 6.6759 | -0.024 | 0.66759 | 0.066759
[5 empty] | I - | 13 | BR | H1 - 4 | 2-3 | 2.47 | 32.62 | 30.15 | 50 | BR | D1 - D |  | 1-2 |  |  | CH 1 - |  |  |  | 273.2 | 27.006 |  | 2.7006 | 0.27006
Coil Outer shell |  | 2 | 1553 | TS2 | V + | 14 | B | H1 - 6 | 6-8 | 2.37 | 32.16 | 29.79 | 51 | B | D1 - F |  | 2-1 | 3 | 1 / 5 |  |  | 1 - 1 | 2:I.ch1+ / 4 | m = 10 b = 1 |  |  |  |  | -0.21713 | 16.946 | -26.1027
[5 empty] | I + | 14 | R | H1 - 7 | 7-9 | 2.35 | 32.1 | 29.75 | 51 | R | D1 - H |  | 3-1 |  |  | CH 2+ |  |  |  | 4.22 | 1.8324 | -0.054 | 0.18324 | 0.018324
[5 empty] | V - | 14 | B | H1 - 8 | 6-9 | 2.37 | 32.11 | 29.74 | 51 | B | D1 - J |  | 2-2 | 2 | 1 / 6 |  |  | 1 - 2 | RTN / 9 | 77.35 | 6.6759 | -0.025 | 0.66759 | 0.066759
[5 empty] | I - | 14 | R | H1 - 9 | 7-8 | 2.38 | 32.14 | 29.76 | 51 | R | D1 - K |  | 3-2 |  |  | CH 2 - |  |  |  | 273.2 | 27.009 |  | 2.7009 | 0.27009
Coil Outer shell |  | 3 | 1554 | TS3 | V + | 15 | B | H1 - 14 | 14-16 | 2.39 | 32.01 | 29.62 | 52 | B | D1 - L |  | 4-1 | 3 | 1 / 7 |  |  | 2 - 1 | 2:I.ch2+ / 6 | m = 10 b = 1 |  |  |  |  | -0.21673 | 16.9387 | -26.2603
[5 empty] | I + | 15 | O | H1 - 15 | 15-17 | 2.34 | 31.96 | 29.62 | 52 | O | D2 - A |  | 5-1 |  |  | CH 3+ |  |  |  | 4.22 | 1.8429 | 0.027 | 0.18429 | 0.018429
[5 empty] | V - | 15 | B | H1 - 16 | 14-17 | 2.39 | 32.01 | 29.62 | 52 | B | D2 - B |  | 4-2 | 2 | 1 / 9 |  |  | 2 - 2 | RTN / 9 | 77.35 | 6.6893 | 0.114 | 0.66893 | 0.066893
[5 empty] | I - | 15 | O | H1 - 17 | 15-16 | 2.35 | 31.97 | 29.62 | 52 | O | D2 - C |  | 5-2 |  |  | CH 3 - |  |  |  | 273.2 | 27.021 |  | 2.7021 | 0.27021
Coil Outer shell |  | 4 | 1556 | TS4 | V + | 16 | B | H1 - 19 | 19-21 | 2.36 | 32.36 | 30 | 53 | B | D2 - D |  | 6-1 | 3 | 1 / 11 |  |  | 3 - 1 | 2:I.ch3+ / 8 | m = 10 b = 1 |  |  |  |  | -0.21719 | 16.9442 | -26.0267
[5 empty] | I + | 16 | Y | H1 - 20 | 20-22 | 2.38 | 32.33 | 29.95 | 53 | Y | D2 - E |  | 7-1 |  |  | CH 4+ |  |  |  | 4.22 | 1.8279 | -0.0115 | 0.18279 | 0.018279
[5 empty] | V - | 16 | B | H1 - 21 | 19-22 | 2.35 | 32.37 | 30.02 | 53 | B | D2 - F |  | 6-2 | 2 | 1 / 12 |  |  | 3 - 2 | RTN / 9 | 77.35 | 6.6715 | -0.07 | 0.66715 | 0.066715
[5 empty] | I - | 16 | Y | H1 -22 | 20-21 | 2.4 | 32.37 | 29.97 | 53 | Y | D2 - H |  | 7-2 |  |  | CH 4 - |  |  |  | 273.2 | 27.012 |  | 2.7012 | 0.27012
Coil Outer shell |  | 5 | 1557 | TS5 | V + | 17 | B | H1 - 11 | 11-23 | 2.39 | 32.61 | 30.22 | 54 | B | D2 - K |  | 8-1 | 3 | 1 / 13 |  |  | 4 - 1 | 2:I.ch4+ / 12 | m = 10 b = 1 |  |  |  |  | -0.21687 | 16.9366 | -25.9815
[5 empty] | I + | 17 | G | H1 - 12 | 12-24 | 2.38 | 32.54 | 30.16 | 54 | G | D2 - L |  | 9-1 |  |  | CH 5+ |  |  |  | 4.22 | 1.8259 | -0.1048 | 0.18259 | 0.018259
[5 empty] | V - | 17 | B | H1 - 23 | 11-24 | 2.38 | 32.58 | 30.2 | 54 | B | D3 - A |  | 8-2 | 2 | 1 / 15 |  |  | 4 - 2 | RTN / 9 | 77.35 | 6.6709 | -0.076 | 0.66709 | 0.066709
[5 empty] | I - | 17 | G | H1 -24 | 12-23 | 2.36 | 32.58 | 30.22 | 54 | G | D3 - B |  | 9-2 |  |  | CH 5 - |  |  |  | 273.2 | 26.998 |  | 2.6998 | 0.26998
Coil Outer shell |  | 6 | 1558 | TS6 | V + | 18 | B | H2 - 1 | 1-3 | 2.34 | 32.27 | 29.93 | 55 | B | D3 - C |  | 10-1 | 3 | 1 / 17 |  |  | 5 - 1 | 2:I.ch5+ / 14 | m = 10 b = 1 |  |  |  |  | -0.21755 | 16.9591 | -26.1271
[5 empty] | I + | 18 | BC | H2 - 2 | 2-4 | 2.43 | 32.33 | 29.9 | 55 | BC | D3 - D |  | 11-1 |  |  | CH 6+ |  |  |  | 4.22 | 1.8325 | -0.053 | 0.18325 | 0.018325
[5 empty] | V - | 18 | B | H2 - 3 | 1-4 | 2.35 | 32.23 | 29.88 | 55 | B | D3 - E |  | 10-2 | 2 | 1 / 18 |  |  | 5 - 2 | RTN / 9 | 77.35 | 6.6727 | -0.058 | 0.66727 | 0.066727
[5 empty] | I - | 18 | BC | H2 - 4 | 2-3 | 2.47 | 32.33 | 29.86 | 55 | BC | D3 - F |  | 11-2 |  |  | CH 6 - |  |  |  | 273.2 | 27.004 |  | 2.7004 | 0.27004
Coil Outer shell |  | 7 | 1559 | TS21 | V + | 19 | R | H2 - 6 | 6-8 | 2.37 | 32.07 | 29.7 | 63 | R | C1 - A |  | 12-1 | 3 | 1 / 19 |  |  | 6 - 1 | 2:I.ch6+ / 16 | m = 10 b = 1 |  |  |  |  | -0.21793 | 16.9696 | -26.0997
[5 empty] | I + | 19 | BR | H2 - 7 | 7-9 | 2.35 | 32.07 | 29.72 | 63 | BR | C1 - B |  | 13-1 |  |  | CH 7+ |  |  |  | 4.22 | 1.8297 | -0.075 | 0.18297 | 0.018297
[5 empty] | V - | 19 | R | H2 - 8 | 6-9 | 2.37 | 32.06 | 29.69 | 63 | R | C1 - C |  | 12-2 | 2 | 1 / 21 |  |  | 6 - 2 | RTN / 9 | 77.35 | 6.667 | -0.116 | 0.6667 | 0.06667
[5 empty] | I - | 19 | BR | H2 - 9 | 7-8 | 2.38 | 32.08 | 29.7 | 63 | BR | C1 - D |  | 13-2 |  |  | CH 7 - |  |  |  | 273.2 | 26.998 |  | 2.6998 | 0.26998
Coil Outer shell |  | 8 | 1562 | TS22 | V + | 20 | R | H2 - 14 | 14-16 | 2.39 | 32.41 | 30.02 | 64 | R | C1 - F |  | 14-1 | 3 | 1 / 23 |  |  | 7 - 1 | 2:I.ch7+ / 18 | m = 10 b = 1 |  |  |  |  | -0.2175 | 16.9583 | -26.0772
[5 empty] | I + | 20 | W | H2 - 15 | 15-17 | 2.34 | 32.41 | 30.07 | 64 | W | C1 - H |  | 15-1 |  |  | CH 8+ |  |  |  | 4.22 | 1.8295 | -0.076 | 0.18295 | 0.018295
[5 empty] | V - | 20 | R | H2 - 16 | 14-17 | 2.39 | 32.42 | 30.03 | 64 | R | C1 - J |  | 14-2 | 2 | 1 / 24 |  |  | 7 - 2 | RTN / 9 | 77.35 | 6.6694 | -0.091 | 0.66694 | 0.066694
[5 empty] | I - | 20 | W | H2 - 17 | 15-16 | 2.35 | 32.42 | 30.07 | 64 | W | C1 - K |  | 15-2 |  |  | CH 8 - |  |  |  | 273.2 | 26.992 |  | 2.6992 | 0.26992
J2
Coil Outer shell |  | 9 | 1564 | TS23 | V + | 21 | R | H2 - 19 | 19-21 | 2.36 | 32.64 | 30.28 | 65 | R | C1 - L |  | 16-1 | 3 | 2 / 1 |  |  | 8 - 1 | 2:I.ch8+ / 20 | m = 10 b = -1 |  |  |  |  | -0.21712 | 16.9568 | -26.3552
[5 empty] | I + | 21 | O | H2 - 20 | 20-22 | 2.38 | 32.69 | 30.31 | 65 | O | C2 - A |  | 17-1 |  |  | CH 1+ |  |  |  | 4.22 | 1.8468 | 0.057 | 0.18468 | 0.018468
[5 empty] | V - | 21 | R | H2 - 21 | 19-22 | 2.35 | 32.64 | 30.29 | 65 | R | C2 - B |  | 16-2 | 2 | 2 / 3 |  |  | 8 - 2 | RTN / 9 | 77.35 | 6.6887 | 0.107 | 0.66887 | 0.066887
[5 empty] | I - | 21 | O | H2 - 22 | 20-21 | 2.4 | 32.67 | 30.27 | 65 | O | C2 - C |  | 17-2 |  |  | CH 1 - |  |  |  | 273.2 | 27 |  | 2.7 | 0.27
Coil Outer shell |  | 10 | 1572 | TS24 | V + | 22 | R | H2 - 11 | 11-23 | 2.39 | 32.31 | 29.92 | 66 | R | C2 - D |  | 18-1 | 3 | 2 / 5 |  |  | 9 -1 | 2:I.ch9+ / 22 | m = 10 b = 1 |  |  |  |  | -0.21709 | 16.9487 | -26.1693
[5 empty] | I + | 22 | V | H2 - 12 | 12-24 | 2.38 | 32.3 | 29.92 | 66 | V | C2 - E |  | 19-1 |  |  | CH 2+ |  |  |  | 4.22 | 1.8362 | -0.0247 | 0.18362 | 0.018362
[5 empty] | V - | 22 | R | H2 - 23 | 11-24 | 2.38 | 32.3 | 29.92 | 66 | R | C2 - F |  | 18-2 | 2 | 2 / 6 |  |  | 9 - 2 | RTN / 9 | 77.35 | 6.6792 | 0.009 | 0.66792 | 0.066792
[5 empty] | I - | 22 | V | H2 - 24 | 12-23 | 2.36 | 32.32 | 29.96 | 66 | V | C2 - H |  | 19-2 |  |  | CH 2 - |  |  |  | 273.2 | 27.002 |  | 2.7002 | 0.27002
Coil Outer shell |  | 11 | 1576 | TS25 | V + | 23 | R | H3 - 1 | 1-3 | 2.34 | 32.23 | 29.89 | 67 | R | C2 - K |  | 20-1 | 3 | 2 / 7 |  |  | 10 -1 | 2:I.ch10+ / 24 | m = 10 b = 1 |  |  |  |  | -0.21762 | 16.9551 | -26.0112
[5 empty] | I + | 23 | G | H3 - 2 | 2-4 | 2.43 | 32.31 | 29.88 | 67 | G | C2 - L |  | 21-1 |  |  | CH 3+ |  |  |  | 4.22 | 1.8258 | -0.1055 | 0.18258 | 0.018258
[5 empty] | V - | 23 | R | H3 - 3 | 1-4 | 2.35 | 32.21 | 29.86 | 67 | R | C3 - A |  | 20-2 | 2 | 2 / 9 |  |  | 10 - 2 | RTN / 9 | 77.35 | 6.6666 | -0.121 | 0.66666 | 0.066666
[5 empty] | I - | 23 | G | H3 - 4 | 2-3 | 2.47 | 32.32 | 29.85 | 67 | G | C3 - B |  | 21-2 |  |  | CH 3 - |  |  |  | 273.2 | 27.013 |  | 2.7013 | 0.27013
Coil Outer shell |  | 12 | 1571 | TS26 | V + | 24 | R | H3 - 6 | 6-8 | 2.37 | 32.6 | 30.23 | 68 | R | C3 - C |  | 22-1 | 3 | 2 / 11 |  |  | 11 - 1 | 2:I.ch11+ / 26 | m = 10 b = 1 |  |  |  |  | -0.21724 | 16.9463 | -25.9931
[5 empty] | I + | 24 | BL | H3 - 7 | 7-9 | 2.35 | 32.58 | 30.23 | 68 | BL | C3 - D |  | 23-1 |  |  | CH 4+ |  |  |  | 4.22 | 1.8256 | -0.1071 | 0.18256 | 0.018256
[5 empty] | V - | 24 | R | H3 - 8 | 6-9 | 2.37 | 32.58 | 30.21 | 68 | R | C3 - E |  | 23-2 | 2 | 2 / 12 |  |  | 11 - 2 | RTN / 9 | 77.35 | 6.6683 | -0.102 | 0.66683 | 0.066683
[5 empty] | I - | 24 | BL | H3 - 9 | 7-8 | 2.38 | 32.6 | 30.22 | 68 | BL | C3 - F |  | 23-2 |  |  | CH 4 - |  |  |  | 273.2 | 27.002 |  | 2.7002 | 0.27002
SOW |  | 13 | 1570 | TS33 | V + | 25 | R | H3 - 14 | 14-16 | 2.39 | 32.7 | 30.31 | 75 | R | E3 - C |  | 24-1 | 3 | 2 / 13 |  |  | 12 - 1 | 2:I.ch12+ / 30 | m = 10 b = 1 |  |  |  |  | -0.21693 | 16.9341 | -25.9623
[5 empty] | I + | 25 | V | H3 - 15 | 15-17 | 2.34 | 32.66 | 30.32 | 75 | V | E3 - D |  | 25-1 |  |  | CH 5+ |  |  |  | 4.22 | 1.825 | -0.111 | 0.1825 | 0.01825
[5 empty] | V - | 25 | R | H3 - 16 | 14-17 | 2.39 | 32.69 | 30.3 | 75 | R | E3 - E |  | 24-2 | 2 | 2 / 15 |  |  | 12 - 2 | RTN / 27 | 77.35 | 6.6709 | -0.077 | 0.66709 | 0.066709
[5 empty] | I - | 25 | V | H3 - 17 | 15-16 | 2.35 | 32.66 | 30.31 | 75 | V | E3 - F |  | 25-2 |  |  | CH 5 - |  |  |  | 273.2 | 27.016 |  | 2.7016 | 0.27016
SOW |  | 14 | 1567 | TS34 | V + | 26 | R | H3 - 19 | 19-21 | 2.36 | 32.28 | 29.92 | 76 | R | E3 - H |  | 26-1 | 3 | 2 / 17 |  |  | 13 - 1 | 2:I.ch13+ / 32 | m = 10 b = 2 |  |  |  |  | -0.21821 | 16.9659 | -25.8559
[5 empty] | I + | 26 | GR | H3 - 20 | 20-22 | 2.38 | 32.34 | 29.96 | 76 | GR | E3 - J |  | 27-1 |  |  | CH 6+ |  |  |  | 4.22 | 1.8151 | -0.19 | 0.18151 | 0.018151
[5 empty] | V - | 26 | R | H3 - 21 | 19-22 | 2.35 | 32.3 | 29.95 | 76 | R | E3 - K |  | 26-2 | 2 | 2 / 18 |  |  | 13 - 2 | RTN / 27 | 77.35 | 6.6523 | -0.268 | 0.66523 | 0.066523
[5 empty] | I - | 26 | GR | H3 - 22 | 20-21 | 2.4 | 32.44 | 30.04 | 76 | GR | E3 - L |  | 27-2 |  |  | CH 6 - |  |  |  | 273.2 | 27.01 |  | 2.701 | 0.2701
IN Helium Return pipe |  | 15 | 1563 | TS7 | V + |  |  |  | 11-23 | 2.39 |  | -2.39 | 56 | O | D3 - H |  | 28-1 | 3 | 2 / 19 |  |  | 14 - 1 |  | m = 10 b = 1 |  |  |  |  | -0.21746 | 16.9645 | -26.2456
[5 empty] | I + |  |  |  | 12-24 | 2.38 |  | -2.38 | 56 | BR | D3 - J |  | 29-1 |  |  | CH 7+ |  |  | 2:I.ch14+ / 34 | 4.22 | 1.8392 | -0.0014 | 0.18392 | 0.018392
[5 empty] | V - |  |  |  | 11-24 | 2.38 |  | -2.38 | 56 | O | D3 - K |  | 28-2 | 2 | 2 / 21 |  |  | 14 - 2 |  | 77.35 | 6.6783 |  | 0.66783 | 0.066783
[5 empty] | I - |  |  |  | 12-23 | 2.36 |  | -2.36 | 56 | BR | D3 - -L |  | 29-2 |  |  | CH 7 - |  |  | RTN / 27 | 273.2 | 26.987 |  | 2.6987 | 0.26987
|  | 16 | 1566 | TS8 | V + | [7 empty] | 57 | O | A1 - A |  | 30-1 | 3 | 2 / 23 |  |  | 15 - 1 |  | m = 10 b = 1 |  |  |  |  | -0.21685 | 16.9433 | -26.215
[5 empty] | I + | [7 empty] | 57 | Y | A1 - B |  | 31-1 |  |  | CH 8+ |  |  | 2:I.ch15+ / 36 | 4.22 | 1.8396 | 0 | 0.18396 | 0.018396
[5 empty] | V - | [7 empty] | 57 | O | A1 - C |  | 30-2 | 2 | 2 / 24 |  |  | 15 - 2 |  | 77.35 | 6.6843 | 0.06 | 0.66843 | 0.066843
[5 empty] | I - | [7 empty] | 57 | Y | A1 - D |  | 31-2 |  |  | CH 8 - |  |  | RTN / 27 | 273.2 | 27.006 |  | 2.7006 | 0.27006
[17 empty] | TS-03 |  |  | J3 |  | TS-PLC-03 | slot 3 / pin #
|  | 17 | 1568 | TS27 | V + | [7 empty] | 69 | O | C3 - H |  | 0-1 | 3 | 3 / 1 |  |  | 0 - 1 | 3:I.ch0+ / 2 | m = 10 b = 1 |  |  |  |  | -0.21598 | 16.917 | -26.16
[5 empty] | I + | [7 empty] | 69 | G | C3 - J |  | 1-1 |  |  | CH 1+ |  |  |  | 4.22 | 1.839 | -0.002 | 0.1839 | 0.01839
[5 empty] | V - | [7 empty] | 69 | O | C3 - K |  | 0-2 | 2 | 3 / 3 |  |  | 0 - 2 | RTN / 9 | 77.35 | 6.6901 | 0.12 | 0.66901 | 0.066901
[5 empty] | I - | [7 empty] | 69 | G | C3 - L |  | 1-2 |  |  | CH 1 - |  |  |  | 273.2 | 27.009 |  | 2.7009 | 0.27009
|  | 18 | 1575 | TS28 | V + | [7 empty] | 70 | O | E1 - A |  | 2-1 | 3 | 3 / 5 |  |  | 1 - 1 | 3:I.ch1+ / 4 | m = 10 b = 1 |  |  |  |  | -0.21678 | 16.9412 | -26.2018
[5 empty] | I + | [7 empty] | 70 | BL | E1 - B |  | 3-1 |  |  | CH 2+ |  |  |  | 4.22 | 1.839 | -0.003 | 0.1839 | 0.01839
[5 empty] | V - | [7 empty] | 70 | O | E1 - C |  | 2-2 | 2 | 3 / 6 |  |  | 1 - 2 | RTN / 9 | 77.35 | 6.6841 | 0.06 | 0.66841 | 0.066841
[5 empty] | I - | [7 empty] | 70 | BL | E1 - D |  | 3-2 |  |  | CH 2 - |  |  |  |  | 27.004
Temperature Sensor N2 Outer Radiation Shield
| > 60 K | 22 | 1555 | TS29 | V + | 27 | R | N1 - 1 | 1-3 | 0.24 | 32.21 | 31.97 | 71 | R | E1- F |  | 5-1 | 3 | 3 / 7 |  |  | 2 - 1 | 3:I.ch2+ / 6 | m = 10 b = 3 |  |  |  |  | -0.21809 | 16.9642 | -25.8176
[5 empty] | I + | 27 | W | N1 - 2 | 2-4 | 0.2 | 32.28 | 32.08 | 71 | W | E1 - H |  | 4-2 |  |  | CH 3+ |  |  |  | 4.22 | 1.8129 | -0.206 | 0.18129 | 0.018129
[5 empty] | V - | 27 | R | N1 - 3 | 1-4 | 0.2 | 32.19 | 31.99 | 71 | R | E1 - J |  | 5-2 | 2 | 3 / 9 |  |  | 2 - 2 | RTN / 9 | 77.35 | 6.65 | -0.292 | 0.665 | 0.0665
[5 empty] | I - | 27 | W | N1 - 4 | 2-3 | 0.22 | 32.29 | 32.07 | 71 | W | E1 - K |  |  |  |  | CH 3 - |  |  |  | 273.2 | 26.994 |  | 2.6994 | 0.26994
6-1
|  | 23 | 1560 | TS30 | V + | 28 | O | N1 - 6 | 6-8 | 0.15 | 32.38 | 32.23 | 72 | O | E1 - L |  | 7-1 | 3 | 3 / 11 |  |  | 3 - 1 | 3:I.ch3+ / 8 | m = 10 b = 3 |  |  |  |  | -0.21839 | 16.9753 | -25.853
[5 empty] | I + | 28 | B | N1 - 7 | 7-9 | 0.15 | 32.35 | 32.2 | 72 | B | E2 - A |  | 6-2 |  |  | CH 4+ |  |  |  | 4.22 | 1.8139 | -0.198 | 0.18139 | 0.018139
[5 empty] | V - | 28 | O | N1 - 8 | 6-9 | 0.21 | 32.35 | 32.14 | 72 | O | E2 - B |  | 7-2 | 2 | 3 / 12 |  |  | 3 - 2 | RTN / 9 | 77.35 | 6.6482 | -0.311 | 0.66482 | 0.066482
[5 empty] | I - | 28 | B | N1 - 9 | 7-8 | 0.15 | 32.37 | 32.22 | 72 | B | E2 - C |  |  |  |  | CH 4 - |  |  |  | 273.2 | 26.985 |  | 2.6985 | 0.26985
8-1
|  | 24 | 1561 | TS31 | V + | 29 | O | N1 - 14 | 14-16 | 0.18 | 32.59 | 32.41 | 73 | O | E2 - E |  | 9-1 | 3 | 3 / 13 |  |  | 4 - 1 | 3:I.ch4+ / 12 | m = 10 b = 2 |  |  |  |  | -0.21785 | 16.9651 | -25.9931
[5 empty] | I + | 29 | BR | N1 - 15 | 15-17 | 0.16 | 32.55 | 32.39 | 73 | BR | E2 - F |  | 8-2 |  |  | CH 5+ |  |  |  | 4.22 | 1.8236 | -0.122 | 0.18236 | 0.018236
[5 empty] | V - | 29 | O | N1 - 16 | 14-17 | 0.11 | 32.57 | 32.46 | 73 | O | E2 - H |  | 9-2 | 2 | 3 / 15 |  |  | 4 - 2 | RTN / 9 | 77.35 | 6.6613 | -0.175 | 0.66613 | 0.066613
[5 empty] | I - | 29 | BR | N1 - 17 | 15-16 | 0.09 | 32.56 | 32.47 | 73 | BR | E2 - J |  |  |  |  | CH 5 - |  |  |  | 273.2 | 26.99 |  | 2.699 | 0.2699
10-1
|  | 25 | 1565 | TS32 | V + | 30 | O | N1 - 19 | 19-21 | 0.1 | 32.34 | 32.24 | 74 | O | E2 - J |  | 11-1 | 3 | 3 / 17 |  |  | 5 - 1 | 3:I.ch5+ / 14 | m = 10 b = 1 |  |  |  |  | -0.21722 | 16.9448 | -25.9437
[5 empty] | I + | 30 | R | N1 - 20 | 20-22 | 0.13 | 32.33 | 32.2 | 74 | R | E2 - K |  | 10-2 |  |  | CH 6+ |  |  |  | 4.22 | 1.8227 | -0.129 | 0.18227 | 0.018227
[5 empty] | V - | 30 | O | N1 - 21 | 19-22 | 0.06 | 32.31 | 32.25 | 74 | O | E3 - A |  | 11-2 | 2 | 3 / 18 |  |  | 5 - 2 | RTN / 9 | 77.35 | 6.6654 | -0.133 | 0.66654 | 0.066654
[5 empty] | I - | 30 | R | N1 - 22 | 20-21 | 0.05 | 32.34 | 32.29 | 74 | R | E3 -- B |  |  |  |  | CH 6 - |  |  |  | 273.2 | 26.997 |  | 2.6997 | 0.26997
Temperature Sensor N2 Inner Radiation Shield | > 60 K | [15 empty] | 12-1
|  | 26 | 1569 | TS12 | V + | 32 | BR | N2 - 6 | 6-8 | 0.15 | 32.26 | 32.11 | 61 | BR | A3 - A |  | 13-1 | 3 | 3 / 19 |  |  | 6 - 1 | 3:I.ch6+ / 16 | m = 10 b = -2 |  |  |  |  | -0.2168 | 16.9534 | -26.5197
[5 empty] | I + | 32 | Y | N2 - 7 | 7-9 | 0.15 | 32.25 | 32.1 | 61 | Y | A3 - B |  | 12-2 |  |  | CH 7+ |  |  |  | 4.22 | 1.8573 | 0.139 | 0.18573 | 0.018573
[5 empty] | V - | 32 | BR | N2 - 8 | 6-9 | 0.21 | 32.24 | 32.03 | 61 | BR | A3 - C |  | 13-2 | 2 | 3 / 21 |  |  | 6 - 2 | RTN / 9 | 77.35 | 6.701 | 0.234 | 0.6701 | 0.06701
[5 empty] | I - | 32 | Y | N2 - 9 | 7-8 | 0.15 | 32.26 | 32.11 | 61 | Y | A3 - D |  |  |  |  | CH 7 - |  |  |  | 273.2 | 27.005 |  | 2.7005 | 0.27005
14-1
|  | 27 | 1573 | TS13 | V + | 33 | BR | N2 - 14 | 14-16 | 0.18 | 32.54 | 32.36 | 62 | BR | A3 - H |  | 15-1 | 3 | 3 / 23 |  |  | 7 - 1 | 3:I.ch7+ / 18 | m = 10 b = -2 |  |  |  |  | -0.21654 | 16.9406 | -26.4275
[5 empty] | I + | 33 | BL | N2 - 15 | 15-17 | 0.16 | 32.54 | 32.38 | 62 | BL | A3 - J |  | 14-2 |  |  | CH 8+ |  |  |  | 4.22 | 1.853 | 0.106 | 0.1853 | 0.01853
[5 empty] | V - | 33 | BR | N2 - 16 | 14-17 | 0.11 | 32.55 | 32.44 | 62 | BR | A3 - K |  | 15-2 | 2 | 3 / 24 |  |  | 7 - 2 | RTN / 9 | 77.35 | 6.6997 | 0.221 | 0.66997 | 0.066997
[5 empty] | I - | 33 | BL | N2 - 17 | 15-16 | 0.09 | 32.55 | 32.46 | 62 | BL | A3 - L |  |  |  |  | CH 8 - |  |  |  | 273.2 | 27.017 |  | 2.7017 | 0.27017
J4
supply |  | 28 | 1574 | TS9 | V + | [7 empty] | 58 | BR | A1 - H |  | 16-1 | 3 | 4 / 1 |  |  | 8 - 1 | 3:I.ch8+ / 20 | m = 10 b = 2 |  |  |  |  | -0.21756 | 16.9545 | -25.9635
[5 empty] | I + | [7 empty] | 58 | GR | A1 - J |  | 17-1 |  |  | CH 1+ |  |  |  | 4.22 | 1.8229 | -0.1281 | 0.18229 | 0.018229
[5 empty] | V - | [7 empty] | 58 | BR | A1 - K |  | 16-2 | 2 | 4 / 3 |  |  | 8 - 2 | RTN / 9 | 77.35 | 6.6633 | -0.154 | 0.66633 | 0.066633
[5 empty] | I - | [7 empty] | 58 | GR | A1 - L |  | 17-2 |  |  | CH 1 - |  |  |  | 273.2 | 26.999 |  | 2.6999 | 0.26999
Return |  | 29 | 1578 | TS10 | V + | [7 empty] | 59 | BR | A2 - A |  | 18-1 | 3 | 4 / 5 |  |  | 9 -1 | 3:I.ch9+ / 22 | m = 10 b = 2 |  |  |  |  | -0.21841 | 16.9779 | -26.0017
[5 empty] | I + | [7 empty] | 59 | W | A2 - B |  | 19-1 |  |  | CH 2+ |  |  |  | 4.22 | 1.8228 | -0.128 | 0.18228 | 0.018228
[5 empty] | V - | [7 empty] | 59 | BR | A2 - C |  | 18-2 | 2 | 4 / 6 |  |  | 9 - 2 | RTN / 9 | 77.35 | 6.6576 | -0.213 | 0.66576 | 0.066576
[5 empty] | I - | [7 empty] | 59 | W | A2 - D |  | 19-2 |  |  | CH 2- |  |  |  | 273.2 | 27.003 |  | 2.7003 | 0.27003
midpoint |  | 30 | 1628 | TS11 | V + | 31 | BR | N2 - 1 | 1-3 | 0.24 | 32.03 | 31.79 | 60 | BR | A2 - H |  | 20-1 | 3 | 4 / 7 |  |  | 10 -1 | 3:I.ch10+ / 24 | [5 empty] | -0.21822 | 16.9773 | -26.1549
[5 empty] | I + | 31 | R | N2 - 2 | 2-4 | 0.2 | 32.1 | 31.9 | 60 | R | A2 - J |  | 21-1 |  |  | CH 3+ |  |  |  | 4.22 | 1.8323 | -0.055 | 0.18323 | 0.018323
[5 empty] | V - | 31 | BR | N2 - 3 | 1-4 | 0.2 | 32 | 31.8 | 60 | BR | A2 - K |  | 20-2 | 2 | 4 / 9 |  |  | 10 - 2 | RTN / 9 | 77.35 | 6.6682 | 0.03 | 0.66682 | 0.066682
[5 empty] | I - | 31 | R | N2 - 4 | 2-3 | 0.22 | 32.12 | 31.9 | 60 | R | A2 - L |  | 21-2 |  |  | CH 3 - |  |  |  | 273.2 | 27.01 |  | 2.701 | 0.2701
Supply |  | 31 |  | AST1a | I+ | [7 empty] | 109 | BR | F2 - J |  | 22-1 |  | 4 / 11 |  |  | 11 - 1 | 3:I.ch11+ / 26
[5 empty] | V+ | [7 empty] | 109 | BR | F2 - K |  | 23-1 |  |  | CH 4+
[5 empty] | I- | [7 empty] | 109 | G | F2 - L |  | 23-2 |  | 4 / 12 |  |  | 11 - 2 | RTN / 9
[17 empty] | 23-2 |  |  | CH 4 -
[17 empty] | 24-1 |  | 4 / 13 |  |  | 12 - 1 | 3:I.ch12+ / 30
Return |  | 32 |  | AST2 | I+ | [7 empty] | 92 | BR | F1 - A |  | 25-1 |  |  | CH 5+
[5 empty] | V+ | [7 empty] | 92 | BR | F1 - B |  | 24-2 |  | 4 / 15 |  |  | 12 - 2 | RTN / 27
[5 empty] | I- | [7 empty] | 92 | G | F1 - C |  | 25-2 |  |  | CH 5 -
| DIODE 2 WIRE | 19 |  | CU1 | I+ | [7 empty] | 94 | BR | F2 - A |  | 26-1 |  | 4 / 17 |  |  | 13 - 1 | 3:I.ch13+ / 32
[5 empty] | I- | [7 empty] | 94 | O | F2 - B |  | 27-1 |  |  | CH 6+
[17 empty] | 26-2 |  | 4 / 18 |  |  | 13 - 2 | RTN / 27
[17 empty] | 27-2 |  |  | CH 6 -
| DIODE 2 WIRE | 20 |  | CU2 | I+ | [7 empty] | 95 | BR | F2- E |  | 28-1 |  | 4 / 19 |  |  | 14 - 1
[5 empty] | I- | [7 empty] | 95 | Y | F2- F |  | 29-1 |  |  | CH 7+ |  |  | 3:I.ch14+ / 34
[17 empty] | 28-2 |  | 4 / 21 |  |  | 14 - 2
[17 empty] | 29-2 |  |  | CH 7 - |  |  | RTN / 27
| 3 WIRE | 21 |  | AST4 | I+ | [7 empty] | 93 | O | F1 - H |  | 30-1 |  | 4 / 23 |  |  | 15 - 1
[5 empty] | I- | [7 empty] | 93 | V | F1 - J |  | 31-1 |  |  | CH 8+ |  |  | 3:I.ch15+ / 36
[5 empty] | V+ | [7 empty] | 93 | O | F1 - K |  | 30-2 |  | 4 / 24 |  |  | 15 - 2
[17 empty] | 31-2 |  |  | CH 8 - |  |  | RTN / 27
ON Helium Return pipe

## Strain Gauges
Signal description | Range | Description |  | connector pin | Wire color | Terminal Strip | Dataforth SCM5B38 - 031D | Dataforth feed out | Wire Color | Terminal Strip | 1756 IF16
Strain Gauges |  | Strain Guages/ Load cells | AWG-24 4-wire shield cable | Tuchel 5 pin socket |  |  | +/-3mV/V +3.33V excitation |  |  |  | slot #/ Channel # / pin #
Radial Gauges | 0 to 5 kgf | Full bridged 350Ohms 1mV/V @ 11,000 lbf |  |  |  |  | pin # | Board # / Pin # |  | TS-PLC-07
A |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 5 / 1 |  | 0-1 | 7:I.ch0+ / 2
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 5 / 3 |  | 0-2 | RTN / 9
|  |  | Supply + | 5 |  |  | 1
B |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 5 / 5 |  | 1-1 | 7:I.ch1+ / 4
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 5 / 6 |  | 1-2 | RTN / 9
|  |  | Supply + | 5 |  |  | 1
C |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 5 / 7 |  | 2-1 | 7:I.ch2+ / 6
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 5 / 9 |  | 2-2 | RTN / 9
|  |  | Supply + | 5 |  |  | 1
D |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 5 / 11 |  | 3-1 | 7:I.ch3+ / 8
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 5 / 12 |  | 3-2 | RTN / 9
|  |  | Supply + | 5 |  |  | 1
E |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 5 / 13 |  | 4-1 | 7:I.ch4+ / 12
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 5 / 15 |  | 4-2 | RTN / 9
|  |  | Supply + | 5 |  |  | 1
F |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 5 / 17 |  | 5-1 | 7:I.ch5+ / 14
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 5 / 18 |  | 5-2 | RTN / 9
|  |  | Supply + | 5 |  |  | 1
G |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 5 / 19 |  | 6-1 | 7:I.ch6+ / 16
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 5 / 21 |  | 6-2 | RTN / 9
|  |  | Supply + | 5 |  |  | 1
H |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 5 / 23 |  | 7-1 | 7:I.ch7+ / 18
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 5 / 24 |  | 7-2 | RTN / 9
|  |  | Supply + | 5 |  |  | 1
J |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 6 / 1 |  | 8-1 | 7:I.ch8+ / 20
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 6 / 3 |  | 8-2 | RTN / 9
|  |  | Supply + | 5 |  |  | 1
K |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 6 / 5 |  | 9-1 | 7:I.ch9+ / 22
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 6 / 6 |  | 9-2 | RTN / 27
|  |  | Supply + | 5 |  |  | 1
L |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 6 / 7 |  | 10-1 | 7:I.ch10+ / 24
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 6 / 9 |  | 10-2 | RTN / 27
|  |  | Supply + | 5 |  |  | 1
M |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 6 / 11 |  | 11-1 | 7:I.ch11+ / 26
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 6 / 12 |  | 11-2 | RTN / 27
|  |  | Supply + | 5 |  |  | 1
N |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 6 / 13 |  | 12-1 | 7:I.ch12+ / 30
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 6 / 15 |  | 12-2 | RTN / 27
|  |  | Supply + | 5 |  |  | 1
P |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 6 / 17 |  | 13-1 | 7:I.ch13+ / 32
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 6 / 18 |  | 13-2 | RTN / 27
|  |  | Supply + | 5 |  |  | 1
R |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 6 / 19 |  | 14-1 | 7:I.ch14+ / 34
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 6 / 21 |  | 14-2 | RTN / 27
|  |  | Supply + | 5 |  |  | 1
S |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 6 / 23 |  | 15-1 | 7:I.ch15+ / 36
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 6 / 24 |  | 15-2 | RTN / 27
|  |  | Supply + | 5 |  |  | 1
Axial Gauges | 0 to 5 kgf | Strainsert FL10U(C)-2SG(W)B | AWG-24 4-wire shield cable | Tuchel 5 pin socket |  |  | pin # | Board # / Pin # |  | TS-PLC-08 | slot 8 / pin #
T |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 7 / 1 |  | 0-1 | 8:I.ch0+ / 2
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 7 / 3 |  | 0-2 | RTN / 9
|  |  | Supply + | 5 |  |  | 1
V |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 7 / 5 |  | 1-1 | 8:I.ch1+ / 4
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 7 / 6 |  | 1-2 | RTN / 9
|  |  | Supply + | 5 |  |  | 1
W |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 7 / 7 |  | 2-1 | 8:I.ch2+ / 6
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 7 / 9 |  | 2-2 | RTN / 9
|  |  | Supply + | 5 |  |  | 1
X |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 7 / 11 |  | 3-1 | 8:I.ch3+ / 8
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 7 / 12 |  | 3-2 | RTN / 9
|  |  | Supply + | 5 |  |  | 1

## Voltage taps
Signal description | VT # | Internal Wire # color | Internal connector / pin | Measured romm temperture resistnce |  |  |  | Internal Cable # Color | Wire color | Vacuum Feed Thur Connector / Pin | Protection Resistor | Wire Color | Terminal Strip | Wire Color | Dataforth | Dataforth feed out | Wire Color | Terminal Strip | 1756 IF16
Voltage Taps |  |  | 25 pin D connectors | Fluke 87 III S# 82800752 with remote term blk db 25 pin patch panel (102FA) | [5 empty] | 10 pin vacuum feedthrough |  |  |  |  | 31-09D | Board # / pin # |  |  | slot # / Channel # / pin #
|  |  |  | Pin to Pin | Residual terminal block resistance |  | Ohms | [5 empty] | TS-09 |  |  |  |  | TS-PLC-09
|  |  |  | Pin to Pin | Voltmeter lead resistance null out | Total resistance | Voltage Tap resistance |  | R | Rho | L | A
|  |  |  | Lead 1 to lead 2 |  | 4.12 | Multimeter |  | 4 | 2.1875e-08 | 12500 | 6.836e-05
Section 7.4 Page 85 |  |  |  | Lead 1 to lead 2 |  | 3.99 | Meager #1 |  | 4 | 2.7e-08 | 10127.4074 | 6.836e-05
|  |  |  | Lead 1 to Cryostat | 280V | >1Tohm | Meager #2 |  | 4.0001 | 2.7e-08 | 12500 | 8.4372e-05
at Top of current lead A Warm end | VT 6 |  |  | VT6 - 7 | 1.35 | 3.88 | 2.53 |  |  |  | 200k Resistors |  | 1-1 |  | 3 | 9 / 1 |  | 1-1 | 9:I.ch0+ / 2
at Bottom of current lead A | VT 4 |  |  | VT4 - 7 | 1.35 | 3.89 | 2.54 | B86 | R | B2 - L | 200k Resistors |  | 2-1 |  | 2 | 9 / 3 |  | 2-1 | 9:I.ch0- / 2
|  |  |  | VT6 - VT7 | 1.37 |  | -1.37
at Top of current lead B Warm end | VT 7 |  |  | VT7 - 7 | 1.35 |  | -1.35 |  |  |  | 200k Resistors |  | 3-1 |  | 3 | 9 / 5 |  | 3-1 | 9:I.ch1+ / 2
at Bottom of current lead B | VT 5 |  |  | VT5 - 7 | 1.35 |  | -1.35 | B91 | B | B3 - L | 200k Resistors |  | 4-1 |  | 2 | 9 / 6 |  | 4-1 | 9:I.ch1- / 2
Coil exit point Inner layer | VT 2 | 5 G | H4 - 25 | 7-25 | 2.41 | 5.17 | 2.76 | 83 G | BL | B2 - C | 200k Resistors |  | 5-1 |  | 3 | 9 / 7 |  | 5-1 | 9:I.ch2+ / 2
Center tap | VT 1 | 3 O | H4 - 7 |  |  |  |  | 79 O | O | B1 - E | 200k Resistors |  | 6-1 |  | 2 | 9 / 9 |  | 6-1 | 9:I.ch2- / 2
Center tap | VT 1 | 3 O | H4 - 7 | 7-1 | 2.42 | 5.02 | 2.6 | 79 O | O | B1 - E |  |  | 6-1 |  | 3 | 9 / 11 |  | 6-1 | 9:I.ch3+ / 2
Coil exit point outer layer | VT 3 | 1 BR | H4 - 1 |  |  |  |  | 77 B | B | B1 - A | 200k Resistors |  | 7-1 |  | 2 | 9 / 12 |  | 7-1 | 9:I.ch3- / 2
Flux Loop 1 inner coil | FL 1 | 10 B | H4 - 5 | 5-9 | 2.93 | 25.02 | 22.09 | 80 B | B | B1 - J | 200k Resistors |  | 8-1 |  | 3 | 9 / 13 |  | 8-1 | 9:I.ch4+ / 2
|  | 10 BR | H4 - 9 |  |  |  |  | 81 BR | BR | B1 - K | 200k Resistors |  | 9-1 |  | 2 | 9 / 15 |  | 9-1 | 9:I.ch4- / 2
Flux loop 2 outer coil | FL 2 | 11 O | H4 - 16 | 16-23 | 2.35 | 24.63 | 22.28 | 85 O | O | B2 - H | 200k Resistors |  | 10-1 |  | 3 | 9 / 17 |  | 10-1 | 9:I.ch5+ / 2
|  | 11 BR | H4 - 23 |  |  |  |  | 90 BR | BR | B3 - H | 200k Resistors |  | 11-1 |  | 2 | 9 / 18 |  | 11-1 | 9:I.ch5- / 2
Quench Detector
at Top of current lead A Warm end | VT 6 | [11 empty] | 1-2 |  | D1H
at Bottom of current lead A | VT 4 | [6 empty] | B86 | R | B2 - L |  |  | 2-2 |  | D1HM
at Top of current lead B Warm end | VT 7 | [11 empty] | 3-2 |  | D4H
at Bottom of current lead B | VT 5 | [6 empty] | B91 | B | B3 - L |  |  | 4-2 |  | D4HM
Coil exit point Inner layer | VT 2 | 5 G | H4 - 25 |  |  |  |  | 83 G | BL | B2 - C |  |  | 5-2 |  | D2H
Center tap | VT 1 | 3 O | H4 - 7 |  |  |  |  | 79 O | O | B1 - E |  |  | 6-2 |  | D2HM
Center tap | VT 1 | 3 O | H4 - 7 |  |  |  |  | 79 O | O | B1 - E |  |  | 6-2 |  | D2LM
Coil exit point outer layer | VT 3 | 1 BR | H4 - 1 |  |  |  |  | 77 B | B | B1 - A |  |  | 7-2 |  | D2L
Flux Loop 1 inner coil | FL 1 | 10 B | H4 - 5 |  |  |  |  | 80 B | B | B1 - J |  |  | 8-2 |  | D3H
20 turns |  | 10 BR | H4 - 9 |  |  |  |  | 81 BR | BR | B1 - K |  |  | 9-2 |  | D3HM
Flux loop 2 outer coil | FL 2 | 11 O | H4 - 16 |  |  |  |  | 85 O | O | B2 - H |  |  | 10-2 |  | D3LM
20 turns |  | 11 BR | H4 - 23 |  |  |  |  | 90 BR | BR | B3 - H |  |  | 11-2 |  | D3L
Heater Mid-plane
6 turn
200Vdc @ 2A 400 Watts Continous
200Vdc @ 4A =800 Watts (10 secs) |  |  |  | room temperature Coil resistance
Al

## other sensors
Signal description | Range |  | CCL sensor # | Description |  | Internal Cable # | Wire color | Internal connector / pin | Internal Cable # | Vacuum Feed Thur Connector / Pin | Wire Color
Pressure |  |  |  | Pressure gauges
LN2 Pressure | 0.4 - 5 Bar | 1 | 0 | P6P/5bar Hottinger Baldwin Gmbh 5V/225Hz strain gauge full bridge |  |  |  |  | Dataforth SCM5B38 - 35D | +/- 10 VDC | 1756 - IF16
LHe Pressure | 0.4 - 5 Bar | 1 | 0 | P6P/5bar Hottinger Baldwin Gmbh 5V/225Hz strain gauge full bridge |  |  |  |  | Dataforth SCM5B38 - 35D | +/- 10 VDC | 1756 - IF16
Vacuum |  |  |  | Vacuum gauges
LH Thermovac TR206 | Vacuum gauge | 1 | 0 | 760 Torr to 1x10-3 Torr |  |  |  |  | Leybold Combivac CM 31 | 0 - 10 VDC | 1756 - IF16
Liquid Level |  |  |  | Level Probes
LN2 Level | 0 - 100 % | 1 | 0 | LN2 - Capitance 3 wire |  |  |  |  | None | 4-20 mA
LHe Level S# 616-412-25-006 | 0 - 100 % | 1 | 0 | 740mm active length, 1340mm overall length, 0 to 0.5V (100% to 0%) | 10 pin 4 used |  |  |  | Dataforth SCM5B31 - 03D | +/- 10 VDC | 1756 - IF16
Gas Flow Meters |  |  |  | He Mass Flow Meters
Lead He Gas Flow | 0-2000 sccm N2 | 2 | 0 | MKS 1749A-02000SV |  |  |  |  | None | 4-20 mA | 1756 - IF16
JT Valves |  |  |  | Valves
JT Valve Manual/Auto Switch |  | 1 |  | Switch for all 11 JT valves | [6 empty] | 1756 - IB16D
Jt Valve Motors |  | 10 |  | Globe Motor (24 VDC @ 0.8A) |  |  |  |  | Dataforth DSCA43-15C | 24V dc out to motor. | 1756 - OW16
LVDT | -6-100% | 10 |  | LVDT |  |  |  |  | Macro Sensors LVC-2402 | 4-20 mA | 1756 - IF16
Quench detector |  | 1 | 1 | 4 channel unit
|  | 9 |  | Interlock Signals |  |  |  |  | None | Digital Signal | 1756 - IB16D
|  | 1 |  | Relay Contact |  |  |  |  | None
|  | 1 |  | Reset Signal |  |  |  |  | None |  | 1756 - OW16
Heaters/Thermostat external |  |  |  | Heaters/Thermostat external
|  | 5 | 0
NMR Unit |  |  |  | NMR
PLC commuications | Ascii | 1 |  | Metrolab Field Regulation RG2040 |  |  |  |  | Flex I/O | RS232 | Control Net
Power on/off relay |  | 1 | [6 empty] | None |  | 1756 - OW16
Power Supply |  |  |  | PSU
PSU Interlocks and Status Signals |  | 16 |  | Interlock and Status Signals from PSU |  |  |  |  | None | Digital Signal | 1756 - IB16D
PSU Interlocks |  | 2 |  | Interlock Signals to the PSU from PLC |  |  |  |  | None |  | 1756 - OW16
PSU Reset Signal |  | 1 |  | Reset Signal to PSU & Diodes |  |  |  |  | None |  | 1756 - OW16
PSU Course Current Signal |  | 1 |  | 3 V dc +/- Dac output from PSU |  |  |  |  | Dataforth SCM5B31 - 03D | +/- 10 VDC | 1756 - IF16

## Signal Conditioners
Date Oct. 26, 2018
|  | DataForth Signal Conditioners |  | 1
Qty | Catalog # | Description | Price | Total Cost | Date | PR/PO
30 | SCM5B-30-01D | Analog Input Narrow bandwidth 10mV in / 10V out | 265 | 7950
1 | SCM5B-31-03D | Analog Voltage Input Module 10V in / 10V out | 160 | 160
8 | SCM5B-31-09D | Analog Voltage Input Module 40V in / 10V out | 160 | 1280
20 | SCM5B-38-31D | Strain Gage Input Module | 162 | 3240
4 | SCM5B-43-05D | General Purpose with excitation 5V in / 10V out | 205 | 820
10 | SCMXCA-004 12 | Ribbon Cable | 35 | 350
10 | SCMPB-05-02 | Backplane-8 channels with DIN mounts | 227 | 2270
10 | SCMXIF-DIN | Breakout Board Din Mounted | 63 | 630
10 | SCMXRK-002 | Rack Mounting Plate 19" | 61 | 610
7 | LVC-2500 | LVDT signal conditioner 24Vdc | 425 | 2975
4 | CMA1D03A | Constant current Source boards 8channels | 410 | 1640
1 | SDP5-5-100T | PSU 5 volts | 92.4 | 92.4
2 | SDN10-24-100P | PSU 24V Psu | 292.3 | 584.6
|  |  | SUM | 22602

## PRs
SHMS Electrical Components |  | 12CCMC | 12GEV | 43-014 & 43-053
PR | Sent Date | PR Estimate | PO /Buyer | Vendor | PO Cost | Description | Delivery Complete | Location
378971 | 2018-08-08 | 55898 | 18-P1254 | EECO | 55755.39 | PLC and I/O modules | 2018-09-26 | TED 1544
378847 | 2018-07-31 | 33030 | 18-01188 | EECO | 33030 | Software | 2018-08-24 | TED 1544
382342 | 2019-01-15 | 172 | 19-M0362 | EECO | 172 | Cable
|  | 89100 |  |  | 88957.39 | Total