## other sensors
Signal description | Range |  | CCL sensor # | Description |  | Internal Cable # | Wire color | Internal connector / pin | Internal Cable # | Vacuum Feed Thur Connector / Pin | Wire Color
Pressure |  |  |  | Pressure gauges
LN2 Pressure | 0.4 - 5 Bar | 1 | 0 | P6P/5bar Hottinger Baldwin Gmbh 5V/225Hz strain gauge full bridge |  |  |  |  | Dataforth SCM5B38 - 35D | +/- 10 VDC | 1756 - IF16
LHe Pressure | 0.4 - 5 Bar | 1 | 0 | P6P/5bar Hottinger Baldwin Gmbh 5V/225Hz strain gauge full bridge |  |  |  |  | Dataforth SCM5B38 - 35D | +/- 10 VDC | 1756 - IF16
Vacuum |  |  |  | Vacuum gauges
LH Thermovac TR206 | Vacuum gauge | 1 | 0 | 760 Torr to 1x10-3 Torr | [5 empty] | 0 - 10 VDC | 1756 - IF16
Liquid Level |  |  |  | Level Probes
LN2 Level | 0 - 100 % | 1 | 0 | LN2 - Capitance 3 wire |  |  |  |  | None | 4-20 mA
LHe Level S# 616-412-25-006 | 0 - 100 % | 1 | 0 | 740mm active length, 1340mm overall length, 0 to 0.5V (100% to 0%) | 10 pin 4 used |  |  |  | Dataforth SCM5B31 - 03D | +/- 10 VDC | 1756 - IF16
Gas Flow Meters |  |  |  | He Mass Flow Meters
Lead He Gas Flow | 0-2000 sccm N2 | 2 | 0 | MKS 1749A-02000SV |  |  |  |  | None | 4-20 mA | 1756 - IF16
JT Valves |  |  |  | Valves
JT Valve Manual/Auto Switch |  | 1 |  | Switch for all 11 JT valves | [6 empty] | 1756 - IB16D
Jt Valve Motors |  | 10 |  | Globe Motor (24 VDC @ 0.8A) |  |  |  |  | Dataforth DSCA43-15C | 24V dc out to motor. | 1756 - OW16
LVDT | -6-100% | 10 |  | LVDT |  |  |  |  | Macro Sensors LVC-2402 | 4-20 mA | 1756 - IF16
Quench detector |  | 1 | 1 | 4 channel unit
|  | 9 |  | Interlock Signals |  |  |  |  | None | Digital Signal | 1756 - IB16D
|  | 1 |  | Relay Contact |  |  |  |  | None
|  | 1 |  | Reset Signal |  |  |  |  | None |  | 1756 - OW16
Heaters/Thermostat external |  |  |  | Heaters/Thermostat external
|  | 5 | 0
NMR Unit |  |  |  | NMR
PLC commuications | Ascii | 1 |  | Metrolab Field Regulation RG2040 |  |  |  |  | Flex I/O | RS232 | Control Net
Power on/off relay |  | 1 | [6 empty] | None |  | 1756 - OW16
Power Supply |  |  |  | PSU
PSU Interlocks and Status Signals |  | 16 |  | Interlock and Status Signals from PSU |  |  |  |  | None | Digital Signal | 1756 - IB16D
PSU Interlocks |  | 2 |  | Interlock Signals to the PSU from PLC |  |  |  |  | None |  | 1756 - OW16
PSU Reset Signal |  | 1 |  | Reset Signal to PSU & Diodes |  |  |  |  | None |  | 1756 - OW16
PSU Course Current Signal |  | 1 |  | 3 V dc +/- Dac output from PSU |  |  |  |  | Dataforth SCM5B31 - 03D | +/- 10 VDC | 1756 - IF16

## Strain Gauges
Signal description | Range | Description |  | connector pin | Wire color | Terminal Strip | Dataforth SCM5B38 - 031D | Dataforth feed out | Wire Color | Terminal Strip | 1756 IF16
Strain Gauges |  | Strain Guages/ Load cells | AWG-24 4-wire shield cable | Tuchel 5 pin socket |  |  | +/-3mV/V +3.33V excitation |  |  |  | slot #/ Channel # / pin #
Radial Gauges | 0 to 5 kgf | Full bridged 350Ohms 1mV/V @ 11,000 lbf |  | sg-x |  | MAG-TS-07 | pin # | Board # / Pin # |  | TS-PLC-07
A |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 5 / 1 |  | 0-1 | 7:I.ch0+ / 2
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 5 / 3 |  | 0-2 | RTN / 9
|  |  | Supply + | 5 |  |  | 1
B |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 5 / 5 |  | 1-1 | 7:I.ch1+ / 4
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 5 / 6 |  | 1-2 | RTN / 9
|  |  | Supply + | 5 |  |  | 1
C |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 5 / 7 |  | 2-1 | 7:I.ch2+ / 6
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 5 / 9 |  | 2-2 | RTN / 9
|  |  | Supply + | 5 |  |  | 1
D |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 5 / 11 |  | 3-1 | 7:I.ch3+ / 8
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 5 / 12 |  | 3-2 | RTN / 9
|  |  | Supply + | 5 |  |  | 1
E |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 5 / 13 |  | 4-1 | 7:I.ch4+ / 12
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 5 / 15 |  | 4-2 | RTN / 9
|  |  | Supply + | 5 |  |  | 1
F |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 5 / 17 |  | 5-1 | 7:I.ch5+ / 14
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 5 / 18 |  | 5-2 | RTN / 9
|  |  | Supply + | 5 |  |  | 1
G |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 5 / 19 |  | 6-1 | 7:I.ch6+ / 16
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 5 / 21 |  | 6-2 | RTN / 9
|  |  | Supply + | 5 |  |  | 1
H |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 5 / 23 |  | 7-1 | 7:I.ch7+ / 18
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 5 / 24 |  | 7-2 | RTN / 9
|  |  | Supply + | 5 |  |  | 1
J |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 6 / 1 |  | 8-1 | 7:I.ch8+ / 20
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 6 / 3 |  | 8-2 | RTN / 9
|  |  | Supply + | 5 |  |  | 1
K |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 6 / 5 |  | 9-1 | 7:I.ch9+ / 22
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 6 / 6 |  | 9-2 | RTN / 27
|  |  | Supply + | 5 |  |  | 1
L |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 6 / 7 |  | 10-1 | 7:I.ch10+ / 24
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 6 / 9 |  | 10-2 | RTN / 27
|  |  | Supply + | 5 |  |  | 1
M |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 6 / 11 |  | 11-1 | 7:I.ch11+ / 26
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 6 / 12 |  | 11-2 | RTN / 27
|  |  | Supply + | 5 |  |  | 1
N |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 6 / 13 |  | 12-1 | 7:I.ch12+ / 30
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 6 / 15 |  | 12-2 | RTN / 27
|  |  | Supply + | 5 |  |  | 1
P |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 6 / 17 |  | 13-1 | 7:I.ch13+ / 32
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 6 / 18 |  | 13-2 | RTN / 27
|  |  | Supply + | 5 |  |  | 1
R |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 6 / 19 |  | 14-1 | 7:I.ch14+ / 34
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 6 / 21 |  | 14-2 | RTN / 27
|  |  | Supply + | 5 |  |  | 1
S |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 6 / 23 |  | 15-1 | 7:I.ch15+ / 36
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 6 / 24 |  | 15-2 | RTN / 27
|  |  | Supply + | 5 |  |  | 1
Axial Gauges | 0 to 5 kgf | Strainsert FL10U(C)-2SG(W)B | AWG-24 4-wire shield cable | Tuchel 5 pin socket |  | MAG-TS-08 | pin # | Board # / Pin # |  | TS-PLC-08 | slot 8 / pin #
T |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 7 / 1 |  | 0-1 | 8:I.ch0+ / 2
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 7 / 3 |  | 0-2 | RTN / 9
|  |  | Supply + | 5 |  |  | 1
V |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 7 / 5 |  | 1-1 | 8:I.ch1+ / 4
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 7 / 6 |  | 1-2 | RTN / 9
|  |  | Supply + | 5 |  |  | 1
W |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 7 / 7 |  | 2-1 | 8:I.ch2+ / 6
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 7 / 9 |  | 2-2 | RTN / 9
|  |  | Supply + | 5 |  |  | 1
X |  |  | Supply - | 1 |  |  | 4
|  |  | Signal P | 2 |  |  | 3 | 7 / 11 |  | 3-1 | 8:I.ch3+ / 8
|  |  | Screen GND | 3
|  |  | Signal N | 4 |  |  | 2 | 7 / 12 |  | 3-2 | RTN / 9
|  |  | Supply + | 5 |  |  | 1

## Temp Sensors
Signal description | Range | Digi-Trend sensor Number | CCL sensor serial # | Description |  | Internal Cable # | Wire color | Internal connector / pin | Measured room temperature resistance |  |  |  | Internal Cable # | Wire Color | Vacuum Feed Thru Connector / Pin | Cable Label | Wire color | Terminal Strip | Dataforth SCM5B30-01D | Dataforth feed out | Cable Label | Wire color | Constant Current Source | Terminal Strip | Cable Label | Wire Color | 1756 IF16 | Temperature | Resistance | Δ T [K] | Voltage from a 100uA | Power | Coefficients |  |  | Changed name
[8 empty] | 25 pin D connectors | Fluke 87 III S# 82800752 with remote term blk db 25 pin patch panel (102FA) | [5 empty] | 10 pin vacuum feedthrough |  |  |  | +/- 10mV in +/- 10V out |  |  |  | 100uA | [7 empty] | mV | uW | Rodium Fe sensors
Temperature |  |  |  | Temperature sensors |  |  |  | 5 | Pin to Pin | Residual terminal block resistance |  | Ohms | [6 empty] | pin # | Board # / Pin # |  |  | 8 Channels per card |  |  |  |  | K | R |  | 0.0001 |  | Table 7.1 Page 75
Temperature Sensor He Coil Shell | 300 - 4 K | [7 empty] | Pin to Pin | Voltmeter lead resistance null out | Total resistance | Probe resistance | [5 empty] | MAG-TS-02 |  |  |  |  | J1 | TS-PLC-02 |  |  | slot #/ Channel # / pin # | [5 empty] | x^2 | x | b
Rhodium/Fe 4-wire PH20002
Coil Outer shell |  | 1 | 1552 | TS1 | V + | 13 | B | H1 - 1 | 1-3 | 2.34 | 32.54 | 30.2 | 50 | B | D1 - A | 400-01 | BK | 0-1 | 3 | 1 / 1 | 210-04 | BL/WT |  | 0 - 1 | 302-01 | BL/WT | 2:I.ch0+ / 2 | m = 10 b = 1 |  |  |  |  | -0.21783 | 16.9703 | -26.2339
[5 empty] | I + | 13 | BR | H1 - 2 | 2-4 | 2.43 | 32.59 | 30.16 | 50 | BR | D1 - B | 400-01 | RD | 1-1 |  |  |  |  | CH 1+ |  |  |  |  | 4.22 | 1.8379 | -0.0115 | 0.18379 | 0.018379
Train gauge |  |  |  |  | V - | 13 | B | H1 - 3 | 1-4 | 2.35 | 32.6 | 30.25 | 50 | B | D1 - C | 400-01 | WT | 0-2 | 2 | 1 / 3 | 210-04 | WT/BL |  | 0 - 2 |  | RTN | RTN / 9 | 77.35 | 6.6759 | -0.024 | 0.66759 | 0.066759
[5 empty] | I - | 13 | BR | H1 - 4 | 2-3 | 2.47 | 32.62 | 30.15 | 50 | BR | D1 - D | 400-01 | GN | 1-2 |  |  |  |  | CH 1 - |  |  |  |  | 273.2 | 27.006 |  | 2.7006 | 0.27006
Coil Outer shell |  | 2 | 1553 | TS2 | V + | 14 | B | H1 - 6 | 6-8 | 2.37 | 32.16 | 29.79 | 51 | B | D1 - F | 400-01 | BL | 2-1 | 3 | 1 / 5 | 210-04 | OR/WT |  | 1 - 1 | 302-01 | WT/BL | 2:I.ch1+ / 4 | m = 10 b = 1 |  |  |  |  | -0.21713 | 16.946 | -26.1027
[5 empty] | I + | 14 | R | H1 - 7 | 7-9 | 2.35 | 32.1 | 29.75 | 51 | R | D1 - H | 400-01 | OR | 3-1 |  |  |  |  | CH 2+ |  |  |  |  | 4.22 | 1.8324 | -0.054 | 0.18324 | 0.018324
[5 empty] | V - | 14 | B | H1 - 8 | 6-9 | 2.37 | 32.11 | 29.74 | 51 | B | D1 - J | 400-01 | YL | 2-2 | 2 | 1 / 6 | 210-04 | WT/OR |  | 1 - 2 |  | RTN | RTN / 9 | 77.35 | 6.6759 | -0.025 | 0.66759 | 0.066759
[5 empty] | I - | 14 | R | H1 - 9 | 7-8 | 2.38 | 32.14 | 29.76 | 51 | R | D1 - K | 400-01 | VT | 3-2 |  |  |  |  | CH 2 - |  |  |  |  | 273.2 | 27.009 |  | 2.7009 | 0.27009
Coil Outer shell |  | 3 | 1554 | TS3 | V + | 15 | B | H1 - 14 | 14-16 | 2.39 | 32.01 | 29.62 | 52 | B | D1 - L | 400-01 | SL | 4-1 | 3 | 1 / 7 | 210-04 | GN/WT |  | 2 - 1 | 302-01 | OR/WT | 2:I.ch2+ / 6 | m = 10 b = 1 |  |  |  |  | -0.21673 | 16.9387 | -26.2603
[5 empty] | I + | 15 | O | H1 - 15 | 15-17 | 2.34 | 31.96 | 29.62 | 52 | O | D2 - A | 400-02 | BK | 5-1 |  |  |  |  | CH 3+ |  |  |  |  | 4.22 | 1.8429 | 0.027 | 0.18429 | 0.018429
[5 empty] | V - | 15 | B | H1 - 16 | 14-17 | 2.39 | 32.01 | 29.62 | 52 | B | D2 - B | 400-02 | RD | 4-2 | 2 | 1 / 9 | 210-04 | WT/GN |  | 2 - 2 |  | RTN | RTN / 9 | 77.35 | 6.6893 | 0.114 | 0.66893 | 0.066893
[5 empty] | I - | 15 | O | H1 - 17 | 15-16 | 2.35 | 31.97 | 29.62 | 52 | O | D2 - C | 400-02 | WT | 5-2 |  |  |  |  | CH 3 - |  |  |  |  | 273.2 | 27.021 |  | 2.7021 | 0.27021
Coil Outer shell |  | 4 | 1556 | TS4 | V + | 16 | B | H1 - 19 | 19-21 | 2.36 | 32.36 | 30 | 53 | B | D2 - D | 400-02 | GN | 6-1 | 3 | 1 / 11 | 210-04 | BN/WT |  | 3 - 1 | 302-01 | WT/OR | 2:I.ch3+ / 8 | m = 10 b = 1 |  |  |  |  | -0.21719 | 16.9442 | -26.0267
[5 empty] | I + | 16 | Y | H1 - 20 | 20-22 | 2.38 | 32.33 | 29.95 | 53 | Y | D2 - E | 400-02 | BN | 7-1 |  |  |  |  | CH 4+ |  |  |  |  | 4.22 | 1.8279 | -0.0115 | 0.18279 | 0.018279
[5 empty] | V - | 16 | B | H1 - 21 | 19-22 | 2.35 | 32.37 | 30.02 | 53 | B | D2 - F | 400-02 | BL | 6-2 | 2 | 1 / 12 | 210-04 | WT/BN |  | 3 - 2 |  | RTN | RTN / 9 | 77.35 | 6.6715 | -0.07 | 0.66715 | 0.066715
[5 empty] | I - | 16 | Y | H1 -22 | 20-21 | 2.4 | 32.37 | 29.97 | 53 | Y | D2 - H | 400-02 | OR | 7-2 |  |  |  |  | CH 4 - |  |  |  |  | 273.2 | 27.012 |  | 2.7012 | 0.27012
Coil Outer shell |  | 5 | 1557 | TS5 | V + | 17 | B | H1 - 11 | 11-23 | 2.39 | 32.61 | 30.22 | 54 | B | D2 - K | 400-02 | VT | 8-1 | 3 | 1 / 13 | 210-04 | SL/WT |  | 4 - 1 | 302-01 | GN/WT | 2:I.ch4+ / 12 | m = 10 b = 1 |  |  |  |  | -0.21687 | 16.9366 | -25.9815
[5 empty] | I + | 17 | G | H1 - 12 | 12-24 | 2.38 | 32.54 | 30.16 | 54 | G | D2 - L | 400-02 | SL | 9-1 |  |  |  |  | CH 5+ |  |  |  |  | 4.22 | 1.8259 | -0.1048 | 0.18259 | 0.018259
[5 empty] | V - | 17 | B | H1 - 23 | 11-24 | 2.38 | 32.58 | 30.2 | 54 | B | D3 - A | 400-03 | BK | 8-2 | 2 | 1 / 15 | 210-04 | WT/SL |  | 4 - 2 |  |  | RTN / 9 | 77.35 | 6.6709 | -0.076 | 0.66709 | 0.066709
[5 empty] | I - | 17 | G | H1 -24 | 12-23 | 2.36 | 32.58 | 30.22 | 54 | G | D3 - B | 400-03 | RD | 9-2 |  |  |  |  | CH 5 - |  |  |  |  | 273.2 | 26.998 |  | 2.6998 | 0.26998
Coil Outer shell |  | 6 | 1558 | TS6 | V + | 18 | B | H2 - 1 | 1-3 | 2.34 | 32.27 | 29.93 | 55 | B | D3 - C | 400-03 | WT | 10-1 | 3 | 1 / 17 | 210-04 | BL/RD |  | 5 - 1 | 302-01 | WT/GN | 2:I.ch5+ / 14 | m = 10 b = 1 |  |  |  |  | -0.21755 | 16.9591 | -26.1271
[5 empty] | I + | 18 | BC | H2 - 2 | 2-4 | 2.43 | 32.33 | 29.9 | 55 | BC | D3 - D | 400-03 | GN | 11-1 |  |  |  |  | CH 6+ |  |  |  |  | 4.22 | 1.8325 | -0.053 | 0.18325 | 0.018325
[5 empty] | V - | 18 | B | H2 - 3 | 1-4 | 2.35 | 32.23 | 29.88 | 55 | B | D3 - E | 400-03 | BN | 10-2 | 2 | 1 / 18 | 210-04 | RD/BL |  | 5 - 2 |  | RTN | RTN / 9 | 77.35 | 6.6727 | -0.058 | 0.66727 | 0.066727
[5 empty] | I - | 18 | BC | H2 - 4 | 2-3 | 2.47 | 32.33 | 29.86 | 55 | BC | D3 - F | 400-03 | BL | 11-2 |  |  |  |  | CH 6 - |  |  |  |  | 273.2 | 27.004 |  | 2.7004 | 0.27004
Coil Outer shell |  | 7 | 1559 | TS21 | V + | 19 | R | H2 - 6 | 6-8 | 2.37 | 32.07 | 29.7 | 63 | R | C1 - A | 400-04 | BK | 12-1 | 3 | 1 / 19 | 210-04 | OR/RD |  | 6 - 1 | 302-01 | BN/WT | 2:I.ch6+ / 16 | m = 10 b = 1 |  |  |  |  | -0.21793 | 16.9696 | -26.0997
[5 empty] | I + | 19 | BR | H2 - 7 | 7-9 | 2.35 | 32.07 | 29.72 | 63 | BR | C1 - B | 400-04 | RD | 13-1 |  |  |  |  | CH 7+ |  |  |  |  | 4.22 | 1.8297 | -0.075 | 0.18297 | 0.018297
[5 empty] | V - | 19 | R | H2 - 8 | 6-9 | 2.37 | 32.06 | 29.69 | 63 | R | C1 - C | 400-04 | WT | 12-2 | 2 | 1 / 21 | 210-04 | RD/OR |  | 6 - 2 |  | RTN | RTN / 9 | 77.35 | 6.667 | -0.116 | 0.6667 | 0.06667
[5 empty] | I - | 19 | BR | H2 - 9 | 7-8 | 2.38 | 32.08 | 29.7 | 63 | BR | C1 - D | 400-04 | GN | 13-2 |  |  |  |  | CH 7 - |  |  |  |  | 273.2 | 26.998 |  | 2.6998 | 0.26998
Coil Outer shell |  | 8 | 1562 | TS22 | V + | 20 | R | H2 - 14 | 14-16 | 2.39 | 32.41 | 30.02 | 64 | R | C1 - F | 400-04 | BL | 14-1 | 3 | 1 / 23 | 210-04 | GN/RD |  | 7 - 1 | 302-01 | WT/BN | 2:I.ch7+ / 18 | m = 10 b = 1 |  |  |  |  | -0.2175 | 16.9583 | -26.0772
[5 empty] | I + | 20 | W | H2 - 15 | 15-17 | 2.34 | 32.41 | 30.07 | 64 | W | C1 - H | 400-04 | OR | 15-1 |  |  |  |  | CH 8+ |  |  |  |  | 4.22 | 1.8295 | -0.076 | 0.18295 | 0.018295
[5 empty] | V - | 20 | R | H2 - 16 | 14-17 | 2.39 | 32.42 | 30.03 | 64 | R | C1 - J | 400-04 | YL | 14-2 | 2 | 1 / 24 | 210-04 | RD/GN |  | 7 - 2 |  | RTN | RTN / 9 | 77.35 | 6.6694 | -0.091 | 0.66694 | 0.066694
[5 empty] | I - | 20 | W | H2 - 17 | 15-16 | 2.35 | 32.42 | 30.07 | 64 | W | C1 - K | 400-04 | VT | 15-2 |  |  |  |  | CH 8 - |  |  |  |  | 273.2 | 26.992 |  | 2.6992 | 0.26992
J2
Coil Outer shell |  | 9 | 1564 | TS23 | V + | 21 | R | H2 - 19 | 19-21 | 2.36 | 32.64 | 30.28 | 65 | R | C1 - L | 400-04 | SL | 16-1 | 3 | 2 / 1 | 210-08 | BL/WT |  | 8 - 1 | 302-01 | SL/WT | 2:I.ch8+ / 20 | m = 10 b = -1 |  |  |  |  | -0.21712 | 16.9568 | -26.3552
[5 empty] | I + | 21 | O | H2 - 20 | 20-22 | 2.38 | 32.69 | 30.31 | 65 | O | C2 - A | 400-05 | BK | 17-1 |  |  |  |  | CH 1+ |  |  |  |  | 4.22 | 1.8468 | 0.057 | 0.18468 | 0.018468
[5 empty] | V - | 21 | R | H2 - 21 | 19-22 | 2.35 | 32.64 | 30.29 | 65 | R | C2 - B | 400-05 | RD | 16-2 | 2 | 2 / 3 | 210-08 | WT/BL |  | 8 - 2 |  | RTN | RTN / 9 | 77.35 | 6.6887 | 0.107 | 0.66887 | 0.066887
[5 empty] | I - | 21 | O | H2 - 22 | 20-21 | 2.4 | 32.67 | 30.27 | 65 | O | C2 - C | 400-05 | WT | 17-2 |  |  |  |  | CH 1 - |  |  |  |  | 273.2 | 27 |  | 2.7 | 0.27
Coil Outer shell |  | 10 | 1572 | TS24 | V + | 22 | R | H2 - 11 | 11-23 | 2.39 | 32.31 | 29.92 | 66 | R | C2 - D | 400-05 | GN | 18-1 | 3 | 2 / 5 | 210-08 | OR/WT |  | 9 -1 | 302-01 | WT/SL | 2:I.ch9+ / 22 | m = 10 b = 1 |  |  |  |  | -0.21709 | 16.9487 | -26.1693
[5 empty] | I + | 22 | V | H2 - 12 | 12-24 | 2.38 | 32.3 | 29.92 | 66 | V | C2 - E | 400-05 | BN | 19-1 |  |  |  |  | CH 2+ |  |  |  |  | 4.22 | 1.8362 | -0.0247 | 0.18362 | 0.018362
[5 empty] | V - | 22 | R | H2 - 23 | 11-24 | 2.38 | 32.3 | 29.92 | 66 | R | C2 - F | 400-05 | BL | 18-2 | 2 | 2 / 6 | 210-08 | WT/OR |  | 9 - 2 | 302-01 | RTN | RTN / 9 | 77.35 | 6.6792 | 0.009 | 0.66792 | 0.066792
[5 empty] | I - | 22 | V | H2 - 24 | 12-23 | 2.36 | 32.32 | 29.96 | 66 | V | C2 - H | 400-05 | OR | 19-2 |  |  |  |  | CH 2 - |  |  |  |  | 273.2 | 27.002 |  | 2.7002 | 0.27002
Coil Outer shell |  | 11 | 1576 | TS25 | V + | 23 | R | H3 - 1 | 1-3 | 2.34 | 32.23 | 29.89 | 67 | R | C2 - K | 400-05 | VT | 20-1 | 3 | 2 / 7 | 210-08 | GN/WT |  | 10 -1 | 302-01 | BL/RD | 2:I.ch10+ / 24 | m = 10 b = 1 |  |  |  |  | -0.21762 | 16.9551 | -26.0112
[5 empty] | I + | 23 | G | H3 - 2 | 2-4 | 2.43 | 32.31 | 29.88 | 67 | G | C2 - L | 400-05 | SL | 21-1 |  |  |  |  | CH 3+ |  |  |  |  | 4.22 | 1.8258 | -0.1055 | 0.18258 | 0.018258
[5 empty] | V - | 23 | R | H3 - 3 | 1-4 | 2.35 | 32.21 | 29.86 | 67 | R | C3 - A | 400-06 | BK | 20-2 | 2 | 2 / 9 | 210-08 | WT/GN |  | 10 - 2 |  | RTN | RTN / 9 | 77.35 | 6.6666 | -0.121 | 0.66666 | 0.066666
[5 empty] | I - | 23 | G | H3 - 4 | 2-3 | 2.47 | 32.32 | 29.85 | 67 | G | C3 - B | 400-06 | RD | 21-2 |  |  |  |  | CH 3 - |  |  |  |  | 273.2 | 27.013 |  | 2.7013 | 0.27013
Coil Outer shell |  | 12 | 1571 | TS26 | V + | 24 | R | H3 - 6 | 6-8 | 2.37 | 32.6 | 30.23 | 68 | R | C3 - C | 400-06 | WT | 22-1 | 3 | 2 / 11 | 210-08 | BN/WT |  | 11 - 1 | 302-01 | RD/BL | 2:I.ch11+ / 26 | m = 10 b = 1 |  |  |  |  | -0.21724 | 16.9463 | -25.9931
[5 empty] | I + | 24 | BL | H3 - 7 | 7-9 | 2.35 | 32.58 | 30.23 | 68 | BL | C3 - D | 400-06 | GN | 23-1 |  |  |  |  | CH 4+ |  |  |  |  | 4.22 | 1.8256 | -0.1071 | 0.18256 | 0.018256
[5 empty] | V - | 24 | R | H3 - 8 | 6-9 | 2.37 | 32.58 | 30.21 | 68 | R | C3 - E | 400-06 | BN | 22-2 | 2 | 2 / 12 | 210-08 | WT/BN |  | 11 - 2 |  | RTN | RTN / 9 | 77.35 | 6.6683 | -0.102 | 0.66683 | 0.066683
[5 empty] | I - | 24 | BL | H3 - 9 | 7-8 | 2.38 | 32.6 | 30.22 | 68 | BL | C3 - F | 400-06 | BL | 23-2 |  |  |  |  | CH 4 - |  |  |  |  | 273.2 | 27.002 |  | 2.7002 | 0.27002
Sorbs SOW |  | 13 | 1570 | TS33 | V + | 25 | R | H3 - 14 | 14-16 | 2.39 | 32.7 | 30.31 | 75 | R | E3 - C | 400-09 | WT | 24-1 | 3 | 2 / 13 | 210-08 | SL/WT |  | 12 - 1 | 302-01 | RD/BL | 2:I.ch12+ / 30 | m = 10 b = 1 |  |  |  |  | -0.21693 | 16.9341 | -25.9623
[5 empty] | I + | 25 | V | H3 - 15 | 15-17 | 2.34 | 32.66 | 30.32 | 75 | V | E3 - D | 400-09 | GN | 25-1 |  |  |  |  | CH 5+ |  |  |  |  | 4.22 | 1.825 | -0.111 | 0.1825 | 0.01825
[5 empty] | V - | 25 | R | H3 - 16 | 14-17 | 2.39 | 32.69 | 30.3 | 75 | R | E3 - E | 400-09 | BN | 24-2 | 2 | 2 / 15 | 210-08 | WT/SL |  | 12 - 2 |  | RD/BN | RTN / 27 | 77.35 | 6.6709 | -0.077 | 0.66709 | 0.066709
[5 empty] | I - | 25 | V | H3 - 17 | 15-16 | 2.35 | 32.66 | 30.31 | 75 | V | E3 - F | 400-09 | BL | 25-2 |  |  |  |  | CH 5 - |  |  |  |  | 273.2 | 27.016 |  | 2.7016 | 0.27016
Sorbs SOW |  | 14 | 1567 | TS34 | V + | 26 | R | H3 - 19 | 19-21 | 2.36 | 32.28 | 29.92 | 76 | R | E3 - H | 400-09 | OR | 26-1 | 3 | 2 / 17 | 210-08 | BL/RD |  | 13 - 1 | 302-01 | RD/OR | 2:I.ch13+ / 32 | m = 10 b = 2 |  |  |  |  | -0.21821 | 16.9659 | -25.8559
[5 empty] | I + | 26 | GR | H3 - 20 | 20-22 | 2.38 | 32.34 | 29.96 | 76 | GR | E3 - J | 400-09 | YL | 27-1 |  |  |  |  | CH 6+ |  |  |  |  | 4.22 | 1.8151 | -0.19 | 0.18151 | 0.018151
[5 empty] | V - | 26 | R | H3 - 21 | 19-22 | 2.35 | 32.3 | 29.95 | 76 | R | E3 - K | 400-09 | VT | 26-2 | 2 | 2 / 18 | 210-08 | RD/BL |  | 13 - 2 |  | RTN | RTN / 27 | 77.35 | 6.6523 | -0.268 | 0.66523 | 0.066523
[5 empty] | I - | 26 | GR | H3 - 22 | 20-21 | 2.4 | 32.44 | 30.04 | 76 | GR | E3 - L | 400-09 | SL | 27-2 |  |  |  |  | CH 6 - |  |  |  |  | 273.2 | 27.01 |  | 2.701 | 0.2701
Supply Helium pipe |  | 15 | 1563 | TS7 | V + |  |  |  | 11-23 | 2.39 |  | -2.39 | 56 | O | D3 - H | 400-03 | OR | 28-1 | 3 | 2 / 19 | 210-08 | OR/RD |  | 14 - 1 | 302-01 | GN/RD | 2:I.ch14+ / 34 | m = 10 b = 1 |  |  |  |  | -0.21746 | 16.9645 | -26.2456
[5 empty] | I + |  |  |  | 12-24 | 2.38 |  | -2.38 | 56 | BR | D3 - J | 400-03 | YL | 29-1 |  |  |  |  | CH 7+ |  |  |  |  | 4.22 | 1.8392 | -0.0014 | 0.18392 | 0.018392
[5 empty] | V - |  |  |  | 11-24 | 2.38 |  | -2.38 | 56 | O | D3 - K | 400-03 | VT | 28-2 | 2 | 2 / 21 | 210-08 | RD/OR |  | 14 - 2 |  | RTN | RTN / 27 | 77.35 | 6.6783 |  | 0.66783 | 0.066783
[5 empty] | I - |  |  |  | 12-23 | 2.36 |  | -2.36 | 56 | BR | D3 - L | 400-03 | SL | 29-2 |  |  |  |  | CH 7 - |  |  |  |  | 273.2 | 26.987 |  | 2.6987 | 0.26987
Return Helium Pipe
|  | 16 | 1566 | TS8 | V + | [7 empty] | 57 | O | A1 - A | 400-10 | BK | 30-1 | 3 | 2 / 23 | 210-08 | GN/RD |  | 15 - 1 | 302-01 | RD/GN | 2:I.ch15+ / 36 | m = 10 b = 1 |  |  |  |  | -0.21685 | 16.9433 | -26.215
[5 empty] | I + | [7 empty] | 57 | Y | A1 - B | 400-10 | RD | 31-1 |  |  |  |  | CH 8+ |  |  |  |  | 4.22 | 1.8396 | 0 | 0.18396 | 0.018396
[5 empty] | V - | [7 empty] | 57 | O | A1 - C | 400-10 | WT | 30-2 | 2 | 2 / 24 | 210-08 | RD/GN |  | 15 - 2 |  | RTN | RTN / 27 | 77.35 | 6.6843 | 0.06 | 0.66843 | 0.066843
[5 empty] | I - | [7 empty] | 57 | Y | A1 - D | 400-10 | GN | 31-2 |  |  |  |  | CH 8 - |  |  |  |  | 273.2 | 27.006 |  | 2.7006 | 0.27006
3
[18 empty] | MAG-TS-03 |  |  |  |  | J3 | TS-PLC-03 |  |  | slot 3 / pin #
Current lead Cold End Neg Lead |  | 17 | 1568 | TS27 | V + | [7 empty] | 69 | O | C3 - H | 400-06 | OR | 0-1 | 3 | 3 / 1 |  |  |  | 0 - 1 | 303-01 | BL/WT | 3:I.ch0+ / 2 | m = 10 b = 1 |  |  |  |  | -0.21598 | 16.917 | -26.16
[5 empty] | I + | [7 empty] | 69 | G | C3 - J | 400-06 | YL | 1-1 |  |  |  |  | CH 1+ |  |  |  |  | 4.22 | 1.839 | -0.002 | 0.1839 | 0.01839
[5 empty] | V - | [7 empty] | 69 | O | C3 - K | 400-06 | WT | 0-2 | 2 | 3 / 3 |  |  |  | 0 - 2 |  | RTN | RTN / 9 | 77.35 | 6.6901 | 0.12 | 0.66901 | 0.066901
[5 empty] | I - | [7 empty] | 69 | G | C3 - L | 400-06 | SL | 1-2 |  |  |  |  | CH 1 - |  |  |  |  | 273.2 | 27.009 |  | 2.7009 | 0.27009
Current Lead Cold End Pos Lead |  | 18 | 1575 | TS28 | V + | [7 empty] | 70 | O | E1 - A | 400-07 | BK | 2-1 | 3 | 3 / 5 |  |  |  | 1 - 1 | 303-01 | WT/BL | 3:I.ch1+ / 4 | m = 10 b = 1 |  |  |  |  | -0.21678 | 16.9412 | -26.2018
[5 empty] | I + | [7 empty] | 70 | BL | E1 - B | 400-07 | RD | 3-1 |  |  |  |  | CH 2+ |  |  |  |  | 4.22 | 1.839 | -0.003 | 0.1839 | 0.01839
[5 empty] | V - | [7 empty] | 70 | O | E1 - C | 400-07 | WT | 2-2 | 2 | 3 / 6 |  |  |  | 1 - 2 |  | RTN | RTN / 9 | 77.35 | 6.6841 | 0.06 | 0.66841 | 0.066841
[5 empty] | I - | [7 empty] | 70 | BL | E1 - D | 400-07 | GN | 3-2 |  |  |  |  | CH 2 - | [5 empty] | 27.004
Temperature Sensor N2 Outer Radiation Shield
| > 60 K | 22 | 1555 | TS29 | V + | 27 | R | N1 - 1 | 1-3 | 0.24 | 32.21 | 31.97 | 71 | R | E1- F | 400-07 | BL | 4-1 (CHANGED) | 3 | 3 / 7 |  |  |  | 2 - 1 | 303-01 | OR/WT | 3:I.ch2+ / 6 | m = 10 b = 3 |  |  |  |  | -0.21809 | 16.9642 | -25.8176
[5 empty] | I + | 27 | W | N1 - 2 | 2-4 | 0.2 | 32.28 | 32.08 | 71 | W | E1 - H | 400-07 | OR | 5-1 |  |  |  |  | CH 3+ |  |  |  |  | 4.22 | 1.8129 | -0.206 | 0.18129 | 0.018129
[5 empty] | V - | 27 | R | N1 - 3 | 1-4 | 0.2 | 32.19 | 31.99 | 71 | R | E1 - J | 400-07 | YL | 4-2 | 2 | 3 / 9 |  |  |  | 2 - 2 |  | RTN | RTN / 9 | 77.35 | 6.65 | -0.292 | 0.665 | 0.0665
[5 empty] | I - | 27 | W | N1 - 4 | 2-3 | 0.22 | 32.29 | 32.07 | 71 | W | E1 - K | 400-07 | VL | 5-2 |  |  |  |  | CH 3 - |  |  |  |  | 273.2 | 26.994 |  | 2.6994 | 0.26994
|  | 23 | 1560 | TS30 | V + | 28 | O | N1 - 6 | 6-8 | 0.15 | 32.38 | 32.23 | 72 | O | E1 - L | 400-08 | SL | 6-1 | 3 | 3 / 11 |  |  |  | 3 - 1 | 303-01 | WT/OR | 3:I.ch3+ / 8 | m = 10 b = 3 |  |  |  |  | -0.21839 | 16.9753 | -25.853
[5 empty] | I + | 28 | B | N1 - 7 | 7-9 | 0.15 | 32.35 | 32.2 | 72 | B | E2 - A | 400-08 | BK | 7-1 |  |  |  |  | CH 4+ |  |  |  |  | 4.22 | 1.8139 | -0.198 | 0.18139 | 0.018139
[5 empty] | V - | 28 | O | N1 - 8 | 6-9 | 0.21 | 32.35 | 32.14 | 72 | O | E2 - B | 400-08 | RD | 6-2 | 2 | 3 / 12 |  |  |  | 3 - 2 |  | RTN | RTN / 9 | 77.35 | 6.6482 | -0.311 | 0.66482 | 0.066482
[5 empty] | I - | 28 | B | N1 - 9 | 7-8 | 0.15 | 32.37 | 32.22 | 72 | B | E2 - C | 400-08 | WT | 7-2 |  |  |  |  | CH 4 - |  |  |  |  | 273.2 | 26.985 |  | 2.6985 | 0.26985
|  | 24 | 1561 | TS31 | V + | 29 | O | N1 - 14 | 14-16 | 0.18 | 32.59 | 32.41 | 73 | O | E2 - E | 400-08 | BN | 8-1 | 3 | 3 / 13 |  |  |  | 4 - 1 | 303-01 | GN/WT | 3:I.ch4+ / 12 | m = 10 b = 2 |  |  |  |  | -0.21785 | 16.9651 | -25.9931
[5 empty] | I + | 29 | BR | N1 - 15 | 15-17 | 0.16 | 32.55 | 32.39 | 73 | BR | E2 - F | 400-08 | BL | 9-1 |  |  |  |  | CH 5+ |  |  |  |  | 4.22 | 1.8236 | -0.122 | 0.18236 | 0.018236
[5 empty] | V - | 29 | O | N1 - 16 | 14-17 | 0.11 | 32.57 | 32.46 | 73 | O | E2 - H | 400-08 | OR | 8-2 | 2 | 3 / 15 |  |  |  | 4 - 2 |  |  | RTN / 9 | 77.35 | 6.6613 | -0.175 | 0.66613 | 0.066613
[5 empty] | I - | 29 | BR | N1 - 17 | 15-16 | 0.09 | 32.56 | 32.47 | 73 | BR | E2 - J | 400-08 | YL | 9-2 |  |  |  |  | CH 5 - |  |  |  |  | 273.2 | 26.99 |  | 2.699 | 0.2699
|  | 25 | 1565 | TS32 | V + | 30 | O | N1 - 19 | 19-21 | 0.1 | 32.34 | 32.24 | 74 | O | E2 - K | 400-08 | VT | 10-1 | 3 | 3 / 17 |  |  |  | 5 - 1 | 303-01 | WT/GN | 3:I.ch5+ / 14 | m = 10 b = 1 |  |  |  |  | -0.21722 | 16.9448 | -25.9437
[5 empty] | I + | 30 | R | N1 - 20 | 20-22 | 0.13 | 32.33 | 32.2 | 74 | R | E2 - L | 400-08 | SL | 11-1 |  |  |  |  | CH 6+ |  |  |  |  | 4.22 | 1.8227 | -0.129 | 0.18227 | 0.018227
[5 empty] | V - | 30 | O | N1 - 21 | 19-22 | 0.06 | 32.31 | 32.25 | 74 | O | E3 - A | 400-09 | BK | 10-2 | 2 | 3 / 18 |  |  |  | 5 - 2 |  | RTN | RTN / 9 | 77.35 | 6.6654 | -0.133 | 0.66654 | 0.066654
[5 empty] | I - | 30 | R | N1 - 22 | 20-21 | 0.05 | 32.34 | 32.29 | 74 | R | E3 - B | 400-09 | RD | 11-2 |  |  |  |  | CH 6 - |  |  |  |  | 273.2 | 26.997 |  | 2.6997 | 0.26997
Temperature Sensor N2 Inner Radiation Shield | > 60 K
|  | 26 | 1569 | TS12 | V + | 32 | BR | N2 - 6 | 6-8 | 0.15 | 32.26 | 32.11 | 61 | BR | A3 - A | 400-12 | BK | 12-1 | 3 | 3 / 19 |  |  |  | 6 - 1 | 303-01 | BN/WT | 3:I.ch6+ / 16 | m = 10 b = -2 |  |  |  |  | -0.2168 | 16.9534 | -26.5197
[5 empty] | I + | 32 | Y | N2 - 7 | 7-9 | 0.15 | 32.25 | 32.1 | 61 | Y | A3 - B | 400-12 | RD | 13-1 |  |  |  |  | CH 7+ |  |  |  |  | 4.22 | 1.8573 | 0.139 | 0.18573 | 0.018573
[5 empty] | V - | 32 | BR | N2 - 8 | 6-9 | 0.21 | 32.24 | 32.03 | 61 | BR | A3 - C | 400-12 | WT | 12-2 | 2 | 3 / 21 |  |  |  | 6 - 2 |  | RTN | RTN / 9 | 77.35 | 6.701 | 0.234 | 0.6701 | 0.06701
[5 empty] | I - | 32 | Y | N2 - 9 | 7-8 | 0.15 | 32.26 | 32.11 | 61 | Y | A3 - D | 400-12 | GN | 13-2 |  |  |  |  | CH 7 - |  |  |  |  | 273.2 | 27.005 |  | 2.7005 | 0.27005
Current Lead Pot Return |  | 27 | 1573 | TS13 | V + | 33 | BR | N2 - 14 | 14-16 | 0.18 | 32.54 | 32.36 | 62 | BR | A3 - H | 400-12 | OR | 14-1 | 3 | 3 / 23 |  |  |  | 7 - 1 | 303-01 | WT/BN | 3:I.ch7+ / 18 | m = 10 b = -2 |  |  |  |  | -0.21654 | 16.9406 | -26.4275
[5 empty] | I + | 33 | BL | N2 - 15 | 15-17 | 0.16 | 32.54 | 32.38 | 62 | BL | A3 - J | 400-12 | YL | 15-1 |  |  |  |  | CH 8+ |  |  |  |  | 4.22 | 1.853 | 0.106 | 0.1853 | 0.01853
[5 empty] | V - | 33 | BR | N2 - 16 | 14-17 | 0.11 | 32.55 | 32.44 | 62 | BR | A3 - K | 400-12 | VT | 14-2 | 2 | 3 / 24 |  |  |  | 7 - 2 |  | RTN | RTN / 9 | 77.35 | 6.6997 | 0.221 | 0.66997 | 0.066997
[5 empty] | I - | 33 | BL | N2 - 17 | 15-16 | 0.09 | 32.55 | 32.46 | 62 | BL | A3 - L | 400-12 | SL | 15-2 |  |  |  |  | CH 8 - |  |  |  |  | 273.2 | 27.017 |  | 2.7017 | 0.27017
J4
Current Lead Pot Supply |  | 28 | 1574 | TS9 | V + | [7 empty] | 58 | BR | A1 - H | 400-10 | OR | 16-1 | 3 | 4 / 1 |  |  |  | 8 - 1 | 303-01 | SL/WT | 3:I.ch8+ / 20 | m = 10 b = 2 |  |  |  |  | -0.21756 | 16.9545 | -25.9635
[5 empty] | I + | [7 empty] | 58 | GR | A1 - J | 400-10 | YL | 17-1 |  |  |  |  | CH 1+ |  |  |  |  | 4.22 | 1.8229 | -0.1281 | 0.18229 | 0.018229
[5 empty] | V - | [7 empty] | 58 | BR | A1 - K | 400-10 | VT | 16-2 | 2 | 4 / 3 |  |  |  | 8 - 2 |  | RTN | RTN / 9 | 77.35 | 6.6633 | -0.154 | 0.66633 | 0.066633
[5 empty] | I - | [7 empty] | 58 | GR | A1 - L | 400-10 | SL | 17-2 |  |  |  |  | CH 1 - |  |  |  |  | 273.2 | 26.999 |  | 2.6999 | 0.26999
LN2 Return |  | 29 | 1578 | TS10 | V + | [7 empty] | 59 | BR | A2 - A | 400-11 | BK |  | 3 | 4 / 5 |  |  |  | 9 -1 | 303-01 | WT/SL | 3:I.ch9+ / 22 | m = 10 b = 2 |  |  |  |  | -0.21841 | 16.9779 | -26.0017
[5 empty] | I + | [7 empty] | 59 | W | A2 - B | 400-11 | RD | 18-1 |  |  |  |  | CH 2+ |  |  |  |  | 4.22 | 1.8228 | -0.128 | 0.18228 | 0.018228
[5 empty] | V - | [7 empty] | 59 | BR | A2 - C | 400-11 | WT | 19-1 | 2 | 4 / 6 |  |  |  | 9 - 2 |  | RTN | RTN / 9 | 77.35 | 6.6576 | -0.213 | 0.66576 | 0.066576
[5 empty] | I - | [7 empty] | 59 | W | A2 - D | 400-11 | GN | 18-2 |  |  |  |  | CH 2- |  |  |  |  | 273.2 | 27.003 |  | 2.7003 | 0.27003
19-2
LN2 midpoint |  | 30 | 1628 | TS11 | V + | 31 | BR | N2 - 1 | 1-3 | 0.24 | 32.03 | 31.79 | 60 | BR | A2 - H | 400-11 | OR | 20-1 | 3 | 4 / 7 |  |  |  | 10 -1 | 303-01 | BL/RD | 3:I.ch10+ / 24 | [5 empty] | -0.21822 | 16.9773 | -26.1549
[5 empty] | I + | 31 | R | N2 - 2 | 2-4 | 0.2 | 32.1 | 31.9 | 60 | R | A2 - J | 400-11 | YL | 21-1 |  |  |  |  | CH 3+ |  |  |  |  | 4.22 | 1.8323 | -0.055 | 0.18323 | 0.018323
[5 empty] | V - | 31 | BR | N2 - 3 | 1-4 | 0.2 | 32 | 31.8 | 60 | BR | A2 - K | 400-11 | VT | 20-2 | 2 | 4 / 9 |  |  |  | 10 - 2 |  | RTN | RTN / 9 | 77.35 | 6.6682 | 0.03 | 0.66682 | 0.066682
[5 empty] | I - | 31 | R | N2 - 4 | 2-3 | 0.22 | 32.12 | 31.9 | 60 | R | A2 - L | 400-11 | SL | 21-2 |  |  |  |  | CH 3 - |  |  |  |  | 273.2 | 27.01 |  | 2.701 | 0.2701
Supply |  | 31 |  | AST1a | I+ | [7 empty] | 109 | BR | F2 - J | 400-14 | YL | 22-1 |  | 4 / 11 |  |  |  | 11 - 1 | 303-01 | RD/BL | 3:I.ch11+ / 26
[5 empty] | V+ | [7 empty] | 109 | BR | F2 - K | 400-14 | VT | 23-1 |  |  |  |  | CH 4+
[5 empty] | I- | [7 empty] | 109 | G | F2 - L | 400-14 | SL | 23-2 |  | 4 / 12 |  |  |  | 11 - 2 |  | RTN | RTN / 9
[18 empty] | 23-2 |  |  |  |  | CH 4 -
[20 empty] | 4 / 13 |  |  |  | 12 - 1 | 303-01 | RD/BL | 3:I.ch12+ / 30
Return |  | 32 |  | AST2 | I+ | [7 empty] | 92 | BR | F1 - A | 400-13 | BK | 24-1 |  |  |  |  | CH 5+
[5 empty] | V+ | [7 empty] | 92 | BR | F1 - B | 400-13 | RD | 25-1 |  | 4 / 15 |  |  |  | 12 - 2 |  | RD/BN | RTN / 27
[5 empty] | I- | [7 empty] | 92 | G | F1 - C | 400-13 | WT | 24-2 |  |  |  |  | CH 5 -
25-2
| DIODE 2 WIRE | 19 |  | CU1 | I+ | [7 empty] | 94 | BR | F2 - A | 400-14 | BK | 26-1 |  | 4 / 17 |  |  |  | 13 - 1 | 303-01 | RD/OR | 3:I.ch13+ / 32
[5 empty] | I- | [7 empty] | 94 | O | F2 - B | 400-14 | RD | 27-1 |  |  |  |  | CH 6+
[18 empty] | 26-2 |  | 4 / 18 |  |  |  | 13 - 2 |  | RTN | RTN / 27
[18 empty] | 27-2 |  |  |  |  | CH 6 -
| DIODE 2 WIRE | 20 |  | CU2 | I+ | [7 empty] | 95 | BR | F2- E | 400-14 | BN | 28-1 |  | 4 / 19 |  |  |  | 14 - 1 | 303-01 | GN/RD
[5 empty] | I- | [7 empty] | 95 | Y | F2- F | 400-14 | BL | 29-1 |  |  |  |  | CH 7+ |  |  |  | 3:I.ch14+ / 34
[18 empty] | 28-2 |  | 4 / 21 |  |  |  | 14 - 2 |  | RTN
[18 empty] | 29-2 |  |  |  |  | CH 7 - |  |  |  | RTN / 27
| 3 WIRE | 21 |  | AST4 | I+ | [7 empty] | 93 | O | F1 - H | 400-13 | OR | 30-1 |  | 4 / 23 |  |  |  | 15 - 1 | 303-01 | RD/GN
[5 empty] | I- | [7 empty] | 93 | V | F1 - J | 400-13 | YL | 31-1 |  |  |  |  | CH 8+ |  |  |  | 3:I.ch15+ / 36
[5 empty] | V+ | [7 empty] | 93 | O | F1 - K | 400-13 | VT | 30-2 |  | 4 / 24 |  |  |  | 15 - 2 |  | RTN
[18 empty] | 31-2 |  |  |  |  | CH 8 - |  |  |  | RTN / 27
ON Helium Return pipe
Sensors in CCR
He 4.5K supply
He Cold return
He warm return
He Mag Supply
He CL-Pot Supply
He CL_Pot Return
LN2 Supply
Ln2 Return
Sensors in HX and transfer lines
LN2 Exhaust
He Cooldown

## Temp resolution
Tolerance
Typical RH-Fe |  | +/- 0.1K
T | R
K | ohms |  |  | Typical fit
4.22 | 1.8379 |  |  | T = -0.2178xR^2 + 16.9703xR - 26.234
77.35 | 6.6759
273.2 | 27.006
Constant Current Source |  | Tolerance |  | delta T
100uA | 0.0001 | 1e-05 |  | 0.0001697
Voltage I/O conversion |  | Tolerance
+/- 10.25V |  | 0.00032 |  | 0.0054305
0 -10.25V |  | 0.00016 |  | 0.0027152
0 to 5.125V |  | 8e-05 |  | 0.0013576
Dataforth Signal
30-01D | accuracy | 3e-06 | 0.03% | 5.0911e-05
| temp drift | 5e-06 | 1uV/C | 8.4851e-05
Sum of delta T
|  |  |  | 0.0029359 | K | Much less than sensor Tolerance
Sum of Squares
0.0027223

## CLEOII-Chassis Layout
PLC I/O Usage Table, Chassis 2, Remote A | [9 empty] | Added signals
Magnets | Quench Detector / PSU | Temperature sensors | Temperature sensors | CL mass Flow Controller | JT motor relays | Reset and Pwr enabled | Load Sensors | Load Sensors | 8 Channels | Changed from diodes to PT-102
Module | IV16D | IF16 | IF16 | OF4 | OW16I | OW16I | IF16 | IF16 | IF16 | Changed name for signal
Configuration |  | Single_Ended voltage | Single_Ended Voltage | 0-5 volts output |  |  | Single_Ended_Mode Voltage | Single_Ended_Mode Voltage/Current | Differential_Mode Voltage
Chasis Slot # | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
Channel
0 | QD_U1 | TS_1 | TS_27_CL_Neg | Setpnt_m_flow_A | MOT_JTV1_Close | Danf_QD_Reset | Radial - A | Axial - T | Current Lead A
1 | QD_L1 | TS_2 | TS_28_CL_POS | Setpnt_m_flow_B | MOT_JTV1_Open | Vacuum_power | Radial - B | Axial - V | Current Lead B
2 | QD_U2 | TS_3 | TS_29 (LN2) |  | MOT_JTV2_Close | Keep_Alive_Reset | Radial - C | Axial - W | Coil 2 Inner
3 | QD_L2 | TS_4 | TS_30 (LN2) |  | MOT_JTV2_Open | Keep_Alive_Control | Radial - D | Axial - X | Coil 1 Outer
4 | QD_U3 | TS_5 | TS_31 (LN2) | N/A | MOT_JTV3_Close | Power_liquid_level_meter | Radial - E |  | Flux Loop 1
5 | QD_L3 | TS_6 | TS_32 (LN2) | N/A | MOT_JTV3_Open | PLC_Fast_Relay | Radial - F |  | Flux Loop 2
6 | QD_U4 | TS_21 | TS_12 (LN2) | N/A | MOT_JTV4_Close | PLC_Slow_Relay | Radial - G |  | Magnet
7 | QD_L4 | TS_22 | TS_13 (LN2) | N/A | MOT_JTV4_Open | MPS_CPU_Reset | Radial - H |  | PSU_Output_Voltage
8 | QD_Sum | TS_23 | TS_9 (LN2_Supply) | N/A | MOT_JTV5_Close | CCS_Control_1 | Radial - J | LL_He | N/A
9 | UPS | TS_24 | TS_10 (LN2_Return) | N/A | MOT_JTV5_Open | CCS_Control_2 | Radial - K | LL_N2 | N/A
10 |  | TS_25 | TS_11 (LN2_Center) | N/A | MOT_JTV6-Close | CCS_Control_3 | Radial - L | Mass_Flow_CL_A | N/A
11 | JVT Mtrs KS remote | TS_26 | ATM1a | N/A | MOT_JTV6_Open | CCS_Control_4 | Radial - M | Mass_Flow_CL_B | N/A
12 | PSU_Control_pwr_on | TS_33 | AST2 | N/A | Mot_JTV7_Close | CCS_Control_5 | Radial - N | Vacuum | N/A
13 | PSU_Main_pwr_on | TS_34 | CU1 | N/A | Mot_JTV7_Open | CCS_Control_6 | Radial - P | Pressure He | N/A
14 | PSU_Sum_Interlocks | TS_7_He_Supply | Cu2 | N/A | Warm_return_Close | MPS_Comm_On | Radial - R | Pressure LN2 | N/A
15 | PSU_ready | TS_8_He_Return | AST4 | N/A | Warm_return_Open |  | Radial - S | I_coarse | N/A
Dataforth Signal Conditioner Table SCM5B modules
Channel -> board # |  | 1 | 3 |  | Label |  | 5 | 7 | 9
1 |  | 30-01D [+/- 10mV +/- 10V] | 30-01D [+/- 10mV +/- 10V] |  | JT1 LN2 Cooldown |  | 38-31D [+/-10mV - +/-3.33V] | 38-31D [+/-10mV - +/-3.33V] | 31-09D [+/-40V - +/-10V]
2 |  | 30-01D [+/- 10mV +/- 10V] | 30-01D [+/- 10mV +/- 10V] |  | JT2 LN2 Top Fill |  | 38-31D [+/-10mV - +/-3.33V] | 38-31D [+/-10mV - +/-3.33V] | 31-09D [+/-40V - +/-10V]
3 |  | 30-01D [+/- 10mV +/- 10V] | 30-01D [+/- 10mV +/- 10V] |  | JT3 He Cooldown |  | 38-31D [+/-10mV - +/-3.33V] | 38-31D [+/-10mV - +/-3.33V] | 31-09D [+/-40V - +/-10V]
4 |  | 30-01D [+/- 10mV +/- 10V] | 30-01D [+/- 10mV +/- 10V] |  | JT4 He Bottom Fill |  | 38-31D [+/-10mV - +/-3.33V] | 38-31D [+/-10mV - +/-3.33V] | 31-09D [+/-40V - +/-10V]
5 |  | 30-01D [+/- 10mV +/- 10V] | 30-01D [+/- 10mV +/- 10V] |  | JT5 He Top Fill |  | 38-31D [+/-10mV - +/-3.33V] |  | 31-09D [+/-40V - +/-10V]
6 |  | 30-01D [+/- 10mV +/- 10V] | 30-01D [+/- 10mV +/- 10V] |  | JT6 He Cold Return |  | 38-31D [+/-10mV - +/-3.33V] |  | 31-09D [+/-40V - +/-10V]
7 |  | 30-01D [+/- 10mV +/- 10V] | 30-01D [+/- 10mV +/- 10V] |  | JT7 CCL Pot |  | 38-31D [+/-10mV - +/-3.33V] |  | 31-09D [+/-40V - +/-10V]
8 |  | 30-01D [+/- 10mV +/- 10V] | 30-01D [+/- 10mV +/- 10V] |  | Warm return Valve |  | 38-31D [+/-10mV - +/-3.33V]
Channel -> board # |  | 2 | 4 |  |  |  | 6 | 8
1 |  | 30-01D [+/- 10mV +/- 10V] | 30-01D [+/- 10mV +/- 10V] |  |  |  | 38-31D [+/-10mV - +/-3.33V]
2 |  | 30-01D [+/- 10mV +/- 10V] | 30-01D [+/- 10mV +/- 10V] |  |  |  | 38-31D [+/-10mV - +/-3.33V]
3 |  | 30-01D [+/- 10mV +/- 10V] | 30-01D [+/- 10mV +/- 10V] |  |  |  | 38-31D [+/-10mV - +/-3.33V] | 31-03D [+/- 10V +/- 10V]
4 |  | 30-01D [+/- 10mV +/- 10V] | 30-01D [+/- 10mV +/- 10V] |  |  |  | 38-31D [+/-10mV - +/-3.33V] | 31-03D [+/- 10V +/- 10V]
5 |  | 30-01D [+/- 10mV +/- 10V] | 30-01D [+/- 10mV +/- 10V] |  |  |  | 38-31D [+/-10mV - +/-3.33V] | 31-03D [+/- 10V +/- 10V]
6 |  | 30-01D [+/- 10mV +/- 10V] | 30-01D [+/- 10mV +/- 10V] |  |  |  | 38-31D [+/-10mV - +/-3.33V] | 38-35D [+/- 10mV +/- 10V]
7 |  | 30-01D [+/- 10mV +/- 10V] | 30-01D [+/- 10mV +/- 10V] |  |  |  | 38-31D [+/-10mV - +/-3.33V] | 38-35D [+/- 10mV +/- 10V]
8 |  | 30-01D [+/- 10mV +/- 10V] | 30-01D [+/- 10mV +/- 10V] |  |  |  | 38-31D [+/-10mV - +/-3.33V] | 31-03D [+/- 10V +/- 10V]
PLC I/O Usage Table, Chassis 2, Remote B
Magnets | Temperature sensors | Relays | LVDT | LVDT
Module | IF16 | OW16I | IF16 | IF16
Configuration | Single_Ended voltage |  | Differential_Mode Current | Differential_Mode Current
Chasis Slot # | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9
Channel
0 | PT_LN2_ Supply | JTV_HX_LN2_Close | LVDT_JTV1 | LVDT_HX_LN2
1 | PT_LN2_Return | JTV_HX_LN2_Open | LVDT_JTV2 | LVDT_HX_GHE
2 | PT_CL_A Temp Warm | JTV_HX_GHe_Close | LVDT_JTV3
3 | PT_CL_B Temp Warm | JTV_HX_GHe_Open | LVDT_JTV4
4 | PT_HX_GN2__exhaust | SORB Heater 1 | LVDT_JTV5
5 | PT_HX_Ghe_Mix | SORB Heater 2 | LVDT_JTV6
6 |  | Current Lead Heater A | LVDT_JTV7
7 |  | Current Lead Heater B | Warm_Return_Position
8 | TD_He_Magnet_supply |  | N/A | N/A
9 | TD_He_Cold_supply |  | N/A | N/A
10 | TD_He_Cold _Return |  | N/A | N/A
11 | TD_He_Warm_Return |  | N/A | N/A
12 | TD_He_CL_Pot_Supply |  | N/A | N/A
13 | TD_He_Reservoir |  | N/A | N/A
14 |  |  | N/A | N/A
15 |  |  | N/A | N/A
Dataforth Signal Conditioner Table SCM5B modules
Channel -> board # | 10 |  | LVDT Modules | LVDT Modules
1 | 30-01D [+/- 10mV +/- 10V] |  | LVC-2500 | LVC-2500
2 | 30-01D [+/- 10mV +/- 10V] |  | LVC-2500 | LVC-2500
3 | 30-01D [+/- 10mV +/- 10V] |  | LVC-2500
4 | 30-01D [+/- 10mV +/- 10V] |  | LVC-2500
5 | 30-01D [+/- 10mV +/- 10V] |  | LVC-2500
6 | 30-01D [+/- 10mV +/- 10V] |  | LVC-2500
7 |  |  | LVC-2500
8
Channel -> board # | 11
1 | 43-05D [+/-5V - +/-10V]
2 | 43-05D [+/-5V - +/-10V]
3 | 43-05D [+/-5V - +/-10V]
4 | 43-05D [+/-5V - +/-10V]
5 | 43-05D [+/-5V - +/-10V]
6 | 43-05D [+/-5V - +/-10V]
7
8
Temp (Rd-Fe) | 30-01D [+/- 10mV +/- 10V]
Strain gauges | 38-31D [+/-10mV +/-3.33V]
Gauges | 31-03D [+/- 10V +/- 10V]
Voltage taps | 31-09D [+/-40V - +/-10V]
PT temps | 34-1562 [RTD +/- 5V]
Diode Temp | 43-05D [+/-5V - +/-10V]

## Voltage taps
Signal description | VT # | Internal Wire # color | Internal connector / pin | Internal Cable # Color | Wire color | Vacuum Feed Thur Connector / Pin | Protection Resistor | Wire Color | Terminal Strip | Wire Color | Dataforth | Dataforth feed out | Wire Color | Terminal Strip | 1756 IF16
Voltage Taps |  |  | 25 pin D connectors |  |  | 10 pin vacuum feedthrough |  |  |  |  | 31-09D | Board # / pin # |  |  | slot # / Channel # / pin #
[9 empty] | MAG-TS-09 |  |  |  |  | TS-PLC-09 |  | Changed name/labels
[5 empty] | R | Rho | L | A
[5 empty] | 4 | 2.1875e-08 | 12500 | 6.836e-05
Section 7.4 Page 85 |  |  |  |  | 4 | 2.7e-08 | 10127.4074 | 6.836e-05
[5 empty] | 4.0001 | 2.7e-08 | 12500 | 8.4372e-05
at Top of current lead A Warm end | VT 6 | [5 empty] | 200k Resistors |  | 1-1 |  | 3 | 9 / 1 |  | 0-1 | 9:I.ch0+ / 2
at Bottom of current lead A | VT 4 |  |  | B86 | R | B2 - L | 200k Resistors |  | 2-1 |  | 2 | 9 / 3 |  | 1-1 | 9:I.ch0- / 2
at Top of current lead B Warm end | VT 7 | [5 empty] | 200k Resistors |  | 3-1 |  | 3 | 9 / 5 |  | 2-1 | 9:I.ch1+ / 2
at Bottom of current lead B | VT 5 |  |  | B91 | B | B3 - L | 200k Resistors |  | 4-1 |  | 2 | 9 / 6 |  | 3-1 | 9:I.ch1- / 2
Coil exit point Inner layer | VT 2 | 5 G | H4 - 25 | 83 G | BL | B2 - C | 200k Resistors |  | 5-1 |  | 3 | 9 / 7 |  | 5-1 | 9:I.ch2+ / 2
Center tap | VT 1 | 3 O | H4 - 7 | 79 O | O | B1 - E | 200k Resistors |  | 6-1 |  | 2 | 9 / 9 |  | 6-1 | 9:I.ch2- / 2
Center tap | VT 1 | 3 O | H4 - 7 | 79 O | O | B1 - E |  |  | 6-1 |  | 3 | 9 / 11 |  | 7-1 | 9:I.ch3+ / 2
Coil exit point outer layer | VT 3 | 1 BR | H4 - 1 | 77 B | B | B1 - A | 200k Resistors |  | 7-1 |  | 2 | 9 / 12 |  | 8-1 | 9:I.ch3- / 2
Flux Loop 1 inner coil | FL 1 | 10 B | H4 - 5 | 80 B | B | B1 - J | 200k Resistors |  | 8-1 |  | 3 | 9 / 13 |  | 9-1 | 9:I.ch4+ / 2
|  | 10 BR | H4 - 9 | 81 BR | BR | B1 - K | 200k Resistors |  | 9-1 |  | 2 | 9 / 15 |  | 10-1 | 9:I.ch4- / 2
Flux loop 2 outer coil | FL 2 | 11 O | H4 - 16 | 85 O | O | B2 - H | 200k Resistors |  | 10-1 |  | 3 | 9 / 17 |  | 11-1 | 9:I.ch5+ / 2
|  | 11 BR | H4 - 23 | 90 BR | BR | B3 - H | 200k Resistors |  | 11-1 |  | 2 | 9 / 18 |  | 12-1 | 9:I.ch5- / 2
Quench Detector
at Top of current lead A Warm end | VT 6 | [7 empty] | 1-2 |  | D1H
at Bottom of current lead A | VT 4 |  |  | B86 | R | B2 - L |  |  | 2-2 |  | D1HM
at Top of current lead B Warm end | VT 7 | [7 empty] | 3-2 |  | D4H
at Bottom of current lead B | VT 5 |  |  | B91 | B | B3 - L |  |  | 4-2 |  | D4HM
Coil exit point Inner layer | VT 2 | 5 G | H4 - 25 | 83 G | BL | B2 - C |  |  | 5-2 |  | D2H
Center tap | VT 1 | 3 O | H4 - 7 | 79 O | O | B1 - E |  |  | 6-2 |  | D2HM
Center tap | VT 1 | 3 O | H4 - 7 | 79 O | O | B1 - E |  |  | 6-2 |  | D2LM
Coil exit point outer layer | VT 3 | 1 BR | H4 - 1 | 77 B | B | B1 - A |  |  | 7-2 |  | D2L
Flux Loop 1 inner coil | FL 1+ | 10 B | H4 - 5 | 80 B | B | B1 - J |  |  | 8-2 |  | D3H
20 turns | FL 1- | 10 BR | H4 - 9 | 81 BR | BR | B1 - K |  |  | 9-2 |  | D3HM
Flux loop 2 outer coil | FL 2+ | 11 O | H4 - 16 | 85 O | O | B2 - H |  |  | 10-2 |  | D3LM
20 turns | FL 2- | 11 BR | H4 - 23 | 90 BR | BR | B3 - H |  |  | 11-2 |  | D3L
Heater Mid-plane
6 turn
200Vdc @ 2A 400 Watts Continous
200Vdc @ 4A =800 Watts (10 secs)

## Coil RT Measurement
|  | Measured |  | Calculated Ohms | Resivity | Length | Area | Ri | R2 | S1/turn | S2/turn | Turns 1 | Turns 2 | L1 | L2 | Ltotal
|  | Ohms |  |  | Ohms-m | m | m^2 | m | m | m | m |  |  | m | m | m
Multimeter | Lead 1 to lead 2 | 4.12 | Al | 4.9905 | 2.724e-08 | 12523.9214 | 6.836e-05 | 1.548 | 1.564 | 9.7264 | 9.8269 | 640 | 641 | 6224.8773 | 6299.0441 | 12523.9214
Meager #1 | Lead 1 to lead 2 | 3.99 | SSC (Cu) | 43.5728 | 1.67e-08 | 12523.9214 | 4.8e-06
|  |  | total | 4.4777
[8 empty] | Area | Insulation | SSC (11) | ssc (9) | Al area
[8 empty] | mm^2 | 0.2mm all round |  |  | m^2
[8 empty] | 16x5.1 | 8.44 | 4.8 |  | 68.36
[8 empty] | 16x4.89 | 8.356 |  | 3.9273 | 65.9567

## PLC Modules Software
|  | Date 6/14/2018 |  | 0.74
Solid - Cleo II Controls
Qty | Catalog # | Description | Price | Ext. Price
|  | Software |  |  |  |  | Date | PR | PO | Ordered | unit Cost | Total Cost
0 | 9324RLD700NXENE | RS-Logix 5000 Pro Edition | 5500 | 0 |  |  | 2018-07-31 | 378847 |  | 33000
1 | 9701-VWSTENE | RSVIEW Studio | 2810 | 2810
Train gauge | 9701-VWSS000LENE | RSView Enterprise unlimited displays *(Trade in not shown) | 16700 | #VALUE!
4 | 9701-VWSCWAENE | SE Client | 2780 | 11120
1 | AB9518HSW2K5 | Historian | 18600 | 18600
1 | AB9518HCALENE | Historian CLIENT | 554 | 554 | [6 empty] | sub-Total | 0
1 | 9301-OPCSRVENE | KPServer | 2400 | 2400
|  |  | Subtotal: | #VALUE! |  |  | Date |  | PO | Ordered | Unit Cost | Total Cost
Summary of Hardware:
|  | Subsystem of New System : Redundant CLX |  |  | 1756-L62M | LOGIX5562 PROCESSOR WITH 4 MBYTE NV MEMORY
2 | 1756-A4 | 4 SLOT CONTROLLOGIX CHASSIS | 290 | 580 | 1757-SRC1 | 1757 REDUNDANCY MODULE CABLE - 1 METER, 3.28FT
4 | 1756-CNB | CONTROLNET BRIDGE MODULE | 1715 | 6860 | 1757-SRM | SYSTEM REDUNDANCY MODULE
2 | 1756-ENBT | CLX ETHERNET/IP 10/100 BRIDGE MODULE - TWISTED PR | 2356 | 4712 | 1756-CNB | CONTROLNET BRIDGE MODULE
2 | 1756-L72 | LOGIX5572 PROCESSOR WITH 4 MBYTE NV MEMORY | 6864 | 13728 | 1756-ENBT | CLX ETHERNET/IP 10/100 BRIDGE MODULE - TWISTED PR
2 | 1756-N2 | EMPTY SLOT FILLER CARD (ONE FILLER PER PACKAGE) | 25 | 50 | 1756-N2 | EMPTY SLOT FILLER CARD (ONE FILLER PER PACKAGE)
2 | 1756-PA75 | 85-265 VAC POWER SUPPLY (5V @ 13 AMP) | 975 | 1950 | 1756-PA72 | 85-265 VAC POWER SUPPLY (5V @ 10 AMP)
1 | 1756RMC1 | 175 REDUNDANCY MODULE CABLE - 1 METER, 3.28FT | 172 | 172 | 1786-XT | CONTROLNET 75 OHM TERMINATOR PLUG
2 | 1757-RM2 | SYSTEM REDUNDANCY MODULE | 4990 | 9980 | 1786-TPS | CONTROLNET T-TAP/STRAIGHT
2 | 1786-XT | CONTROLNET 75 OHM TERMINATOR PLUG | 11 | 22 | 1756-TBCH | 36 PIN SCREW CLAMP BLOCK WITH STANDARD HOUSING
14 | 1786-TPS | CONTROLNET T-TAP/STRAIGHT | 130 | 1820 | 1756-TBNH | 20 NEMA Screw Clamp Block
1 | 1786-RG6 | CONTROLNET RG6 SHIELD HIGH FLEX CABLE-1000FT ROLL | 568 | 568 | 1756-IV16 | DC (10-30V) Sourcing Input Module
|  |  | Subtotal: | 40442 | 1756-DNB | Device net Bridge
[5 empty] | 1756-OW16I | N.O. ISOLATED RELAY OUTPUT 16 PTS (36 PIN)
|  | Subsystem of New System : Cleo II magnet I/O |  |  | 1756-IF16 | ANALOG INPUT - CURRENT/VOLTAGE 16 PTS (36 PIN)
1 | 1756-A10 | 10 SLOT CONTROLLOGIX CHASSIS | 540 | 540 | 1756-A4 | 4 slot chassis
1 | 1756-CNB | CONTROLNET BRIDGE MODULE | 1312 | 1312 | 1756-A10 | 10 slot chassis
1 | 1756-IV16 | 79-132 VAC INPUT 16 PTS (20 PIN) | 315 | 315 | 1756-OG16 | 0-5 TTL output
1 | 1756-IB16D | 16 channel Digital Input Module | 1397 | 1397 | MVI94-GSC | Prosoft ASCII Generic Serial Communication Module
2 | 1756-OW16I | N.O. ISOLATED RELAY OUTPUT 16 PTS (36 PIN) | 550 | 1100 | 1794-ACN15 | SINGLE MEDIA ADAPTER FOR CONTROLNET
5 | 1756-IF16 | ANALOG INPUT - CURRENT/VOLTAGE 16 PTS (36 PIN) | 1352 | 6760 | 1794-PS13 | 85-264 VAC TO 24 VDC 1.3A POWER SUPPLY
1 | 1756-PA72 | 85-265 VAC POWER SUPPLY (5V @ 10 AMP) | 805 | 805 | 1756-IF4FXOF2F | 4 channel Hi speed buffered Input
9 | 1756-TBCH | 36 PIN SCREW CLAMP BLOCK WITH STANDARD HOUSING | 75 | 675 | 800T-J91A | 3 position switch
1 | 1794-ACN15 | SINGLE MEDIA ADAPTER FOR CONTROLNET | 671 | 671 | 800T-H33B | 2 position
1 | 1794-PS13 | 85-264 VAC TO 24 VDC 1.3A POWER SUPPLY | 188 | 188 | 1489-A2C020 | Circuit Breaker 2A
1 | MVI94-GSC | Prosoft ASCII Generic Serial Communication Module | 1056 | 1056 | 1489-A2C050 | Circuit Breaker 5A
|  |  | Subtotal: | 14819 | 1489-A2C100 | Circuit Breaker 10A
[5 empty] | 1492-JD3 | AB end anchor
|  | Subsystem of New System : Cryogenics/Other |  |  | 1492-EAJ35 | AB double terminal block
1 | 1756-A10 | 10 SLOT CONTROLLOGIX CHASSIS | 542 | 542 | 1783-US08T | Ethernet switch
1 | 1756-CNB | CONTROLNET BRIDGE MODULE | 1312 | 1312 | 1756-EN2T | Control Logix Ethernet/IP module
2 | 1756-IB16D | 16 channel Digital Input Module | 542 | 1084 | 1756-OF4 | 4 output voltage or current |  |  |  |  | Total | 0
5 | 1756-IF16 | ANALOG INPUT - CURRENT/VOLTAGE 16 PTS (36 PIN) | 1397 | 6985
2 | 1756-OW16I | N.O. ISOLATED RELAY OUTPUT 16 PTS (36 PIN) | 559 | 1118
1 | 1756-PA72 | 85-265 VAC POWER SUPPLY (5V @ 10 AMP) | 805 | 805
9 | 1756-TBCH | 36 PIN SCREW CLAMP BLOCK WITH STANDARD HOUSING | 67 | 603
1 | 1794-ACN15 | SINGLE MEDIA ADAPTER FOR CONTROLNET | 671 | 671
1 | 1794-PS13 | 85-264 VAC TO 24 VDC 1.3A POWER SUPPLY | 190 | 190
1 | MVI94-GSC | Prosoft ASCII Generic Serial Communication Module | 1056 | 1056
|  |  | Subtotal: | 14366
PLC Controller Hardware:
1 |  | Server Windows based | 2875 | 2875
1 |  | Client PC | 2000 | 2000
2 |  | 10-slot chasis | 582 | 1164
2 |  | Redundant PLC L55m23 | 3316.5 | 6633
2 |  | CNB/D | 1732 | 3464
2 |  | PSU | 531 | 1062
2 |  | Taps | 71 | 142
2 |  | Terminators | 5 | 10
1 |  | Redundancy cable | 104 | 104
2 |  | System Redundancy Module | 3206 | 6412
1 |  | Progamming cable | 48 | 48
2 |  | Empty Slot Filler | 13.5 | 27
|  |  | Subtotal | 23941
Note on D25: Steven Lassiter:
Note on D26: Steven Lassiter:

## Signal Conditioners
Date Oct. 26, 2018
|  | DataForth Signal Conditioners |  | 1
Qty | Catalog # | Description | Price | Total Cost | Date | PR/PO
45 | SCM5B-30-01D | Analog Input Narrow bandwidth 10mV in / 10V out | 164 | 7380 |  | 395689
10 | SCM5B-31-03D | Analog Voltage Input Module 10V in / 10V out | 164 | 1640 |  | 395689
8 | SCM5B-31-09D | Analog Voltage Input Module 40V in / 10V out | 164 | 1312 |  | 395689
20 | SCM5B-38-31D | Strain Gage Input Module | 195 | 3900 |  | 395689
0 | SCM5B-43-05D | General Purpose with excitation 5V in / 10V out | 205 | 0
5 | SCM5B-38-35D | 10mV in / 10 V out | 195 | 975 |  | 395689
11 | SCMXCA-004 12 | Ribbon Cable | 39 | 429 |  | 395689
11 | SCMPB-05-02 | Backplane-8 channels with DIN mounts | 279 | 3069 |  | 395689
11 | SCMXIF-DIN | Breakout Board Din Mounted | 69 | 759 |  | 395689
11 | SCMXRK-002 | Rack Mounting Plate 19" | 67 | 737
7 | LVC-2500 | LVDT signal conditioner 24Vdc | 425 | 2975
2 | CMA1D03A | Constant current Source boards 8channels | 410 | 820
2 | SDP5-5-100T | PSU 5 volts | 92.4 | 184.8
2 | SDN10-24-100P | PSU 24V Psu | 292.3 | 584.6
|  |  | SUM | 24765.4
|  |  | procured | 17586
|  |  | outstanding | 7179.4

## Rack Termination cabling
Date 2/07/2007
|  | I/O Rack / Hardware / Cable per Magnet |  |  |  |  | Date | PR/PO | Ordered | unit Cost | Total Cost
Qty per magnet/rack | Catalog # | Description | Price | Disc. Price | Ext. Price
1
|  | Rack Chassis |  | 0.737
1 | XFR-317824 | Vertical Frame 31.562 deep 82.375 Tall 27.562 wide | 663.52 | 663.52 | 663.52
4 | XMA-78 | Mounting Angle Pair | 45 | 45 | 180
6 | XPN-7324 | Panel 73.5" Tall 24" wide | 105 | 105 | 630
1 | DS-XSR-7824 | Front Door | 160 | 160 | 160
1 | DS-XSL-7824 | Rear Door | 160 | 160 | 160
1 | XES-2024 | Equipment Shelf | 36 | 36 | 36
2 | XSP-3178-S | Side Panel | 130 | 130 | 260
1 | XPN-2819 | Top Panel | 40 | 40 | 40
2 | XPN-0324-CAP | Data Cable Access Panel | 28 | 28 | 56
100 | HW-092 | Threaded captive clip nut | 0.4 | 0.4 | 40
100 | HW-103 | #10-32 thread truss head machine screw | 0.28 | 0.28 | 28
1 | 446149 | Latch/Slotted | 4 | 4 | 4
1 | A423008LP | SINGLE DOOR ENCLOSURE HOFFMAN | 750 | 750 | 750
1 | a42p30 | PANEL FOR ENCLOSURE | 175 | 175 | 175
3 | ABA-05 | Panel reducers | 51 | 51 | 153
|  |  | Subtotal: |  | 3335.52 | 16677.6
|  | Rack hardware |  |  |  |  | Date | PR/PO | Ordered | unit Cost | Total Cost
800 | 1492-W4 | Terminal Blocks Single Level | 1.55 | 1.55 | 1240 |  | 2010-03-03 | 289237 / 10-P1267 | 2000 | 0.71 | 1420
100 | 1492-WD3 | Terminal Blocks Double Level | 3.4 | 3.4 | 340 |  | 2010-03-03 | 289237 / 10-P1267 | 400 | 2.07 | 828
200 | 1492-WTF3 | Terminal Blocks Triple Level | 5.05 | 5.05 | 1010 |  | 2010-03-03 | 289237 / 10-P1267 | 900 | 4.7 | 4230
20 | 1492-EB3 | End barrier for W4 | 0.6 | 0.6 | 12 |  | 2010-03-03 | 289237 / 10-P1267 | 50 | 0.6 | 30
30 | 1492-EA35 | End Anchor for W4 | 0.9 | 0.9 | 27 |  | 2010-03-03 | 289237 / 10-P1267 | 50 | 1.22 | 61
30 | 1492-GM35 | Terminal Group marker | 1.85 | 1.85 | 55.5 |  | 2010-03-03 | 289237 / 10-P1267 | 100 | 2.2 | 220
20 | 1492-EBD3 | End Barrier for WD3 | 1.01 | 1.01 | 20.2 |  | 2010-03-03 | 289237 / 10-P1267 | 200 | 0.77 | 154
30 | 1492-EAH35 | End Anchor for WD3 | 3.4 | 3.4 | 102 |  | 2010-03-03 | 289237 / 10-P1267 | 50 | 2.42 | 121
20 | 1492-EBFT3 | End Barrier for WTF3 | 1 | 1 | 20 |  | 2010-03-03 | 289237 / 10-P1267 | 50 | 0.97 | 48.5
15 | 1492-SJT5-20-B | JUMPERS 20-POLE PACAKGE OF 10 | 5.5 | 5.5 | 82.5 |  | 2010-03-03 | 289237 / 10-P1267 | 60 | 5.12 | 307.2
1 | 800T-H33B | KEYED SELECTOR SWITCH 2-POLE 2NO / 2NC | 133 | 133 | 133 |  | 2010-03-03 | 289237 / 10-P1267 | 6 | 87.15 | 522.9
11 | 800T-J91A | SWITCH 3-POSITION STC, 2NO, 2NC | 62 | 62 | 682 |  | 2010-03-03 | 289237 / 10-P1267 | 60 | 56.9 | 3414
1 | 700-ha33z24 | CONTROL RELAY 2PDT 24VDC 8-PIN | 13 | 13 | 13 |  | 2010-03-03 | 289237 / 10-P1267 | 6 | 12.37 | 74.22
1 | 700-HN126 | SOCKET FOR CONTROL RELAY | 7 | 7 | 7 |  | 2010-03-03 | 289237 / 10-P1267 | 6 | 5.76 | 34.56
2 | 1492-CB1G100 | CIRCUIT BREAKER 10 AMPS | 44 | 44 | 88 |  | 2010-03-03 | 289237 / 10-P1267 | 12 | 44.93 | 539.16
10 | 1492-CB1G050 | CIRCUIT BREAKER 5 AMPS | 44 | 44 | 440 |  | 2010-03-03 | 289237 / 10-P1267 | 6 | 44.93 | 269.58
10 | 1492-CB1G020 | CIRCUIT BREAKER 2 AMPS | 44 | 44 | 440 |  | 2010-03-03 | 289237 / 10-P1267 | 96 | 44.93 | 4313.28
10 | 199-DR1 | DIN MOUNTING RAIL | 6.3 | 6.3 | 63
10 | 885-2269 | BANANA JACK DOUBLE, BLACK | 2.4 | 2.4 | 24
1 | 90F7644 | RELAY TIMER | 84 | 84 | 84 |  | 2010-03-03 | 289237 / 10-P1267 | 6 | 173.91 | 1043.46
|  |  | Subtotal: |  | 4883.2 | 24416 |  |  |  |  | 17630.86
Connectors & hardware
30 | D2x4LKG6 | Wire Duct 2" x4" x 6ft length | 8 | 8 | 240
30 | C2LG6 | 2"x 4" Wire Duct cover 6ft length | 2 | 2 | 60
30 | D1.5x4LG6 | Wirer Duct 1.5'x4'x6ft length | 7.5 | 7.5 | 225 |  |  | Estimate of Various reqs for hardware |  |  | 4000
30 | C1.5LG6 | 1.5"x4" wire Duct cover 6ft length | 1.5 | 1.5 | 45
20 | 534-1301 | TERMINATION SUPPLIES, FERRULES, ETC | 50 | 50 | 1000
|  |  | Subtotal: |  | 1570 | 7850
Cables
1 | ALPHA 1736C/R1000 | CABLE 2 COND 22AWG 500 ft | 200 | 200 | 200 |  | 2012-07-12 | 319742/12-P1962-0 |  |  | 64396
7 | ALPHA 5470/25C | CABLE 25 COND 20AWG 500 ft | 4560 | 4560 | 31920 |  | 2013-01-24 | 326569/13-M0462 |  |  | 230
2 | ALPHA 45470/20C | CABLE 20COND 20 AWG 500 ft | 3125 | 3125 | 6250
|  | Rotation cable |  | 0 | 0 |  | 2013-03-08 | 326872 / 13-0532 |  |  | 3421.86
|  |  |  | 0 | 0 |  | 2010-07-13 | 293267 / 10-M2704-0 |  |  | 1145
|  |  |  | 0 | 0
|  |  |  | 0 | 0
|  |  |  | 0 | 0
|  |  | Subtotal: |  | 38370 | 191850 |  | Cables Ordered | 5 magnets |  | 69192.86
|  |  |  | Total | 48158.72 | 240793.6 |  |  |  | total | 90823.72 | -62.28%
|  |  |  | less cables | 9788.72
Note on D32: :
Note on D42: POMONA 2269-0
Note on D43: CNT-35-96 TYCO ELECTRONICS
Note on D52: 22AWG FERRULES
Note on A57: FEET

## PRs
SHMS Electrical Components |  | 12CCMC | 12GEV | 43-014 & 43-053
PR | Sent Date | PR Estimate | PO /Buyer | Vendor | PO Cost | Description | Delivery Complete | Location
378971 | 2018-08-08 | 55898 | 18-P1254 | EECO | 55755.39 | PLC and I/O modules | 2018-09-26 | TED 1544
378847 | 2018-07-31 | 33030 | 18-01188 | EECO | 33030 | Software | 2018-08-24 | TED 1544
382342 | 2019-01-15 | 172 | 19-M0362 | EECO | 172 | Cable
395689 | 2020-07-13 | 18879 |  | Dataforth |  | Signal Conditioners
395691 | 2020-07-13 | 3750 | 20-M1085 | Cryomagnetics | 3430 | Liquid Level Control and probes LM-501 | 2020-09-22
|  | 111729 |  |  | 92387.39 | Total

## PLC IO
PLC 2018 cost |  | Cost
PLC | [5 empty] | Magnet I/O Rack
[5 empty] | Qty |  | Unit Cost | Total Cost | PO#
Software: | Qty | Unit Cost | Total Cost | PO# |  | PLC Hardware:
RSLogix 5000 Pro edition | 0 | 5500 | 0 |  | 2 | Taps | 71 | 142
RSView Server unlimited | 1 | 16700 | 16700 | 18-P1188 | 18 | Terminators | 5 | 90
RSView Client | 4 | 2780 | 11120 | 18-P1188 | 3 | Screw Clamp Blocks | 45 | 135
RSview Studio for FT View | 1 | 2810 | 2810 | 18-P1188 | 1 | PSU-PS13 | 122 | 122
Kepserver Enterprise ESD | 1 | 2400 | 2400 | 18-P1188 | 1 | MVI-94 gsc | 1056 | 1056
TechConnect Support 1 yr | 0 | 2385 | 0 |  | 1 | FLEX ACN15 | 442 | 442
Subtotal |  |  | 33030 |  | 1 | A/B Breakers | 1043 | 1043
[5 empty] | 1 | PSU 5V | 87 | 87
PLC Controller Hardware: |  |  |  |  | 1 | PSU 24V | 280 | 280
Server Windows based | 0 | 2875 | 0 |  | 3 | CNB/D | 866 | 2598
Client PC | 0 | 2000 | 0 |  |  | Subtotal |  | 5995
4-slot chasis | 2 | 290 | 580 | 18-P1254
Redundant PLC AB1756182E | 2 | 8004 | 16008 | 18-P1254 |  | Hardware and Testing:
CNB/D | 5 | 1720 | 8600 | 18-P1254 | 0 | Omega Signal Conditioners | 1470 | 0
System Redundancy Module | 2 | 5000 | 10000 | 18-P1254 | 1 | Dataforth Signal Conditioners | 10986 | 10986
10-slot Chasis | 3 | 545 | 1635 | 18-P1254 | 1 | Macro_sensors LVDT | 3840 | 3840
PSU 5V | 5 | 975 | 4875 | 18-P1254 | 1 | Relay Board JT | 448 | 448
OF4 | 1 | 1070 | 1070 | 18-P1254 | 1 | Rack Hardware | 2801 | 2801
OW16I | 3 | 550 | 1650 | 18-P1254 | 1 | Quench Detector | 1250 | 1250
IF16 | 8 | 1355 | 10840 | 18-P1254 | 1 | Current Boards | 263 | 263
IV16 | 2 | 320 | 640 | 18-P1254 | 1 | Dataforth cable | 300 | 300
Redundancy cable | 1 | 172 | 172 | Credit Card | 1 | Rack Hardware | 4287 | 4287
Subtotal |  |  | 56070
[6 empty] | Subtotal |  | 24175
Total |  |  | 131565
Cables:
[5 empty] | 1 | Belden 23524 12twpr 18awg 500ft | 1000 | 1000
[5 empty] | 1 | Belden 1058A 12twpr 22awg 500ft | 580 | 580
[5 empty] | 5 | Belden 1060A 24twpr 22awg 500ft | 1575 | 7875
[5 empty] | 2 | Belden 9949 37cond 18awg 500ft | 1420 | 2840
[6 empty] | Subtotal |  | 12295

## Instrumentation
Remaining Instrumentation for CLEO_II Magnet
| 2007 prices and part number | 2007 prices | 2020 prices estimate
qty | part # | unit cost | total | Actual part #
2 | Pressure Gauge (Helium and LN2) | 565 | 1500 | OmegaDyne PX35D1-100GV
2 | Vaccum Gauge | 230 | 500 | WRG-S-DN40CF Wide Range Vacuum Gauge
1 | L He Meter | 3356 | 4500 | Model LM-500 with model 3DA sensor + cables
1 | L N2 Meter | 2461 | 1500 | Model LM-500 with model LN2 sensor + cables
2 | Mass Flow controller | 3250 | 7000 | MCR-500SLPM-D_remote/5M,5IN,Gas:He
2 | Relay Board | 0 | 500 | ICC-ER-B24
4 | Constant Current Source Board | 0 | 500 | CMA1D03A
1 | UPS power protection | 2500 | 6000
22000
Hardware inhouse
Parts at JLAB Spares or surplus items
| JT actuactors with LVDT sensors |  | 0
| LVDT sensor modules |  | 0
| Vacuum pumps |  | 0
| Turbo Pump |  | 0
| Vacumm hoses |  | 0
| Power supply bus lines |  | 0
| Power supply LCW hoses |  | 0
| Ball Valve Actuator (controller) |  | 0 | CE5-C-TX-HZ-G
Hardware to buy (or borrow)
| Cryo interface connection |  | waiting on design
| U-Tubes |  | waiting on quote
| Rack Hardware (terminal blocks, relays, breakers, etc) |  | 15000
| I&C cables (magnet to control rack) |  | 10000
| UPS |  | 6000
| Vacuum gate valve |  | 5000
| Hardware to re-solder current leads to magnet |  | 5000
| Support stand for heat excahnger and working platform |  | 15000
| IOC hardware (remaining items) |  | 2000
58000
Note on C5: Honeywell Web page http://www.sensotec.com/pressure_transducer_instock.shtml 3/7/07
Note on C6: From PR 191738 8/23/2000
Note on C7: American Magnetics Model 135 Monitor and sensor probe
Note on C8: American Magnetics
Note on C9: Flow Controllers PR# 255827 Model #1559A Calibrated to He gas 4-20mA output

## JT POSP
JT # | Desc | EUMax | EUMin | Cycle time | Open Rate | Close Rate | Max On time | Min on time
1 | He Return | 100 | -6 | 1 | 100 | 100 | 0.5 | 0.0009
2 | Cool Down | 100 | -6 | 1 | 100 | 100 | 0.5 | 0.001
3 | He Supply Top | 100 | -6 | 1 | 200 | 200 | 0.1 | 0.0005
4 | LN2 Suppply Top | 100 | -6 | 1 | 100 | 100 | 0.2 | 0.001
5 | He Supply Bottom | 100 | -6 | 1 | 100 | 100 | 0.5 | 0.0008
6 | LN2 Bottom | 100 | -6 | 1 | 100 | 100 | 0.1 | 0.001
7 | CL Pot | 100 | -6 | 1 | 100 | 100 | 0.5 | 0.001
WR | Warm return | 100 | 0 | 1 | 200 | 200 | 0.5 | 0.01
8 | HX LN2 | 100 | -6 | 3 | 200 | 200 | 0.1 | 0.005
9 | HX Ghe | 100 | -6 | 1 | 100 | 100 | 0.1 | 0.009