## Slide 1: LTCC Leakage
Amanda Hoebel

## Slide 2: Calculating Leak Rate from Pressure Decay
Leak causes sector’s internal pressure to decrease when there is no flow.
Pressure drop can be used to estimate leak rate.
Leak rate formula used:

## Slide 3: Python Script Parameters
Script looks for times where gas flow into a sector stops and then when flow starts again.
  Time interval between flow stop and start = “decay period”
Data is discarded if decay period is longer than 2.5 hours.
Data taken from 2017-04-25 17:04:16 to 2017-05-30 08:45:33
  From Marc’s cRIO data logging program.
  Time from 2017-05-09 11:00:00 to 2017-05-11 19:00:00 not used
    Time corresponds to dead band experimentation for solenoid valves.
Average ambient pressure for decay period used in calculation.
Volume assumed to be constant at 2700 L for each sector.

## Slide 4: Leak Rate from Pressure Decay Results
Tyler’s Calculations
Brian’s Calculations

## Slide 5: Fitting an Exponential Decay to Decay Periods
Pressure during each decay period was fit with an exponential decay.
  Exponential decay formula:
Pressure was determined at beginning and end of each decay period using resulting fit equation.
Leak rate was calculated with same formula:

## Slide 6: Leak Rate for Pressure Decay Fit Results
Tyler’s Calculations
Brian’s Calculations

## Slide 7: Pressure Decay Calculation Conclusion
Average leak from data and from fit do not match leak rates as determined by Brian.
Potential causes of discrepancies:
  Volume of sector taken as constant.
    Volume of sector is effected by pressure.
  Ambient pressure varied during decay periods.
    Ideally for pressure drop calculation, pressure outside of vessel should be constant.
  “Good” sectors only had a few decay periods under 2.5 hours long compared to “bad” sectors.
    Sector 5: 16 of 81 decay periods under 150 minutes
    Sector 3: 180 of 249 decay periods under 150 minutes

## Slide 8: Calculating Leak Rate from Flow Rate
Leak rate estimated using flow.
Data taken from 2017-04-25 17:04:16 to 2017-05-30 08:45:33
  From Marc’s cRIO data logging program.
  Time from 2017-05-10 to 2017-05-11 not used.

## Slide 9: Python Program
Hours
Fill Frequency Histogram for S3
Program looks for times where gas starts flowing into a sector, and then when gas stops.
Histograms made for all 6 sectors.
Sector 3
  Flow time: 31.34 minutes
  Number of flows: 7.82 times a day
  Loss: 122.56 L/day

## Slide 10: Results of Leakage calculated using Flow
Amanda’s Flow Calculations
Brian’s Flow Calculations

## Slide 11: Leakage Comparison
Chart: LTCC Leakage Comparison (Flow)
| 1 | 2 | 3 | 4 | 5 | 6
Brian | 112.5368 | 51.5899 | 125.7957 | 54.0051 | 43.5183 | 139.8941
Amanda | 119.0111 | 59.0872 | 134.4467 | 61.3283 | 53.2583 | 142.3367

## Slide 12: Leak Flow Rate Over Ambient Pressure

## Slide 13: Leakage Conclusion
Average leak from flow matches leak rates determined by Brian.
Average leak from pressure does not match rates determined by flow.
All calculations determine sectors 1, 3, and 6 as leakiest sectors.

## Slide 14: Thank You