## LV Chassis Layout
2016-10-04
Chassis Layout -- view is looking at the back of the chassis
CHASSIS 1
| TR |  | PT |  | SG |  | LC |  | TR | PT | SG | LC/HS
| J1 | J5 | J9 | J13 | J17 | J21 | J25 | SPARE CONNECTOR | 0 | 2 | 4 | 0
| J2 | J6 | J10 | J14 | J18 | J22 | J26 | SPARE CHANNELS | 0 | 4 | 8 | 0
| J3 | J7 | J11 | J15 | J19 | J23 | J27
| J4 | J8 | J12 | J16 | J20 | J24 | J28
CHASSIS 2
| TR |  | PT |  | SG |  | LC
| J1 | J5 | J9 | J13 | J17 | J21 | J25 |  | TR | PT | SG | LC/HS
| J2 | J6 | J10 | J14 | J18 | J22 | J26 | SPARE CONNECTOR | 0 | 5 | 1 | 1
| J3 | J7 | J11 | J15 | J19 | J23 | J27 | SPARE CHANNELS | 0 | 10 | 2 | 2
| J4 | J8 | J12 | J16 | J20 | J24 | J28
CHASSIS 3
| TR |  | PT |  | SG |  | LC
| J1 | J5 | J9 | J13 | J17 | J21 | J25 |  | TR | PT | SG | LC/HS
| J2 | J6 | J10 | J14 | J18 | J22 | J26 | SPARE CONNECTOR | 1 | 5 | 3 | 1
| J3 | J7 | J11 | J15 | J19 | J23 | J27 | SPARE CHANNELS | 2 | 10 | 6 | 2
| J4 | J8 | J12 | J16 | J20 | J24 | J28
CHASSIS 4
| TR |  | PT |  | SG |  | LC
| J1 | J5 | J9 | J13 | J17 | J21 | J25 |  | TR | PT | SG | LC/HS
| J2 | J6 | J10 | J14 | J18 | J22 | J26 | SPARE CONNECTOR | 2 | 5 | 3 | 1
| J3 | J7 | J11 | J15 | J19 | J23 | J27 | SPARE CHANNELS | 4 | 10 | 6 | 2
| J4 | J8 | J12 | J16 | J20 | J24 | J28
CHASSIS 5
| TR |  | PT |  | SG |  | LC
| J1 | J5 | J9 | J13 | J17 | J21 | J25 |  | TR | PT | SG | LC/HS
| J2 | J6 | J10 | J14 | J18 | J22 | J26 | SPARE CONNECTOR | 1 | 5 | 2 | 1
| J3 | J7 | J11 | J15 | J19 | J23 | J27 | SPARE CHANNELS | 2 | 10 | 4 | 2
| J4 | J8 | J12 | J16 | J20 | J24 | J28
CHASSIS 6
| TR |  | PT |  | SG |  | LC
| J1 | J5 | J9 | J13 | J17 | J21 | J25 |  | TR | PT | SG | LC/HS
| J2 | J6 | J10 | J14 | J18 | J22 | J26 | SPARE CONNECTOR | 1 | 5 | 4 | 1
| J3 | J7 | J11 | J15 | J19 | J23 | J27 | SPARE CHANNELS | 2 | 10 | 8 | 2
| J4 | J8 | J12 | J16 | J20 | J24 | J28 |  | 10 | 50 | 26 | 10
Each connector has 2 channels (A, B) - cryocon pinout
LC channel can also be HS channel. Excitation is the same
Note on J6: Probir Ghoshal: No more spare now as on 29th Sept 2016 after moving the cernoxes from cryocon readout

## Voltage Inputs
Version: | 1
By: | Joshua Ballard
Date: | 2015-03-17
| Dec1st 2016 | PkG
| 2016-12-13 | SLS/PKG
NI
Signal Name | Type | Input Range (V) |  | Actual Range (V) |  | Module # | Channel # | EPICS PV | PLC Tag (DINT) | PLC Tag Type | PLC Units
|  | Min | Max | Min | Max
VT1 | Volts | -10 | 10 | -10 | 10 | 1 | 0 | hallb-crio-tor-fast:IOC:VT1 | VT_Array[0] | DINT | uV | voltage tap (VTTs)
VT2 | Volts | -10 | 10 | -10 | 10 | 1 | 1 | hallb-crio-tor-fast:IOC:VT2 | VT_Array[1] | DINT | uV | voltage tap
VT3 | Volts | -10 | 10 | -10 | 10 | 1 | 2 | hallb-crio-tor-fast:IOC:VT3 | VT_Array[2] | DINT | uV | voltage tap
VT4 | Volts | -10 | 10 | -10 | 10 | 1 | 3 | hallb-crio-tor-fast:IOC:VT4 | VT_Array[3] | DINT | uV | voltage tap
VT5 | Volts | -10 | 10 | -10 | 10 | 2 | 0 | hallb-crio-tor-fast:IOC:VT5 | VT_Array[4] | DINT | uV | voltage tap
VT6 | Volts | -10 | 10 | -100 | 100 | 2 | 1 | hallb-crio-tor-fast:IOC:VT6 | VT_Array[5] | DINT | uV | voltage tap
VT7 | Volts | -10 | 10 | -10 | 10 | 2 | 2 | hallb-crio-tor-fast:IOC:VT7 | VT_Array[6] | DINT | uV | voltage tap
VT8 | Volts | -10 | 10 | -100 | 100 | 2 | 3 | hallb-crio-tor-fast:IOC:VT8 | VT_Array[7] | DINT | uV | voltage tap
VT9 | Volts | -10 | 10 | -10 | 10 | 3 | 0 | hallb-crio-tor-fast:IOC:VT9 | VT_Array[8] | DINT | uV | voltage tap
VT10 | Volts | -10 | 10 | -100 | 100 | 3 | 1 | hallb-crio-tor-fast:IOC:VT10 | VT_Array[9] | DINT | uV | voltage tap
VT11 | Volts | -10 | 10 | -10 | 10 | 3 | 2 | hallb-crio-tor-fast:IOC:VT11 | VT_Array[10] | DINT | uV | voltage tap
VT12 | Volts | -10 | 10 | -100 | 100 | 4 | 1 | hallb-crio-tor-fast:IOC:VT12 | VT_Array[11] | DINT | uV | voltage tap | channeling is skewed….JTB See B00000-09-00-0183
VT13 | Volts | -10 | 10 | -10 | 10 | 4 | 2 | hallb-crio-tor-fast:IOC:VT13 | VT_Array[12] | DINT | uV | voltage tap
VT14 | Volts | -10 | 10 | -100 | 100 | 4 | 3 | hallb-crio-tor-fast:IOC:VT14 | VT_Array[13] | DINT | uV | voltage tap
VT15 | Volts | -10 | 10 | -10 | 10 | 5 | 0 | hallb-crio-tor-fast:IOC:VT15 | VT_Array[14] | DINT | uV | voltage tap
VT16 | Volts | -10 | 10 | -100 | 100 | 5 | 1 | hallb-crio-tor-fast:IOC:VT16 | VT_Array[15] | DINT | uV | voltage tap
VT17 | Volts | -10 | 10 | -10 | 10 | 5 | 2 | hallb-crio-tor-fast:IOC:VT17 | VT_Array[16] | DINT | uV | voltage tap
VT18 | Volts | -10 | 10 | -10 | 10 | 5 | 3 | hallb-crio-tor-fast:IOC:VT18 | VT_Array[17] | DINT | uV | voltage tap
VT19 | Volts | -10 | 10 | -10 | 10 | 6 | 0 | hallb-crio-tor-fast:IOC:VT19 | VT_Array[18] | DINT | uV | voltage tap
VT20 | Volts | -10 | 10 | -10 | 10 | 6 | 1 | hallb-crio-tor-fast:IOC:VT20 | VT_Array[19] | DINT | uV | voltage tap
VT21 | Volts | -10 | 10 | -10 | 10 | 6 | 2 | hallb-crio-tor-fast:IOC:VT21 | VT_Array[20] | DINT | uV | voltage tap
VT22 | Volts | -10 | 10 | -600 | 600 | 6 | 3 | hallb-crio-tor-fast:IOC:VT22 | VT_Array[22] | DINT | uV | voltage tap
VT23 | Volts | -10 | 10 | -10 | 10 | 3 | 3 | hallb-crio-tor-fast:IOC:VT23 | VT_Array[23] | DINT | uV | voltage tap | added on 4th Oct 2016 across splice S7
VT24 | Volts | -10 | 10 | -10 | 10 | 7 | 0 | hallb-crio-tor-fast:IOC:VT24 | VT_Array[24] | DINT | uV | voltage tap | Added across ONE LCW flexible bus lead B (towards the vaporizer)
* | * | -10 | 10 | -10 | 10 | 7 | 1 | hallb-crio-tor-fast:IOC | VT_Array[24] | DINT | uV | voltage tap | Parallel split from the MPS BNC for fDAQ
G1 | Volts | 0 | 1 | 0 | 1 | 8 | 0 |  | GF_Array[0] | DINT | uV | ground fault GF)
G2 | Volts | 0 | 1 | 0 | 1 | 8 | 1 |  | GF_Array[1] | DINT | uV | ground fault GF)
G3 | Volts | 0 | 1 | 0 | 1 | 8 | 2 |  | GF_Array[2] | DINT | uV | ground fault GF)
G3 | Volts | 0 | 1 | 0 | 1 | 8 | 3 |  | GF_Array[2] | DINT | uV | ground fault GF)
V_Dump | Volts | -10 | 10 | -500 | 500 | 7 | 0 |  | PS_Array[0] | DINT | uV | Status voltage from power supply
ZFCT1 | Volts | 0 | 10 | 0 | 4000 | 7 | 1 | hallb-crio-tor-fast:IOC:IDCCT1 | PS_Array[1] | DINT | uV | Current Monitor output
ZFCT2 | Volts | 0 | 1 | 0 | 1 | 7 | 2 |  | PS_Array[2] | DINT | uV
V_PS | Volts | 0 | 6 | 0 | 6 | 7 | 3 |  | PS_Array[3] | DINT | uV

## RS232 Inputs
2016-10-03
| 2016-11-10 | PKG |  | PLC Arrays | [15 empty] | CHANGE THIS DATA TO ADJUST THE DEFAULT VALUES FOR THE SENSOR TYPES!
| 2016-11-15 | RB/PKG |  | Name | Size* | *element 0 of each array is a counter -- cRIO to increment by 1 each time the array is updated -- roll over to 0 at 2^31 | [14 empty] | Excitation | AC/DC | Excitation | Calibration | PLC Units
| 2016-12-13 | SLS/PKG |  | LC_From_cRIO | 49 | [14 empty] | TR | I | AC | .3-20 | LOOKUP | mK
|  |  |  | SG_From_cRIO | 97 | [14 empty] | SG | V | DC | 2.5 | FORMULA | nV/V
|  |  |  | PT_From_cRIO | 97 | [14 empty] | PT | I | DC | 2500 | FORMULA | mK
|  |  |  | TR_From_cRIO | 97 | [14 empty] | LC | V | DC | 2.5 | LOOKUP | lb
|  |  |  | HS_From_cRIO | 7 | [14 empty] | HS | I | DC | 100000 | LOOKUP | mG
[20 empty] | ELSE | ????? | ????? | ????? | ????? | ?????
Dwg. # | PV Name | Chassis | Connector | Channel | Type (TR, SG,…) | Serial # | Nomenclature | Excitation | AC/DC | Excitation (uA) | Excitation (V) | Calibrated? | Read Register Address | Write Register Address | Calib. Type | PLC Tag (DINT) | PLC Tag Type | PLC Units | Notes
-0162-3 | HS817A1 | 3 | J27 | A | HS | 056294 | CCM A, Front | I | DC | 100000 |  |  | 6c |  | LOOKUP | HS_From_cRIO[3] | DINT | mG | excitation is provided from an external source
-0162-4 | HS817B1 | 4 | J27 | A | HS | 057740 | CCM B, Front | I | DC | 100000 |  |  | 6c |  | LOOKUP | HS_From_cRIO[4] | DINT | mG | excitation is provided from an external source
-0162-5 | HS817C1 | 5 | J27 | A | HS | 058798 | CCM C, Front | I | DC | 100000 |  |  | 6c |  | LOOKUP | HS_From_cRIO[5] | DINT | mG | excitation is provided from an external source
-0162-6 | HS817D1 | 6 | J27 | A | HS | 057894B | CCM D, Front | I | DC | 100000 |  |  | 6c |  | LOOKUP | HS_From_cRIO[6] | DINT | mG | excitation is provided from an external source
-0162-1 | HS817E1 | 1 | J27 | A | HS | 058771 | CCM E, Front | I | DC | 100000 |  |  | 6c |  | LOOKUP | HS_From_cRIO[1] | DINT | mG | excitation is provided from an external source
-0162-2 | HS817F1 | 2 | J27 | A | HS | 058789 | CCM F, Front | I | DC | 100000 |  |  | 6c |  | LOOKUP | HS_From_cRIO[2] | DINT | mG | excitation is provided from an external source
-0162-3 | LC817A1 | 3 | J25 | A | LC | 588903 | CCM A, Front Upstream | V | DC |  | 2.5 |  | 68 | 30 | LOOKUP | LC_From_cRIO[1] | DINT | lb | Voltage
-0162-3 | LC817A2 | 3 | J25 | B | LC | 611619 | CCM A, Front Downstream | V | DC |  | 2.5 |  | 69 | 31 | LOOKUP | LC_From_cRIO[2] | DINT | lb | Voltage
-0162-3 | LC817A3 | 3 | J26 | A | LC | 588905 | CCM A, Back Upstream | V | DC |  | 2.5 |  | 6a | 32 | LOOKUP | LC_From_cRIO[3] | DINT | lb | Voltage
-0162-3 | LC817A4 | 3 | J26 | B | LC | 588891 | CCM A, Back Downstream | V | DC |  | 2.5 |  | 6b | 33 | LOOKUP | LC_From_cRIO[4] | DINT | lb | Voltage
-0162-4 | LC817B1 | 4 | J25 | A | LC | 611618 | CCM B, Front Upstream | V | DC |  | 2.5 |  | 68 | 30 | LOOKUP | LC_From_cRIO[5] | DINT | lb | Voltage
-0162-4 | LC817B2 | 4 | J25 | B | LC | 588876 | CCM B, Front Downstream | V | DC |  | 2.5 |  | 69 | 31 | LOOKUP | LC_From_cRIO[6] | DINT | lb | Voltage
-0162-4 | LC817B3 | 4 | J26 | A | LC | 588893 | CCM B, Back Upstream | V | DC |  | 2.5 |  | 6a | 32 | LOOKUP | LC_From_cRIO[7] | DINT | lb | Voltage
-0162-4 | LC817B4 | 4 | J26 | B | LC | 588874 | CCM B, Back Downstream | V | DC |  | 2.5 |  | 6b | 33 | LOOKUP | LC_From_cRIO[8] | DINT | lb | Voltage
-0162-5 | LC817C1 | 5 | J25 | A | LC | 588878 | CCM C, Front Upstream | V | DC |  | 2.5 |  | 68 | 30 | LOOKUP | LC_From_cRIO[9] | DINT | lb | Voltage
-0162-5 | LC817C2 | 5 | J25 | B | LC | 588889 | CCM C, Front Downstream | V | DC |  | 2.5 |  | 69 | 31 | LOOKUP | LC_From_cRIO[10] | DINT | lb | Voltage
-0162-5 | LC817C3 | 5 | J26 | A | LC | 588879 | CCM C, Back Upstream | V | DC |  | 2.5 |  | 6a | 32 | LOOKUP | LC_From_cRIO[11] | DINT | lb | Voltage
-0162-5 | LC817C4 | 5 | J26 | B | LC | 588897 | CCM C, Back Downstream | V | DC |  | 2.5 |  | 6b | 33 | LOOKUP | LC_From_cRIO[12] | DINT | lb | Voltage
-0162-6 | LC817D1 | 6 | J25 | A | LC | 588873 | CCM D, Front Upstream | V | DC |  | 2.5 |  | 68 | 30 | LOOKUP | LC_From_cRIO[13] | DINT | lb | Voltage
-0162-6 | LC817D2 | 6 | J25 | B | LC | 588896 | CCM D, Front Downstream | V | DC |  | 2.5 |  | 69 | 31 | LOOKUP | LC_From_cRIO[14] | DINT | lb | Voltage
-0162-6 | LC817D3 | 6 | J26 | A | LC | 588881 | CCM D, Back Upstream | V | DC |  | 2.5 |  | 6a | 32 | LOOKUP | LC_From_cRIO[15] | DINT | lb | Voltage
-0162-6 | LC817D4 | 6 | J26 | B | LC | 611621 | CCM D, Back Downstream | V | DC |  | 2.5 |  | 6b | 33 | LOOKUP | LC_From_cRIO[16] | DINT | lb | Voltage
-0162-1 | LC817E1 | 1 | J25 | A | LC | 588880 | CCM E, Front Upstream | V | DC |  | 2.5 |  | 68 | 30 | LOOKUP | LC_From_cRIO[17] | DINT | lb | Voltage
-0162-1 | LC817E2 | 1 | J25 | B | LC | 611620 | CCM E, Front Downstream | V | DC |  | 2.5 |  | 69 | 31 | LOOKUP | LC_From_cRIO[18] | DINT | lb | Voltage
-0162-1 | LC817E3 | 1 | J26 | A | LC | 588887 | CCM E, Back Upstream | V | DC |  | 2.5 |  | 6a | 32 | LOOKUP | LC_From_cRIO[19] | DINT | lb | Voltage
-0162-1 | LC817E4 | 1 | J26 | B | LC | 588877 | CCM E, Back Downstream | V | DC |  | 2.5 |  | 6b | 33 | LOOKUP | LC_From_cRIO[20] | DINT | lb | Voltage
-0162-6 | LC817H1 | 1 | J28 | A | LC | 588884 | Hub Lateral Support Sector 1 | V | DC |  | 2.5 |  | 6e | 36 | LOOKUP | LC_From_cRIO[25] | DINT | lb
-0162-6 | LC817H4 | 1 | J28 | B | LC | 542981 | Hub Lateral Support Sector 4 | V | DC |  | 2.5 |  | 6f | 37 | LOOKUP | LC_From_cRIO[26] | DINT | lb
-0162-2 | LC817F1 | 2 | J25 | A | LC | 588899 | CCM F, Front Upstream | V | DC |  | 2.5 |  | 68 | 30 | LOOKUP | LC_From_cRIO[21] | DINT | lb | Voltage
-0162-2 | LC817F2 | 2 | J25 | B | LC | 588890 | CCM F, Front Downstream | V | DC |  | 2.5 |  | 69 | 31 | LOOKUP | LC_From_cRIO[22] | DINT | lb | Voltage
-0162-2 | LC817F3 | 2 | J26 | A | LC | 588875 | CCM F, Back Upstream | V | DC |  | 2.5 |  | 6a | 32 | LOOKUP | LC_From_cRIO[23] | DINT | lb | Voltage
-0162-2 | LC817F4 | 2 | J26 | B | LC | 588883 | CCM F, Back Downstream | V | DC |  | 2.5 |  | 6b | 33 | LOOKUP | LC_From_cRIO[24] | DINT | lb | Voltage
-0165-1 | PT817A5 | 3 | J9 | A, B | PT |  | CCM A, Shield Back 1 | I | DC | 2500 |  |  | 48,49 | 10,11 | FORMULA | PT_From_cRIO[5] | DINT | mK
-0165-2 | PT817A1 | 3 | J10 | A | PT |  | CCM A, Shield Front 1 | I | DC | 2500 |  |  | 4a | 12 | FORMULA | PT_From_cRIO[1] | DINT | mK
-0165-2 | PT817A2 | 3 | J10 | B | PT |  | CCM A, Shield Front 2 | I | DC | 2500 |  |  | 4b | 13 | FORMULA | PT_From_cRIO[2] | DINT | mK
-0165-2 | PT817A3 | 3 | J11 | A | PT |  | CCM A, Shield Front 3 | I | DC | 2500 |  |  | 4c | 14 | FORMULA | PT_From_cRIO[3] | DINT | mK
-0165-2 | PT817A4 | 3 | J11 | B | PT |  | CCM A, Shield Front 4 | I | DC | 2500 |  |  | 4d | 15 | FORMULA | PT_From_cRIO[4] | DINT | mK
-0166-1 | PT817B5 | 4 | J9 | A, B | PT | 199, 200(r) | CCM B, Shield Back 1 | I | DC | 2500 |  |  | 48,49 | 10,11 | FORMULA | PT_From_cRIO[10] | DINT | mK
-0166-2 | PT817B1 | 4 | J10 | A | PT | 201, 202(r) | CCM B, Shield Front 1 | I | DC | 2500 |  |  | 4a | 12 | FORMULA | PT_From_cRIO[6] | DINT | mK
-0166-2 | PT817B2 | 4 | J10 | B | PT | 181, 182(r) | CCM B, Shield Front 2 | I | DC | 2500 |  |  | 4b | 13 | FORMULA | PT_From_cRIO[7] | DINT | mK
-0166-2 | PT817B3 | 4 | J11 | A | PT | 203, 204(r) | CCM B, Shield Front 3 | I | DC | 2500 |  |  | 4c | 14 | FORMULA | PT_From_cRIO[8] | DINT | mK
-0166-2 | PT817B4 | 4 | J11 | B | PT | 177, 178(r) | CCM B, Shield Front 4 | I | DC | 2500 |  |  | 4d | 15 | FORMULA | PT_From_cRIO[9] | DINT | mK
-0167-1 | PT817C5 | 5 | J9 | A, B | PT | 155, 156(r) | CCM C, Shield Back 1 | I | DC | 2500 |  |  | 48,49 | 10,11 | FORMULA | PT_From_cRIO[15] | DINT | mK
-0167-2 | PT817C1 | 5 | J10 | A | PT | 127, 128(r) | CCM C, Shield Front 1 | I | DC | 2500 |  |  | 4a | 12 | FORMULA | PT_From_cRIO[11] | DINT | mK
-0167-2 | PT817C2 | 5 | J10 | B | PT | 157, 158(r) | CCM C, Shield Front 2 | I | DC | 2500 |  |  | 4b | 13 | FORMULA | PT_From_cRIO[12] | DINT | mK
-0167-2 | PT817C3 | 5 | J11 | A | PT | 105, 106(r) | CCM C, Shield Front 3 | I | DC | 2500 |  |  | 4c | 14 | FORMULA | PT_From_cRIO[13] | DINT | mK
-0167-2 | PT817C4 | 5 | J11 | B | PT | 107, 108(r) | CCM C, Shield Front 4 | I | DC | 2500 |  |  | 4d | 15 | FORMULA | PT_From_cRIO[14] | DINT | mK
-0168-1 | PT817D5 | 6 | J9 | A, B | PT | 167, 168(r) | CCM D, Shield Back 1 | I | DC | 2500 |  |  | 48,49 | 10,11 | FORMULA | PT_From_cRIO[20] | DINT | mK
-0168-2 | PT817D1 | 6 | J10 | A | PT | 117, 118(r) | CCM D, Shield Front 1 | I | DC | 2500 |  |  | 4a | 12 | FORMULA | PT_From_cRIO[16] | DINT | mK
-0168-2 | PT817D2 | 6 | J10 | B | PT | 171, 172(r) | CCM D, Shield Front 2 | I | DC | 2500 |  |  | 4b | 13 | FORMULA | PT_From_cRIO[17] | DINT | mK
-0168-2 | PT817D3 | 6 | J11 | A | PT | 163, 164(r) | CCM D, Shield Front 3 | I | DC | 2500 |  |  | 4c | 14 | FORMULA | PT_From_cRIO[18] | DINT | mK
-0168-2 | PT817D4 | 6 | J11 | B | PT | 139, 140(r) | CCM D, Shield Front 4 | I | DC | 2500 |  |  | 4d | 15 | FORMULA | PT_From_cRIO[19] | DINT | mK
-0169-1 | PT817E5 | 1 | J9 | A, B | PT |  | CCM E, Shield Back 1 | I | DC | 2500 |  |  | 48,49 | 10,11 | FORMULA | PT_From_cRIO[25] | DINT | mK
-0169-2 | PT817E1 | 1 | J10 | A | PT | 189, 190(r) | CCM E, Shield Front 1 | I | DC | 2500 |  |  | 4a | 12 | FORMULA | PT_From_cRIO[21] | DINT | mK
-0169-2 | PT817E2 | 1 | J10 | B | PT | 147, 136(r) | CCM E, Shield Front 2 | I | DC | 2500 |  |  | 4b | 13 | FORMULA | PT_From_cRIO[22] | DINT | mK
-0169-2 | PT817E3 | 1 | J11 | A | PT | 135, 136(r) | CCM E, Shield Front 3 | I | DC | 2500 |  |  | 4c | 14 | FORMULA | PT_From_cRIO[23] | DINT | mK
-0169-2 | PT817E4 | 1 | J11 | B | PT | 153, 154(r) | CCM E, Shield Front 4 | I | DC | 2500 |  |  | 4d | 15 | FORMULA | PT_From_cRIO[24] | DINT | mK
-0178-10 | TP8114R | 1 | J13 | A | PT | 229, 230(r) | Safety valve vent_RV8114R | I | DC | 2500 |  |  | 50 | 18 | FORMULA | PT_From_cRIO[31] | DINT | mK | New sensors added in order to monitor the vent temp and closing of SV after SV operation (during commisioning of torus)
-0178-10 | TP8120 | 1 | J13 | B | PT | 159, 160(r) | Safety valve vent_RV8120 | I | DC | 2500 |  |  | 51 | 19 | FORMULA | PT_From_cRIO[32] | DINT | mK | New sensors added in order to monitor the vent temp and closing of SV after SV operation (during commisioning of torus)
-0178-10 | TP8112 | 1 | J14 | A | PT | 161, 162(r) | Safety valve vent_RV8112 | I | DC | 2500 |  |  | 52 | 1a | FORMULA | PT_From_cRIO[33] | DINT | mK | New sensors added in order to monitor the vent temp and closing of SV after SV operation (during commisioning of torus)
-0164 | TP811B | 1 | J14 | B | PT | 219, 220(r) | Safety valve vent_RV811B | I | DC | 2500 |  |  | 53 | 1b | FORMULA | PT_From_cRIO[34] | DINT | mK | New sensors added in order to monitor the vent temp and closing of SV after SV operation (during commisioning of torus)
-0164 | TP811D | 1 | J15 | A | PT | 231, 232(r) | Safety valve vent_RV811D | I | DC | 2500 |  |  | 54 | 1c | FORMULA | PT_From_cRIO[35] | DINT | mK | New sensors added in order to monitor the vent temp and closing of SV after SV operation (during commisioning of torus)
-0164 | TP811F | 1 | J15 | B | PT | 239, 240(r) | Safety valve vent_RV811F | I | DC | 2500 |  |  | 55 | 1d | FORMULA | PT_From_cRIO[36] | DINT | mK | New sensors added in order to monitor the vent temp and closing of SV after SV operation (during commisioning of torus)
-0170-1 | PT817F5 | 2 | J9 | A, B | PT | 173, 174(r) | CCM F, Shield Back 1 | I | DC | 2500 |  |  | 48,49 | 10,11 | FORMULA | PT_From_cRIO[30] | DINT | mK
-0170-2 | PT817F1 | 2 | J10 | A | PT | 121, 122(r) | CCM F, Shield Front 1 | I | DC | 2500 |  |  | 4a | 12 | FORMULA | PT_From_cRIO[26] | DINT | mK
-0170-2 | PT817F2 | 2 | J10 | B | PT | 133, 134(r) | CCM F, Shield Front 2 | I | DC | 2500 |  |  | 4b | 13 | FORMULA | PT_From_cRIO[27] | DINT | mK
-0170-2 | PT817F3 | 2 | J11 | A | PT | 185, 186(r) | CCM F, Shield Front 3 | I | DC | 2500 |  |  | 4c | 14 | FORMULA | PT_From_cRIO[28] | DINT | mK
-0170-2 | PT817F4 | 2 | J11 | B | PT | 187, 188(r) | CCM F, Shield Front 4 | I | DC | 2500 |  |  | 4d | 15 | FORMULA | PT_From_cRIO[29] | DINT | mK
-0165-1 | SG817A1 | 3 | J17 | A | SG |  | CCM A, Front 1 | V | DC |  | 2.5 |  | 58 | 20 | FORMULA | SG_From_cRIO[1] | DINT | nV/V
-0165-1 | SG817A2 | 3 | J17 | B | SG |  | CCM A, Front 2 | V | DC |  | 2.5 |  | 59 | 21 | FORMULA | SG_From_cRIO[2] | DINT | nV/V
-0165-1 | SG817A3 | 3 | J18 | A | SG |  | CCM A, Back 1 | V | DC |  | 2.5 |  | 5a | 22 | FORMULA | SG_From_cRIO[3] | DINT | nV/V
-0165-1 | SG817A4 | 3 | J18 | B | SG |  | CCM A, Back 2 | V | DC |  | 2.5 |  | 5b | 23 | FORMULA | SG_From_cRIO[4] | DINT | nV/V
-0174-1 | SG817U3HB1 | 3 | J19 | A | SG |  | S3 US, Cold Beam, Upstream Face | V | DC |  | 2.5 |  | 5c | 24 | FORMULA | SG_From_cRIO[35] | DINT | nV/V
-0174-1 | SG817U3HB2 | 3 | J19 | B | SG |  | S3 US, Cold Beam, Downstream Face | V | DC |  | 2.5 |  | 5d | 25 | FORMULA | SG_From_cRIO[36] | DINT | nV/V
-0177-1 | SG817D3HB1 | 3 | J20 | A | SG |  | S3 DS, Cold Beam, Upstream Face | V | DC |  | 2.5 |  | 5e | 26 | FORMULA | SG_From_cRIO[47] | DINT | nV/V
-0177-1 | SG817D3HB2 | 3 | J20 | B | SG |  | S3 DS, Cold Beam, Downstream Face | V | DC |  | 2.5 |  | 5f | 27 | FORMULA | SG_From_cRIO[48] | DINT | nV/V
-0177-1 | SG817D4HB1 | 3 | J21 | A | SG |  | S4 DS, Cold Beam, Upstream Face | V | DC |  | 2.5 |  | 60 | 28 | FORMULA | SG_From_cRIO[49] | DINT | nV/V
-0177-1 | SG817D4HB2 | 3 | J21 | B | SG |  | S4 DS, Cold Beam, Downstream Face | V | DC |  | 2.5 |  | 61 | 29 | FORMULA | SG_From_cRIO[50] | DINT | nV/V
-0166-1 | SG817B1 | 4 | J17 | A | SG |  | CCM B, Front 1 | V | DC |  | 2.5 |  | 58 | 20 | FORMULA | SG_From_cRIO[5] | DINT | nV/V
-0166-1 | SG817B2 | 4 | J17 | B | SG |  | CCM B, Front 2 | V | DC |  | 2.5 |  | 59 | 21 | FORMULA | SG_From_cRIO[6] | DINT | nV/V
-0166-1 | SG817B3 | 4 | J18 | A | SG |  | CCM B, Back 1 | V | DC |  | 2.5 |  | 5a | 22 | FORMULA | SG_From_cRIO[7] | DINT | nV/V
-0166-1 | SG817B4 | 4 | J18 | B | SG |  | CCM B, Back 2 | V | DC |  | 2.5 |  | 5b | 23 | FORMULA | SG_From_cRIO[8] | DINT | nV/V
-0171-1 | SG817B5 | 4 | J19 | A | SG |  | CCM B Axial Support, Front Side | V | DC |  | 2.5 |  | 5c | 24 | FORMULA | SG_From_cRIO[25] | DINT | nV/V | Front/Back corresponds with the CCM Front/Back…
-0171-1 | SG817B6 | 4 | J19 | B | SG |  | CCM B Axial Support, Back Side | V | DC |  | 2.5 |  | 5d | 25 | FORMULA | SG_From_cRIO[26] | DINT | nV/V | Front/Back corresponds with the CCM Front/Back…
-0172-2 | SG817U4HB3 | 4 | J20 | A | SG |  | S4 US, Y Support, Front Side | V | DC |  | 2.5 |  | 5e | 26 | FORMULA | SG_From_cRIO[57] | DINT | nV/V | Front/Back corresponds with the CCM Front/Back…
-0172-2 | SG817U4HB4 | 4 | J20 | B | SG |  | S4 US, Y Support, Back Side | V | DC |  | 2.5 |  | 5f | 27 | FORMULA | SG_From_cRIO[58] | DINT | nV/V | Front/Back corresponds with the CCM Front/Back…
-0172-2 | SG817D4HB3 | 4 | J21 | A | SG |  | S4 DS, Y Support, Front Side | V | DC |  | 2.5 |  | 60 | 28 | FORMULA | SG_From_cRIO[61] | DINT | nV/V | Front/Back corresponds with the CCM Front/Back…
-0172-2 | SG817D4HB4 | 4 | J21 | B | SG |  | S4 DS, Y Support, Back Side | V | DC |  | 2.5 |  | 61 | 29 | FORMULA | SG_From_cRIO[62] | DINT | nV/V | Front/Back corresponds with the CCM Front/Back…
-0167-1 | SG817C1 | 5 | J17 | A | SG |  | CCM C, Front 1 | V | DC |  | 2.5 |  | 58 | 20 | FORMULA | SG_From_cRIO[9] | DINT | nV/V
-0167-1 | SG817C2 | 5 | J17 | B | SG |  | CCM C, Front 2 | V | DC |  | 2.5 |  | 59 | 21 | FORMULA | SG_From_cRIO[10] | DINT | nV/V
-0167-1 | SG817C3 | 5 | J18 | A | SG |  | CCM C, Back 1 | V | DC |  | 2.5 |  | 5a | 22 | FORMULA | SG_From_cRIO[11] | DINT | nV/V
-0167-1 | SG817C4 | 5 | J18 | B | SG |  | CCM C, Back 2 | V | DC |  | 2.5 |  | 5b | 23 | FORMULA | SG_From_cRIO[12] | DINT | nV/V
-0175-1 | SG817U4HB1 | 5 | J19 | A | SG |  | S4 US, Cold Beam, Upstream Face | V | DC |  | 2.5 |  | 5c | 24 | FORMULA | SG_From_cRIO[37] | DINT | nV/V
-0175-1 | SG817U4HB2 | 5 | J19 | B | SG |  | S4 US, Cold Beam, Downstream Face | V | DC |  | 2.5 |  | 5d | 25 | FORMULA | SG_From_cRIO[38] | DINT | nV/V
-0175-1 | SG817U5HB1 | 5 | J20 | A | SG |  | S5 US, Cold Beam, Upstream Face | V | DC |  | 2.5 |  | 5e | 26 | FORMULA | SG_From_cRIO[39] | DINT | nV/V
-0175-1 | SG817U5HB2 | 5 | J20 | B | SG |  | S5 US, Cold Beam, Downstream Face | V | DC |  | 2.5 |  | 5f | 27 | FORMULA | SG_From_cRIO[40] | DINT | nV/V
-0177-2 | SG817D5HB1 | 5 | J21 | A | SG |  | S5 DS, Cold Beam, Upstream Face | V | DC |  | 2.5 |  | 60 | 28 | FORMULA | SG_From_cRIO[51] | DINT | nV/V
-0177-2 | SG817D5HB2 | 5 | J21 | B | SG |  | S5 DS, Cold Beam, Downstream Face | V | DC |  | 2.5 |  | 61 | 29 | FORMULA | SG_From_cRIO[52] | DINT | nV/V
-0177-2 | SG817D6HB1 | 5 | J22 | A | SG |  | S6 DS, Cold Beam, Upstream Face | V | DC |  | 2.5 |  | 62 | 2a | FORMULA | SG_From_cRIO[53] | DINT | nV/V
-0177-2 | SG817D6HB2 | 5 | J22 | B | SG |  | S6 DS, Cold Beam, Downstream Face | V | DC |  | 2.5 |  | 63 | 2b | FORMULA | SG_From_cRIO[54] | DINT | nV/V
-0168-1 | SG817D1 | 6 | J17 | A | SG |  | CCM D, Front 1 | V | DC |  | 2.5 |  | 58 | 20 | FORMULA | SG_From_cRIO[13] | DINT | nV/V
-0168-1 | SG817D2 | 6 | J17 | B | SG |  | CCM D, Front 2 | V | DC |  | 2.5 |  | 59 | 21 | FORMULA | SG_From_cRIO[14] | DINT | nV/V
-0168-1 | SG817D3 | 6 | J18 | A | SG |  | CCM D, Back 1 | V | DC |  | 2.5 |  | 5a | 22 | FORMULA | SG_From_cRIO[15] | DINT | nV/V
-0168-1 | SG817D4 | 6 | J18 | B | SG |  | CCM D, Back 2 | V | DC |  | 2.5 |  | 5b | 23 | FORMULA | SG_From_cRIO[16] | DINT | nV/V
-0171-1 | SG817D5 | 6 | J19 | A | SG |  | CCM D Axial Support, Front Side | V | DC |  | 2.5 |  | 5c | 24 | FORMULA | SG_From_cRIO[27] | DINT | nV/V | Front/Back corresponds with the CCM Front/Back…
-0171-1 | SG817D6 | 6 | J19 | B | SG |  | CCM D Axial Support, Back Side | V | DC |  | 2.5 |  | 5d | 25 | FORMULA | SG_From_cRIO[28] | DINT | nV/V | Front/Back corresponds with the CCM Front/Back…
-0176-1 | SG817U6HB1 | 6 | J20 | A | SG |  | S6 US, Cold Beam, Upstream Face | V | DC |  | 2.5 |  | 5e | 26 | FORMULA | SG_From_cRIO[41] | DINT | nV/V
-0176-1 | SG817U6HB2 | 6 | J20 | B | SG |  | S6 US, Cold Beam, Downstream Face | V | DC |  | 2.5 |  | 5f | 27 | FORMULA | SG_From_cRIO[42] | DINT | nV/V
-0169-1 | SG817E1 | 1 | J17 | A | SG |  | CCM E, Front 1 | V | DC |  | 2.5 |  | 58 | 20 | FORMULA | SG_From_cRIO[17] | DINT | nV/V
-0169-1 | SG817E2 | 1 | J17 | B | SG |  | CCM E, Front 2 | V | DC |  | 2.5 |  | 59 | 21 | FORMULA | SG_From_cRIO[18] | DINT | nV/V
-0169-1 | SG817E3 | 1 | J18 | A | SG |  | CCM E, Back 1 | V | DC |  | 2.5 |  | 5a | 22 | FORMULA | SG_From_cRIO[19] | DINT | nV/V
-0169-1 | SG817E4 | 1 | J18 | B | SG |  | CCM E, Back 2 | V | DC |  | 2.5 |  | 5b | 23 | FORMULA | SG_From_cRIO[20] | DINT | nV/V
-0172-1 | SG817U1HB3 | 1 | J19 | A | SG |  | S1 US, Y Support, Front Side | V | DC |  | 2.5 |  | 5c | 24 | FORMULA | SG_From_cRIO[55] | DINT | nV/V | Front/Back corresponds with the CCM Front/Back…
-0172-1 | SG817U1HB4 | 1 | J19 | B | SG |  | S1 US, Y Support, Back Side | V | DC |  | 2.5 |  | 5d | 25 | FORMULA | SG_From_cRIO[56] | DINT | nV/V | Front/Back corresponds with the CCM Front/Back…
-0172-1 | SG817D1HB3 | 1 | J20 | A | SG |  | S1 DS, Y Support, Front Side | V | DC |  | 2.5 |  | 5e | 26 | FORMULA | SG_From_cRIO[59] | DINT | nV/V | Front/Back corresponds with the CCM Front/Back…
-0172-1 | SG817D1HB4 | 1 | J20 | B | SG |  | S1 DS, Y Support, Back Side | V | DC |  | 2.5 |  | 5f | 27 | FORMULA | SG_From_cRIO[60] | DINT | nV/V | Front/Back corresponds with the CCM Front/Back…
-0170-1 | SG817F1 | 2 | J17 | A | SG |  | CCM F, Front 1 | V | DC |  | 2.5 |  | 58 | 20 | FORMULA | SG_From_cRIO[21] | DINT | nV/V
-0170-1 | SG817F2 | 2 | J17 | B | SG |  | CCM F, Front 2 | V | DC |  | 2.5 |  | 59 | 21 | FORMULA | SG_From_cRIO[22] | DINT | nV/V
-0170-1 | SG817F3 | 2 | J18 | A | SG |  | CCM F, Back 1 | V | DC |  | 2.5 |  | 5a | 22 | FORMULA | SG_From_cRIO[23] | DINT | nV/V
-0170-1 | SG817F4 | 2 | J18 | B | SG |  | CCM F, Back 2 | V | DC |  | 2.5 |  | 5b | 23 | FORMULA | SG_From_cRIO[24] | DINT | nV/V
-0171-2 | SG817F5 | 2 | J19 | A | SG |  | CCM F Axial Support, Front Side | V | DC |  | 2.5 |  | 5c | 24 | FORMULA | SG_From_cRIO[29] | DINT | nV/V | Front/Back corresponds with the CCM Front/Back…
-0171-2 | SG817F6 | 2 | J19 | B | SG |  | CCM F Axial Support, Back Side | V | DC |  | 2.5 |  | 5d | 25 | FORMULA | SG_From_cRIO[30] | DINT | nV/V | Front/Back corresponds with the CCM Front/Back…
-0173-1 | SG817U1HB1 | 2 | J20 | A | SG |  | S1 US, Cold Beam, Upstream Face | V | DC |  | 2.5 |  | 5e | 26 | FORMULA | SG_From_cRIO[31] | DINT | nV/V
-0173-1 | SG817U1HB2 | 2 | J20 | B | SG |  | S1 US, Cold Beam, Downstream Face | V | DC |  | 2.5 |  | 5f | 27 | FORMULA | SG_From_cRIO[32] | DINT | nV/V
-0173-1 | SG817U2HB1 | 2 | J21 | A | SG |  | S2 US, Cold Beam, Upstream Face | V | DC |  | 2.5 |  | 60 | 28 | FORMULA | SG_From_cRIO[33] | DINT | nV/V
-0173-1 | SG817U2HB2 | 2 | J21 | B | SG |  | S2 US, Cold Beam, Downstream Face | V | DC |  | 2.5 |  | 61 | 29 | FORMULA | SG_From_cRIO[34] | DINT | nV/V
-0177-3 | SG817D1HB1 | 2 | J22 | A | SG |  | S1 DS, Cold Beam, Upstream Face | V | DC |  | 2.5 |  | 62 | 2a | FORMULA | SG_From_cRIO[43] | DINT | nV/V
-0177-3 | SG817D1HB2 | 2 | J22 | B | SG |  | S1 DS, Cold Beam, Downstream Face | V | DC |  | 2.5 |  | 63 | 2b | FORMULA | SG_From_cRIO[44] | DINT | nV/V
-0177-3 | SG817D2HB1 | 2 | J23 | A | SG |  | S2 DS, Cold Beam, Upstream Face | V | DC |  | 2.5 |  | 64 | 2c | FORMULA | SG_From_cRIO[45] | DINT | nV/V
-0177-3 | SG817D2HB2 | 2 | J23 | B | SG |  | S2 DS, Cold Beam, Downstream Face | V | DC |  | 2.5 |  | 65 | 2d | FORMULA | SG_From_cRIO[46] | DINT | nV/V
-0165-1 | TR811AR | 3 | J1 | A | TR | X95459 | CCM A, Return | I | AC | .3-20 |  |  | 38 | 0 | LOOKUP | TR_From_cRIO[10] | DINT | mK
-0165-1 | TR811AS | 3 | J1 | B | TR | X95525 | CCM A, Supply | I | AC | .3-20 |  |  | 39 | 1 | LOOKUP | TR_From_cRIO[11] | DINT | mK
-0165-1 | TR817A5 | 3 | J2 | A | TR | X95531 | CCM A, 5 Front | I | AC | .3-20 |  |  | 3a | 2 | LOOKUP | TR_From_cRIO[5] | DINT | mK
-0165-3 | TR817U2HB1 | 3 | J2 | B | TR | X100305 | S2 Splice | I | AC | .3-20 |  |  | 3b | 3 | LOOKUP | TR_From_cRIO[69] | DINT | mK
-0165-3 | TR817A1 | 3 | J3 | A | TR | X95462 | CCM A, 1 Front | I | AC | .3-20 |  | #REF! | 3c | 4 | LOOKUP | TR_From_cRIO[1] | DINT | mK
-0165-3 | TR817A2 | 3 | J3 | B | TR | X95597 | CCM A, 2 Front | I | AC | .3-20 |  |  | 3d | 5 | LOOKUP | TR_From_cRIO[2] | DINT | mK
-0165-3 | TR817A3 | 3 | J4 | A | TR | X95477 | CCM A, 3 Front | I | AC | .3-20 |  |  | 3e | 6 | LOOKUP | TR_From_cRIO[3] | DINT | mK
-0165-3 | TR817A4 | 3 | J4 | B | TR | X95488 | CCM A, 4 Front | I | AC | .3-20 |  |  | 3f | 7 | LOOKUP | TR_From_cRIO[4] | DINT | mK
-0165-3 | TR817A6 | 3 | J5 | A | TR | X95600 | CCM A, 1 Back | I | AC | .3-20 |  |  | 40 | 8 | LOOKUP | TR_From_cRIO[6] | DINT | mK
-0165-3 | TR817A7 | 3 | J5 | B | TR | X95656 | CCM A, 2 Back | I | AC | .3-20 |  |  | 41 | 9 | LOOKUP | TR_From_cRIO[7] | DINT | mK
-0165-3 | TR817A8 | 3 | J6 | A | TR | X95546 | CCM A, 3 Back | I | AC | .3-20 |  |  | 42 | a | LOOKUP | TR_From_cRIO[8] | DINT | mK
-0165-3 | TR817A9 | 3 | J6 | B | TR | X95475 | CCM A, 4 Back | I | AC | .3-20 |  |  | 43 | b | LOOKUP | TR_From_cRIO[9] | DINT | mK
-0174-1 | TR817U3HB2 | 3 | J7 | A | TR | X95532 | S3 US Cold Beam | I | AC | .3-20 |  |  | 44 | c | LOOKUP | TR_From_cRIO[72] | DINT | mK
-0177-1 | TR817D3HB1 | 3 | J7 | B | TR | X100373 | S3 DS, Cold Beam | I | AC | .3-20 |  |  | 45 | d | LOOKUP | TR_From_cRIO[80] | DINT | mK
-0166-1 | TR811BR | 4 | J1 | A | TR | X95649 | CCM B, Return | I | AC | .3-20 |  |  | 38 | 0 | LOOKUP | TR_From_cRIO[21] | DINT | mK
-0166-1 | TR811BS | 4 | J1 | B | TR | X95653 | CCM B, Supply | I | AC | .3-20 |  |  | 39 | 1 | LOOKUP | TR_From_cRIO[22] | DINT | mK
-0166-1 | TR817B5 | 4 | J2 | A | TR | X95468 | CCM B, 5 Front | I | AC | .3-20 |  |  | 3a | 2 | LOOKUP | TR_From_cRIO[16] | DINT | mK
-0166-3 | TR817U3HB1 | 4 | J2 | B | TR | X100356 | S3 Splice | I | AC | .3-20 |  |  | 3b | 3 | LOOKUP | TR_From_cRIO[71] | DINT | mK
-0166-3 | TR817B1 | 4 | J3 | A | TR | X95472 | CCM B, 1 Front | I | AC | .3-20 |  |  | 3c | 4 | LOOKUP | TR_From_cRIO[12] | DINT | mK
-0166-3 | TR817B2 | 4 | J3 | B | TR | X95473 | CCM B, 2 Front | I | AC | .3-20 |  |  | 3d | 5 | LOOKUP | TR_From_cRIO[13] | DINT | mK
-0166-3 | TR817B3 | 4 | J4 | A | TR | X95470 | CCM B, 3 Front | I | AC | .3-20 |  |  | 3e | 6 | LOOKUP | TR_From_cRIO[14] | DINT | mK
-0166-3 | TR817B4 | 4 | J4 | B | TR | X95469 | CCM B, 4 Front | I | AC | .3-20 |  |  | 3f | 7 | LOOKUP | TR_From_cRIO[15] | DINT | mK
-0166-3 | TR817B6 | 4 | J5 | A | TR | X91891 | CCM B, 1 Back | I | AC | .3-20 |  |  | 40 | 8 | LOOKUP | TR_From_cRIO[17] | DINT | mK
-0166-3 | TR817B7 | 4 | J5 | B | TR | X91886 | CCM B, 2 Back | I | AC | .3-20 |  |  | 41 | 9 | LOOKUP | TR_From_cRIO[18] | DINT | mK
-0166-3 | TR817B8 | 4 | J6 | A | TR | X91936 | CCM B, 3 Back | I | AC | .3-20 |  |  | 42 | a | LOOKUP | TR_From_cRIO[19] | DINT | mK
-0166-3 | TR817B9 | 4 | J6 | B | TR | X91940 | CCM B, 4 Back | I | AC | .3-20 |  |  | 43 | b | LOOKUP | TR_From_cRIO[20] | DINT | mK
-0167-1 | TR811CR | 5 | J1 | A | TR | X95474 | CCM C, Return | I | AC | .3-20 |  |  | 38 | 0 | LOOKUP | TR_From_cRIO[32] | DINT | mK
-0167-1 | TR811CS | 5 | J1 | B | TR | X95476 | CCM C, Supply | I | AC | .3-20 |  |  | 39 | 1 | LOOKUP | TR_From_cRIO[33] | DINT | mK
-0167-1 | TR817C5 | 5 | J2 | A | TR | X95478 | CCM C, 5 Front | I | AC | .3-20 |  |  | 3a | 2 | LOOKUP | TR_From_cRIO[27] | DINT | mK
-0167-3 | TR817U4HB1 | 5 | J2 | B | TR | X95651 | S4 Splice | I | AC | .3-20 |  |  | 3b | 3 | LOOKUP | TR_From_cRIO[73] | DINT | mK
-0167-3 | TR817C1 | 5 | J3 | A | TR | X95496 | CCM C, 1 Front | I | AC | .3-20 |  |  | 3c | 4 | LOOKUP | TR_From_cRIO[23] | DINT | mK
-0167-3 | TR817C2 | 5 | J3 | B | TR | X95498 | CCM C, 2 Front | I | AC | .3-20 |  |  | 3d | 5 | LOOKUP | TR_From_cRIO[24] | DINT | mK
-0167-3 | TR817C3 | 5 | J4 | A | TR | X95499 | CCM C, 3 Front | I | AC | .3-20 |  |  | 3e | 6 | LOOKUP | TR_From_cRIO[25] | DINT | mK
-0167-3 | TR817C4 | 5 | J4 | B | TR | X95500 | CCM C, 4 Front | I | AC | .3-20 |  |  | 3f | 7 | LOOKUP | TR_From_cRIO[26] | DINT | mK
-0167-3 | TR817C6 | 5 | J5 | A | TR | X95488 | CCM C, 1 Back | I | AC | .3-20 |  |  | 40 | 8 | LOOKUP | TR_From_cRIO[28] | DINT | mK
-0167-3 | TR817C7 | 5 | J5 | B | TR | X95487 | CCM C, 2 Back | I | AC | .3-20 |  |  | 41 | 9 | LOOKUP | TR_From_cRIO[29] | DINT | mK
-0167-3 | TR817C8 | 5 | J6 | A | TR | X95490 | CCM C, 3 Back | I | AC | .3-20 |  |  | 42 | a | LOOKUP | TR_From_cRIO[30] | DINT | mK
-0167-3 | TR817C9 | 5 | J6 | B | TR | X95495 | CCM C, 4 Back | I | AC | .3-20 |  |  | 43 | b | LOOKUP | TR_From_cRIO[31] | DINT | mK
-0175-1 | TR817U4HB2 | 5 | J7 | A | TR | X95492 | S4 US Cold Beam | I | AC | .3-20 |  |  | 44 | c | LOOKUP | TR_From_cRIO[74] | DINT | mK
-0175-1 | TR817U5HB2 | 5 | J7 | B | TR | X95736 | S5 US Cold Beam | I | AC | .3-20 |  |  | 45 | d | LOOKUP | TR_From_cRIO[76] | DINT | mK
-0168-1 | TR811DR | 6 | J1 | A | TR | X95596 | CCM D, Return | I | AC | .3-20 |  |  | 38 | 0 | LOOKUP | TR_From_cRIO[43] | DINT | mK
-0168-1 | TR811DS | 6 | J1 | B | TR | X95598 | CCM D, Supply | I | AC | .3-20 |  |  | 39 | 1 | LOOKUP | TR_From_cRIO[44] | DINT | mK
-0168-1 | TR817D5 | 6 | J2 | A | TR | X95602 | CCM D, 5 Front | I | AC | .3-20 |  |  | 3a | 2 | LOOKUP | TR_From_cRIO[38] | DINT | mK
-0168-3 | TR817U5HB1 | 6 | J2 | B | TR | X95530 | S5 Splice | I | AC | .3-20 |  |  | 3b | 3 | LOOKUP | TR_From_cRIO[75] | DINT | mK
-0168-3 | TR817D1 | 6 | J3 | A | TR | X95599 | CCM D, 1 Front | I | AC | .3-20 |  |  | 3c | 4 | LOOKUP | TR_From_cRIO[34] | DINT | mK
-0168-3 | TR817D2 | 6 | J3 | B | TR | X95594 | CCM D, 2 Front | I | AC | .3-20 |  |  | 3d | 5 | LOOKUP | TR_From_cRIO[35] | DINT | mK
-0168-3 | TR817D3 | 6 | J4 | A | TR | X95528 | CCM D, 3 Front | I | AC | .3-20 |  |  | 3e | 6 | LOOKUP | TR_From_cRIO[36] | DINT | mK
-0168-3 | TR817D4 | 6 | J4 | B | TR | X95601 | CCM D, 4 Front | I | AC | .3-20 |  |  | 3f | 7 | LOOKUP | TR_From_cRIO[37] | DINT | mK
-0168-3 | TR817D6 | 6 | J5 | A | TR | X95524 | CCM D, 1 Back | I | AC | .3-20 |  |  | 40 | 8 | LOOKUP | TR_From_cRIO[39] | DINT | mK
-0168-3 | TR817D7 | 6 | J5 | B | TR | X95529 | CCM D, 2 Back | I | AC | .3-20 |  |  | 41 | 9 | LOOKUP | TR_From_cRIO[40] | DINT | mK
-0168-3 | TR817D8 | 6 | J6 | A | TR | X95593 | CCM D, 3 Back | I | AC | .3-20 |  |  | 42 | a | LOOKUP | TR_From_cRIO[41] | DINT | mK
-0168-3 | TR817D9 | 6 | J6 | B | TR | X95595 | CCM D, 4 Back | I | AC | .3-20 |  |  | 43 | b | LOOKUP | TR_From_cRIO[42] | DINT | mK
-0176-1 | TR817U6HB2 | 6 | J7 | A | TR | X95728 | S6 US Cold Beam | I | AC | .3-20 |  |  | 44 | c | LOOKUP | TR_From_cRIO[78] | DINT | mK
-0169-1 | TR811ER | 1 | J1 | A | TR | X95585 | CCM E, Return | I | AC | .3-20 |  |  | 38 | 0 | LOOKUP | TR_From_cRIO[54] | DINT | mK
-0169-1 | TR811ES | 1 | J1 | B | TR | X95534 | CCM E, Supply | I | AC | .3-20 |  |  | 39 | 1 | LOOKUP | TR_From_cRIO[55] | DINT | mK
-0169-1 | TR817E5 | 1 | J2 | A | TR | X95533 | CCM E, 5 Front | I | AC | .3-20 |  |  | 3a | 2 | LOOKUP | TR_From_cRIO[49] | DINT | mK
-0169-3 | TR817U6HB1 | 1 | J2 | B | TR | X100386 | S6 Splice | I | AC | .3-20 |  |  | 3b | 3 | LOOKUP | TR_From_cRIO[77] | DINT | mK
-0169-3 | TR817E1 | 1 | J3 | A | TR | X91885 | CCM E, 1 Front | I | AC | .3-20 |  |  | 3c | 4 | LOOKUP | TR_From_cRIO[45] | DINT | mK
-0169-3 | TR817E2 | 1 | J3 | B | TR | X91887 | CCM E, 2 Front | I | AC | .3-20 |  |  | 3d | 5 | LOOKUP | TR_From_cRIO[46] | DINT | mK
-0169-3 | TR817E3 | 1 | J4 | A | TR | X91937 | CCM E, 3 Front | I | AC | .3-20 |  |  | 3e | 6 | LOOKUP | TR_From_cRIO[47] | DINT | mK
-0169-3 | TR817E4 | 1 | J4 | B | TR | X91945 | CCM E, 4 Front | I | AC | .3-20 |  |  | 3f | 7 | LOOKUP | TR_From_cRIO[48] | DINT | mK
-0169-3 | TR817E6 | 1 | J5 | A | TR | X91946 | CCM E, 1 Back | I | AC | .3-20 |  |  | 40 | 8 | LOOKUP | TR_From_cRIO[50] | DINT | mK
-0169-3 | TR817E7 | 1 | J5 | B | TR | X91947 | CCM E, 2 Back | I | AC | .3-20 |  |  | 41 | 9 | LOOKUP | TR_From_cRIO[51] | DINT | mK
-0169-3 | TR817E8 | 1 | J6 | A | TR | X95485 | CCM E, 3 Back | I | AC | .3-20 |  |  | 42 | a | LOOKUP | TR_From_cRIO[52] | DINT | mK
-0169-3 | TR817E9 | 1 | J6 | B | TR | X95497 | CCM E, 4 Back | I | AC | .3-20 |  |  | 43 | b | LOOKUP | TR_From_cRIO[53] | DINT | mK
-0178-7 | TR8122A | 1 | J7 | A | TR | X100362 | Chimney splice top | I | AC | .3-20 |  |  | 44 | c | LOOKUP | TR_From_cRIO[81] | DINT | mK | moved from CRYOCON
-0178-7 | TR8122B | 1 | J7 | B | TR | X100302 | Chimney splice bottom | I | AC | .3-20 |  |  | 45 | d | LOOKUP | TR_From_cRIO[82] | DINT | mK | moved from CRYOCON
-0178-8 | TR8114S | 1 | J8 | A | TR | X103563 | SUPPLY ds HEX RING | I | AC | .3-20 |  |  | 46 | e | LOOKUP | TR_From_cRIO[84] | DINT | mK | moved from CRYOCON
-0178-8 | TR8114R | 1 | J8 | B | TR | X100335 | RETURN DS HEX RING | I | AC | .3-20 |  |  | 47 | f | LOOKUP | TR_From_cRIO[83] | DINT | mK | moved from CRYOCON
-0170-1 | TR811FR | 2 | J1 | A | TR | X95663 | CCM F, Return | I | AC | .3-20 |  |  | 38 | 0 | LOOKUP | TR_From_cRIO[65] | DINT | mK
-0170-1 | TR811FS | 2 | J1 | B | TR | X95554 | CCM F, Supply | I | AC | .3-20 |  |  | 39 | 1 | LOOKUP | TR_From_cRIO[66] | DINT | mK
-0170-1 | TR817F5 | 2 | J2 | A | TR | X95730 | CCM F, 5 Front | I | AC | .3-20 |  |  | 3a | 2 | LOOKUP | TR_From_cRIO[60] | DINT | mK
-0170-3 | TR817U1HB1 | 2 | J2 | B | TR | X95467 | S1 Splice | I | AC | .3-20 |  |  | 3b | 3 | LOOKUP | TR_From_cRIO[67] | DINT | mK
-0170-3 | TR817F1 | 2 | J3 | A | TR | X95733 | CCM F, 1 Front | I | AC | .3-20 |  |  | 3c | 4 | LOOKUP | TR_From_cRIO[56] | DINT | mK
-0170-3 | TR817F2 | 2 | J3 | B | TR | X95747 | CCM F, 2 Front | I | AC | .3-20 |  |  | 3d | 5 | LOOKUP | TR_From_cRIO[57] | DINT | mK
-0170-3 | TR817F3 | 2 | J4 | A | TR | X95535 | CCM F, 3 Front | I | AC | .3-20 |  |  | 3e | 6 | LOOKUP | TR_From_cRIO[58] | DINT | mK
-0170-3 | TR817F4 | 2 | J4 | B | TR | X95545 | CCM F, 4 Front | I | AC | .3-20 |  |  | 3f | 7 | LOOKUP | TR_From_cRIO[59] | DINT | mK
-0170-3 | TR817F6 | 2 | J5 | A | TR | X95547 | CCM F, 1 Back | I | AC | .3-20 |  |  | 40 | 8 | LOOKUP | TR_From_cRIO[61] | DINT | mK
-0170-3 | TR817F7 | 2 | J5 | B | TR | X95549 | CCM F, 2 Back | I | AC | .3-20 |  |  | 41 | 9 | LOOKUP | TR_From_cRIO[62] | DINT | mK
-0170-3 | TR817F8 | 2 | J6 | A | TR | X95550 | CCM F, 3 Back | I | AC | .3-20 |  |  | 42 | a | LOOKUP | TR_From_cRIO[63] | DINT | mK
-0170-3 | TR817F9 | 2 | J6 | B | TR | X95553 | CCM F, 4 Back | I | AC | .3-20 |  |  | 43 | b | LOOKUP | TR_From_cRIO[64] | DINT | mK
-0173-1 | TR817U1HB2 | 2 | J7 | A | TR | X95479 | S1 US Cold Beam | I | AC | .3-20 |  |  | 44 | c | LOOKUP | TR_From_cRIO[68] | DINT | mK
-0173-1 | TR817U2HB2 | 2 | J7 | B | TR | X95471 | S2 US Cold Beam | I | AC | .3-20 |  |  | 45 | d | LOOKUP | TR_From_cRIO[70] | DINT | mK
-0177-3 | TR817D2HB1 | 2 | J8 | A | TR | X95463 | S2 DS, Cold Beam | I | AC | .3-20 |  |  | 46 | e | LOOKUP | TR_From_cRIO[79] | DINT | mK