## Summary
Connector | Rack Loc. | Device/Sensor Name | Device/Sensor Type | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 1, 1A-01 | FL03B25 | (Piezo Actuator) FastTuner-1-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-1-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
|  | (Piezo Actuator) FastTuner-1-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-1-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-1 | Piezo Fast Tuner, Interlock |  |  | J-K & L-M
CM 1, 1A-02 | FL03B25 | RTD_StepperMotor-1 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler1a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler1b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 1A-03 | FL02B27 | RTD_HeVessel-1 Top | Cernox CX-1030-SD-HT-1.4L | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_HeVessel-1 Bottom | Cernox CX-1030-SD-HT-1.4L |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-1-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-1-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 1B-01 | FL03B25 | (Stepper Motor) SlowTuner-1 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
|  | Tuner limit Switch-1-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-1-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 1, 1B-02 | FL03B25 | C1_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, 1B-03 | FL03B24 | C1_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 1B-04 | FL03B24 | C1_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 1B-05 | FL03B25 | Cav_1_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 1B-06 | FL03B25 | Cav_1_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 1B-07 | FL02B27 | StrpHtr-HeVessel-1 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 1, Cav 1-F | FL03B25 | Cavity 1 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, Cav 1-R | FL03B25 | Cavity 1 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
Connector |  | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 1, 2A-01 | FL03B25 | (Piezo Actuator) FastTuner-2-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-2-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
|  | (Piezo Actuator) FastTuner-2-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-2-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-2 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 1, 2A-02 | FL03B25 | RTD_StepperMotor-2 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler2a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler2b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 2A-03 | FL02B27 | Spare | N/A | 4 pr 22 ga | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-2-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-2-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 2B-01 | FL03B25 | (Stepper Motor) SlowTuner-2 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
|  | Tuner limit Switch-2-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-2-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 1, 2B-02 | FL03B25 | C2_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, 2B-03 | FL03B24 | C2_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 2B-04 | FL03B24 | C2_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 2B-05 | FL03B25 | Cav_2_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 2B-06 | FL03B25 | Cav_2_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 2B-07 | FL02B27 | StrpHtr-HeVessel-2 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 1, Cav 2-F | FL03B25 | Cavity 2 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, Cav 2-R | FL03B25 | Cavity 2 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
Connector |  | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 1, 3A-01 | FL03B25 | (Piezo Actuator) FastTuner-3-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-3-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
|  | (Piezo Actuator) FastTuner-3-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-3-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-3 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 1, 3A-02 | FL03B25 | RTD_StepperMotor-3 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler2a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler2b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 3A-03 | FL02B27 | Spare | N/A | 4 pr 22 ga | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-3-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-3-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 3B-01 | FL03B25 | (Stepper Motor) SlowTuner-3 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
|  | Tuner limit Switch-3-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-3-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 1, 3B-02 | FL03B25 | C3_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, 3B-03 | FL03B24 | C3_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 3B-04 | FL03B24 | C3_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 3B-05 | FL03B25 | Cav_3_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 3B-06 | FL03B25 | Cav_3_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 3B-07 | FL02B27 | StrpHtr-HeVessel-3 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 1, Cav 3-F | FL03B25 | Cavity 3 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, Cav 3-R | FL03B25 | Cavity 3 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
Connector |  | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 1, 4A-01 | FL03B25 | (Piezo Actuator) FastTuner-4-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-4-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
|  | (Piezo Actuator) FastTuner-4-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-4-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-4 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 1, 4A-02 | FL03B25 | RTD_StepperMotor-4 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler4a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler4b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 4A-03 | FL02B27 | Spare | N/A | 4 pr 22 ga | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-4-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-4-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 4B-01 | FL03B25 | (Stepper Motor) SlowTuner-4 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
|  | Tuner limit Switch-4-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-4-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 1, 4B-02 | FL03B25 | C4_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, 4B-03 | FL03B24 | C4_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 4B-04 | FL03B24 | C4_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 4B-05 | FL03B25 | Cav_4_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 4B-06 | FL03B25 | Cav_4_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 4B-07 | FL02B27 | StrpHtr-HeVessel-4 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 1, Cav 4-F | FL03B25 | Cavity 4 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, Cav 4-R | FL03B25 | Cavity 4 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
Connector |  | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 1, 5A-01 |  | (Piezo Actuator) FastTuner-5-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-5-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
| FL03B34 | (Piezo Actuator) FastTuner-5-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-5-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-5 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 1, 5A-02 |  | RTD_StepperMotor-5 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL03B34 | RTD_MainCoupler5a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler5b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 5A-03 |  | RTD_HeVessel-5 Top | Cernox CX-1030-SD-HT-1.4L | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL02B27 | RTD_HeVessel-5 Bottom | Cernox CX-1030-SD-HT-1.4L |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-5-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-5-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 5B-01 |  | (Stepper Motor) SlowTuner-5 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
| FL03B34 | Tuner limit Switch-5-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-5-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 1, 5B-02 | FL03B34 | C5_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, 5B-03 | FL03B24 | C5_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 5B-04 | FL03B24 | C5_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 5B-05 | FL03B34 | Cav_5_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 5B-06 | FL03B34 | Cav_5_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 5B-07 | FL02B27 | StrpHtr-HeVessel-5 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 1, Cav 5-F | FL03B34 | Cavity 5 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, Cav 5-R | FL03B34 | Cavity 5 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
Connector |  | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 1, 6A-01 |  | (Piezo Actuator) FastTuner-6-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-6-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
| FL03B34 | (Piezo Actuator) FastTuner-6-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-6-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-6 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 1, 6A-02 |  | RTD_StepperMotor-6 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL03B34 | RTD_MainCoupler6a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler6b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 6A-03 |  | Spare | N/A | 4 pr 22 ga | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL02B27 | Spare | N/A |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-6-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-6-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 6B-01 |  | (Stepper Motor) SlowTuner-6 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
| FL03B34 | Tuner limit Switch-6-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-6-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 1, 6B-02 | FL03B34 | C6_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, 6B-03 | FL03B24 | C6_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 6B-04 | FL03B24 | C6_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 6B-05 | FL03B34 | Cav_6_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 6B-06 | FL03B34 | Cav_6_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 6B-07 | FL02B27 | StrpHtr-HeVessel-6 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 1, Cav 6-F | FL03B34 | Cavity 6 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, Cav 6-R | FL03B34 | Cavity 6 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
Connector |  | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 1, 7A-01 |  | (Piezo Actuator) FastTuner-7-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-7-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
| FL03B34 | (Piezo Actuator) FastTuner-7-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-7-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-7 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 1, 7A-02 |  | RTD_StepperMotor-7 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL03B34 | RTD_MainCoupler7a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler7b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 7A-03 |  | Spare | N/A | 4 pr 22 ga | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL02B27 | Spare | N/A |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-7-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-7-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 7B-01 |  | (Stepper Motor) SlowTuner-7 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
| FL03B34 | Tuner limit Switch-7-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-7-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 1, 7B-02 | FL03B34 | C7_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, 7B-03 | FL03B24 | C7_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 7B-04 | FL03B24 | C7_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 7B-05 | FL03B34 | Cav_7_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 7B-06 | FL03B34 | Cav_7_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 7B-07 | FL02B27 | StrpHtr-HeVessel-7 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 1, Cav 7-F | FL03B34 | Cavity 7 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, Cav 7-R | FL03B34 | Cavity 7 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 1, 8A-01 |  | (Piezo Actuator) FastTuner-8-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-8-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
| FL03B34 | (Piezo Actuator) FastTuner-8-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-8-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-8 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 1, 8A-02 |  | RTD_StepperMotor-8 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL03B34 | RTD_MainCoupler8a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler8b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 8A-03 |  | Spare | N/A | 4 pr 22 ga | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL02B27 | Spare | N/A |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-8-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-8-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 8B-01 |  | (Stepper Motor) SlowTuner-8 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
| FL03B34 | Tuner limit Switch-8-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-8-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 1, 8B-02 | FL03B34 | C8_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, 8B-03 | FL03B24 | C8_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 8B-04 | FL03B24 | C8_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 8B-05 | FL03B34 | Cav_8_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 8B-06 | FL03B34 | Cav_8_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 8B-07 | FL02B27 | StrpHtr-HeVessel-8 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 1, Cav 8-F | FL03B34 | Cavity 8 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, Cav 8-R | FL03B34 | Cavity 8 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection (Warm Side) | Feed Through Connector Pin
CM 1, C-01 |  | VD VT Positive Lead (Whole_Coil-VT6) | Vertical Dipole Coil Voltage Tap (+) | 8cond 22ga Alpha 1298C | MS3126F16-8S | Pin-A
|  | VD VT Negative Lead (Whole_Coil-VT5) | Vertical Dipole Coil Voltage Tap (-) |  |  | Pin-B
|  | HD VT Positive Lead (Whole_Coil-VT4) | Horizontal Dipole Coil Voltage Tap (+) |  |  | Pin-C
|  | HD VT Negative Lead (Whole_Coil-VT3) | Horizontal Dipole Coil Voltage Tap (-) |  |  | Pin-D
|  | Quad VT Positive Lead (Whole_Coil-VT1) | Split Quadrupole Coil Voltage Tap (+) |  |  | Pin-E
|  | Quad VT Negative Lead (Whole_Coil-VT2) | Split Quadrupole Coil Voltage Tap (-) |  |  | Pin-F
|  | Spare | Spare |  |  | Pin-G
|  | Spare | Spare |  |  | Pin-H
CM 1, C-02 |  | Vertical Dipole VTRes28 (PosHalf Coil) | Vertical Dipole Coil Voltage Tap (+) | 8cond 22ga Alpha 1298C | MS3126F16-8S | Pin-A
|  | Vertical Dipole VTRes29 (PosHalf Coil) | Vertical Dipole Coil Voltage Tap (-) |  |  | Pin-B
|  | Vertical Dipole VTRes30 (NegHalf Coil) | Vertical Dipole Coil Voltage Tap (+) |  |  | Pin-C
|  | Vertical Dipole VTRes31 (NegHalf Coil) | Vertical Dipole Coil Voltage Tap (-) |  |  | Pin-D
|  | Horizontal Dipole VTRes22 (PosHalf Coil) | Horizontal Dipole Coil Voltage Tap (+) |  |  | Pin-E
|  | Horizontal Dipole VTRes23 (PosHalf Coil) | Horizontal Dipole Coil Voltage Tap (-) |  |  | Pin-F
|  | Horizontal Dipole VTRes24 (NegHalf Coil) | Horizontal Dipole Coil Voltage Tap (+) |  |  | Pin-G
|  | Horizontal Dipole VTRes25 (NegHalf Coil) | Horizontal Dipole Coil Voltage Tap (-) |  |  | Pin-H
CM 1, C-03 |  | Quadrupole VTRes16 (PosHalf Coil) | Split Quadrupole Coil Voltage Tap (+) | 8cond 22ga Alpha 1298C | MS3126F16-8S | Pin-A
|  | Quadrupole VTRes17 (PosHalf Coil) | Split Quadrupole Coil Voltage Tap (-) |  |  | Pin-B
|  | Quadrupole VTRes18 (NegHalf Coil) | Split Quadrupole Coil Voltage Tap (+) |  |  | Pin-C
|  | Quadrupole VTRes19 (NegHalf Coil) | Split Quadrupole Coil Voltage Tap (-) |  |  | Pin-D
|  | Spare | Spare |  |  | Pin-E
|  | Spare | Spare |  |  | Pin-F
|  | Spare | Spare |  |  | Pin-G
|  | Spare | Spare |  |  | Pin-H
CM 1, C-04 | FL02B27 | Heater Q1 & Q2 | Split Quadrupole and Corrector Coils Persistant Current Magnet Heaters | 4 cond 18 ga Alpha 2424C | MS24266R14B4SN | 1(+), 2(-)
|  | Heater Q3 & Q4 |  |  |  | 3(+), 4(-)
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 1, D-01 |  | BPM-Top | Beam Position Monitor |  | N 00 | Single Pin,Shield
CM 1, D-02 |  | BPM-Right | Beam Position Monitor |  | N 00 | Single Pin,Shield
CM 1, D-03 |  | BPM-Bot | Beam Position Monitor |  | N 00 | Single Pin,Shield
CM 1, D-04 |  | BPM-Left | Beam Position Monitor |  | N 00 | Single Pin,Shield
CM 1, D-05 |  | Fluxgate_Cav1 | Fluxgate | 6pr 20ga | MS3126F14-19S | A,B,C,D (S+,S-,E+,E-)
|  | Fluxgate_Cav2 | Fluxgate |  |  | E,F,G,H (S+,S-,E+,E-)
|  | Fluxgate_Cav5 | Fluxgate |  |  | J,K,L,M (S+,S-,E+,E-)
|  | spare | Spare |  |  | N,P,R,S
CM 1, D-06 |  | Fluxgate_Cav7 | Fluxgate | 6pr 20ga | MS3126F14-19S | A,B,C,D (S+,S-,E+,E-)
|  | Fluxgate_Cav8 | Fluxgate |  |  | E,F,G,H (S+,S-,E+,E-)
|  | spare | Spare |  |  | J,K,L,M
|  | spare | Spare |  |  | N,P,R,S
CM 1, D-07 | FL02B27 | CoolDownHtr-1 | Primary Cool Down Heater #1 Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) , B(-)
|  | CoolDownHtr-2 | Primary Cool Down Heater #2 Kapton Polymide Strip Heater, 46.3 Ω |  |  | C(+) , D(-)
CM 1, D-08 | FL02B27 | CoolDownHtr-3 | Primary Cool Down Heater #3 Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) , B(-)
|  | Heater in Downstream Level Can | W4C050 50V, 100 Watt Cartridge Heater |  |  | C(+) , D(-)
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 1, E1-01 | FL02B27 | LL_Downstream | Cryomagnetics 3D-031-033-PB-RAD-2K 12" Liquid Level Probe | 4 cond 18 ga; Alpha 2424C | 16060-02-A | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | Spare | Jumper/ Potential for Interlock |  |  | E,F
| FL02B27 | Baratron
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 1, E2-01 | FL02B27 | LL_Downstream | Cryomagnetics 3D-031-033-PB-RAD-2K 12" Liquid Level Probe | 4 cond 18 ga; Alpha 2424C | MS3126F10-6S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | Spare | Jumper/ Potential for Interlock |  |  | E,F
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 1, F-01 | FL02B27 | RTD_MP_1 | Cernox CX-1030-SD-HT-1.4L | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_MP_2 | DT-670A1-SD |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MP_3 | DT-670A1-SD |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_MP_4 | DT-670A1-SD |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, F-02 | FL02B27 | RTD_50K_LeadThermIntcp | DT-670A1-SD | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_5K_LeadThermIntcp | DT-670A1-SD |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_2K_LeadThermIntcp_1 | DT-670A1-SD |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_2K_LeadThermIntcp_2 | DT-670A1-SD |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 1, K-01 | FL02B27 | RTD_Line_A1 | DT-670A1-CU | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_Line_B1 | DT-670A1-CU |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_Line_B2 | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_Line_C1 | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, K-02 | FL02B27 | RTD_Line_D1 | DT-670A1-CU | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_Line_E1 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_Line_F1 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_Shield_1 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, K-03 | FL02B27 | RTD_Cooldown_1 | Cernox CX-1030-CU-HT-1.4L | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_Cooldown_2 | Cernox CX-1030-CU-HT-1.4L |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_Faraday Window 1 Port | Cernox CX-1030-SD-HT |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_Faraday Window 2 Port | Cernox CX-1030-SD-HT |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, K-04 | FL02B27 | Heater in Upstream Level Can | W4C050 50V, 100 Watt Cartridge Heater | 4cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 1, L1-01 | FL02B27 | LL_Upstream | Cryomagnetics 3D-031-033-PB-RAD-2K 12" Liquid Level Probe | 4 cond 18 ga; Alpha 2424C | 16060-02-A | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | Spare | Jumper/ Potential for Interlock |  |  | E, F
CM 1, L1-BAR | FL02B27 | Baratron
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
L2-01 | FL02B27 | LL_Upstream | Cryomagnetics 3D-031-033-PB-RAD-2K 12" Liquid Level Probe | 4 cond 18 ga; Alpha 2424C | MS3126F10-6S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | Spare | Jumper/ Potential for Interlock |  |  | E,F
Connector | Rack | Device/Sensor Name | Device/Sensor Type | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 2, 1A-01 |  | (Piezo Actuator) FastTuner-1-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-1-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
| FL04B34 | (Piezo Actuator) FastTuner-1-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-1-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-1 | Piezo Fast Tuner, Interlock |  |  | J-K & L-M
CM 2, 1A-02 |  | RTD_StepperMotor-1 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL04B34 | RTD_MainCoupler1a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler1b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 1A-03 |  | RTD_HeVessel-1 Top | Cernox CX-1030-SD-HT-1.4L | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL02B27 | RTD_HeVessel-1 Bottom | Cernox CX-1030-SD-HT-1.4L |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-1-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-1-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 1B-01 |  | (Stepper Motor) SlowTuner-1 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
| FL04B34 | Tuner limit Switch-1-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-1-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 2, 1B-02 | FL04B34 | C1_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, 1B-03 | FL03B24 | C1_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 1B-04 | FL03B24 | C1_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 1B-05 | FL04B34 | Cav_1_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 1B-06 | FL04B34 | Cav_1_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 1B-07 | FL02B27 | StrpHtr-HeVessel-1 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 2, Cav 1-F | FL04B34 | Cavity 1 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, Cav 1-R | FL04B34 | Cavity 1 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 2, 2A-01 |  | (Piezo Actuator) FastTuner-2-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-2-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
| FL04B34 | (Piezo Actuator) FastTuner-2-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-2-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-2 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 2, 2A-02 |  | RTD_StepperMotor-2 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL04B34 | RTD_MainCoupler2a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler2b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 2A-03 |  | Spare | N/A | 4 pr 22 ga | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL02B27 | Spare | N/A |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-2-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-2-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 2B-01 |  | (Stepper Motor) SlowTuner-2 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
| FL04B34 | Tuner limit Switch-2-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-2-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 2, 2B-02 | FL04B34 | C2_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, 2B-03 | FL03B24 | C2_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 2B-04 | FL03B24 | C2_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 2B-05 | FL04B34 | Cav_2_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 2B-06 | FL04B34 | Cav_2_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 2B-07 | FL02B27 | StrpHtr-HeVessel-2 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 2, Cav 2-F | FL04B34 | Cavity 2 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, Cav 2-R | FL04B34 | Cavity 2 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 2, 3A-01 |  | (Piezo Actuator) FastTuner-3-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-3-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
| FL04B34 | (Piezo Actuator) FastTuner-3-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-3-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-3 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 2, 3A-02 |  | RTD_StepperMotor-3 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL04B34 | RTD_MainCoupler2a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler2b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 3A-03 |  | Spare | N/A | 4 pr 22 ga | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL02B27 | Spare | N/A |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-3-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-3-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 3B-01 |  | (Stepper Motor) SlowTuner-3 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
| FL04B34 | Tuner limit Switch-3-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-3-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 2, 3B-02 | FL04B34 | C3_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, 3B-03 | FL03B24 | C3_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 3B-04 | FL03B24 | C3_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 3B-05 | FL04B34 | Cav_3_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 3B-06 | FL04B34 | Cav_3_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 3B-07 | FL02B27 | StrpHtr-HeVessel-3 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 2, Cav 3-F | FL04B34 | Cavity 3 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, Cav 3-R | FL04B34 | Cavity 3 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 2, 4A-01 |  | (Piezo Actuator) FastTuner-4-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-4-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
| FL04B34 | (Piezo Actuator) FastTuner-4-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-4-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-4 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 2, 4A-02 |  | RTD_StepperMotor-4 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL04B34 | RTD_MainCoupler4a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler4b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 4A-03 |  | Spare | N/A | 4 pr 22 ga | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL02B27 | Spare | N/A |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-4-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-4-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 4B-01 |  | (Stepper Motor) SlowTuner-4 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
| FL04B34 | Tuner limit Switch-4-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-4-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 2, 4B-02 | FL04B34 | C4_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, 4B-03 | FL03B24 | C4_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 4B-04 | FL03B24 | C4_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 4B-05 | FL04B34 | Cav_4_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 4B-06 | FL04B34 | Cav_4_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 4B-07 | FL02B27 | StrpHtr-HeVessel-4 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 2, Cav 4-F | FL04B34 | Cavity 4 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, Cav 4-R | FL04B34 | Cavity 4 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 2, 5A-01 |  | (Piezo Actuator) FastTuner-5-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-5-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
| FL04B25 | (Piezo Actuator) FastTuner-5-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-5-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-5 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 2, 5A-02 |  | RTD_StepperMotor-5 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL04B25 | RTD_MainCoupler5a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler5b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 5A-03 |  | RTD_HeVessel-5 Top | Cernox CX-1030-SD-HT-1.4L | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL02B25 | RTD_HeVessel-5 Bottom | Cernox CX-1030-SD-HT-1.4L |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-5-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-5-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 5B-01 |  | (Stepper Motor) SlowTuner-5 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
| FL04B25 | Tuner limit Switch-5-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-5-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 2, 5B-02 | FL04B25 | C5_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, 5B-03 | FL03B24 | C5_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 5B-04 | FL03B24 | C5_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 5B-05 | FL04B25 | Cav_5_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 5B-06 | FL04B25 | Cav_5_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 5B-07 | FL02B25 | StrpHtr-HeVessel-5 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 2, Cav 5-F | FL04B25 | Cavity 5 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, Cav 5-R | FL04B25 | Cavity 5 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 2, 6A-01 |  | (Piezo Actuator) FastTuner-6-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-6-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
| FL04B25 | (Piezo Actuator) FastTuner-6-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-6-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-6 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 2, 6A-02 |  | RTD_StepperMotor-6 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL04B25 | RTD_MainCoupler6a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler6b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 6A-03 |  | Spare | N/A | 4 pr 22 ga | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL02B25 | Spare | N/A |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-6-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-6-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 6B-01 |  | (Stepper Motor) SlowTuner-6 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
| FL04B25 | Tuner limit Switch-6-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-6-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 2, 6B-02 | FL04B25 | C6_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, 6B-03 | FL03B24 | C6_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 6B-04 | FL03B24 | C6_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 6B-05 | FL04B25 | Cav_6_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 6B-06 | FL04B25 | Cav_6_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 6B-07 | FL02B25 | StrpHtr-HeVessel-6 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 2, Cav 6-F | FL04B25 | Cavity 6 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 6, Cav 6-R | FL04B25 | Cavity 6 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 2, 7A-01 |  | (Piezo Actuator) FastTuner-7-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-7-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
| FL04B25 | (Piezo Actuator) FastTuner-7-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-7-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-7 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 2, 7A-02 |  | RTD_StepperMotor-7 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL04B25 | RTD_MainCoupler7a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler7b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 7A-03 |  | Spare | N/A | 4 pr 22 ga | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL02B25 | Spare | N/A |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-7-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-7-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 7B-01 |  | (Stepper Motor) SlowTuner-7 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
| FL04B25 | Tuner limit Switch-7-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-7-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 2, 7B-02 | FL04B25 | C7_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, 7B-03 | FL03B24 | C7_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 7B-04 | FL03B24 | C7_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 7B-05 | FL04B25 | Cav_7_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 7B-06 | FL04B25 | Cav_7_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 7B-07 | FL02B25 | StrpHtr-HeVessel-7 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 2, Cav 7-F | FL04B25 | Cavity 7 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, Cav 7-R | FL04B25 | Cavity 7 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 2, 8A-01 |  | (Piezo Actuator) FastTuner-8-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-8-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
| FL04B25 | (Piezo Actuator) FastTuner-8-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-8-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-8 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 2, 8A-02 |  | RTD_StepperMotor-8 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL04B25 | RTD_MainCoupler8a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler8b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 8A-03 |  | Spare | N/A | 4 pr 22 ga | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL02B25 | Spare | N/A |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-8-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-8-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 8B-01 |  | (Stepper Motor) SlowTuner-8 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
| FL04B25 | Tuner limit Switch-8-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-8-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 2, 8B-02 | FL04B25 | C8_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, 8B-03 | FL03B24 | C8_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 8B-04 | FL03B24 | C8_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 8B-05 | FL04B25 | Cav_8_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 8B-06 | FL04B25 | Cav_8_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 8B-07 | FL02B25 | StrpHtr-HeVessel-8 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 2, Cav 8-F | FL04B25 | Cavity 8 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, Cav 8-R | FL04B25 | Cavity 8 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection (Warm Side) | Feed Through Connector Pin
CM 2, C-01 |  | VD VT Positive Lead (Whole_Coil-VT6) | Vertical Dipole Coil Voltage Tap (+) | 8cond 22ga Alpha 1298C | MS3126F16-8S | Pin-A
|  | VD VT Negative Lead (Whole_Coil-VT5) | Vertical Dipole Coil Voltage Tap (-) |  |  | Pin-B
|  | HD VT Positive Lead (Whole_Coil-VT4) | Horizontal Dipole Coil Voltage Tap (+) |  |  | Pin-C
|  | HD VT Negative Lead (Whole_Coil-VT3) | Horizontal Dipole Coil Voltage Tap (-) |  |  | Pin-D
|  | Quad VT Positive Lead (Whole_Coil-VT1) | Split Quadrupole Coil Voltage Tap (+) |  |  | Pin-E
|  | Quad VT Negative Lead (Whole_Coil-VT2) | Split Quadrupole Coil Voltage Tap (-) |  |  | Pin-F
|  | Spare | Spare |  |  | Pin-G
|  | Spare | Spare |  |  | Pin-H
CM 2, C-02 |  | Vertical Dipole VTRes28 (PosHalf Coil) | Vertical Dipole Coil Voltage Tap (+) | 8cond 22ga Alpha 1298C | MS3126F16-8S | Pin-A
|  | Vertical Dipole VTRes29 (PosHalf Coil) | Vertical Dipole Coil Voltage Tap (-) |  |  | Pin-B
|  | Vertical Dipole VTRes30 (NegHalf Coil) | Vertical Dipole Coil Voltage Tap (+) |  |  | Pin-C
|  | Vertical Dipole VTRes31 (NegHalf Coil) | Vertical Dipole Coil Voltage Tap (-) |  |  | Pin-D
|  | Horizontal Dipole VTRes22 (PosHalf Coil) | Horizontal Dipole Coil Voltage Tap (+) |  |  | Pin-E
|  | Horizontal Dipole VTRes23 (PosHalf Coil) | Horizontal Dipole Coil Voltage Tap (-) |  |  | Pin-F
|  | Horizontal Dipole VTRes24 (NegHalf Coil) | Horizontal Dipole Coil Voltage Tap (+) |  |  | Pin-G
|  | Horizontal Dipole VTRes25 (NegHalf Coil) | Horizontal Dipole Coil Voltage Tap (-) |  |  | Pin-H
CM 2, C-03 |  | Quadrupole VTRes16 (PosHalf Coil) | Split Quadrupole Coil Voltage Tap (+) | 8cond 22ga Alpha 1298C | MS3126F16-8S | Pin-A
|  | Quadrupole VTRes17 (PosHalf Coil) | Split Quadrupole Coil Voltage Tap (-) |  |  | Pin-B
|  | Quadrupole VTRes18 (NegHalf Coil) | Split Quadrupole Coil Voltage Tap (+) |  |  | Pin-C
|  | Quadrupole VTRes19 (NegHalf Coil) | Split Quadrupole Coil Voltage Tap (-) |  |  | Pin-D
|  | Spare | Spare |  |  | Pin-E
|  | Spare | Spare |  |  | Pin-F
|  | Spare | Spare |  |  | Pin-G
|  | Spare | Spare |  |  | Pin-H
CM 2, C-04 |  | Heater Q1 & Q2 | Split Quadrupole and Corrector Coils Persistant Current Magnet Heaters | 4 cond 18 ga Alpha 2424C | MS24266R14B4SN | 1(+), 2(-)
|  | Heater Q3 & Q4 |  |  |  | 3(+), 4(-)
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 2, D-01 |  | BPM-Top | Beam Position Monitor |  | N 00 | Single Pin,Shield
CM 2, D-02 |  | BPM-Right | Beam Position Monitor |  | N 00 | Single Pin,Shield
CM 2, D-03 |  | BPM-Bot | Beam Position Monitor |  | N 00 | Single Pin,Shield
CM 2, D-04 |  | BPM-Left | Beam Position Monitor |  | N 00 | Single Pin,Shield
CM 2, D-05 |  | Fluxgate_Cav1 | Fluxgate |  | MS3126F14-19S | A,B,C,D (S+,S-,E+,E-)
|  | Fluxgate_Cav2 | Fluxgate |  |  | E,F,G,H (S+,S-,E+,E-)
|  | Fluxgate_Cav5 | Fluxgate |  |  | J,K,L,M (S+,S-,E+,E-)
|  | spare | Spare |  |  | N,P,R,S
CM 2, D-06 |  | Fluxgate_Cav7 | Fluxgate |  | MS3126F14-19S | A,B,C,D (S+,S-,E+,E-)
|  | Fluxgate_Cav8 | Fluxgate |  |  | E,F,G,H (S+,S-,E+,E-)
|  | spare | Spare |  |  | J,K,L,M
|  | spare | Spare |  |  | N,P,R,S
CM 2, D-07 | FL02B25 | CoolDownHtr-1 | Primary Cool Down Heater #1 Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) , B(-)
|  | CoolDownHtr-2 | Primary Cool Down Heater #2 Kapton Polymide Strip Heater, 46.3 Ω |  |  | C(+) , D(-)
CM 2, D-08 | FL02B25 | CoolDownHtr-3 | Primary Cool Down Heater #3 Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) , B(-)
|  | Heater in Downstream Level Can | W4C050 50V, 100 Watt Cartridge Heater |  |  | C(+) , D(-)
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 2, E1-01 | FL02B25 | LL_Downstream | Cryomagnetics 3D-031-033-PB-RAD-2K 12" Liquid Level Probe | 2 pr 20 ga | 16060-02-A | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | Spare | Jumper/ Potential for Interlock |  |  | E,F
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 2, E2-01 | FL02B25 | LL_Downstream | Cryomagnetics 3D-031-033-PB-RAD-2K 12" Liquid Level Probe | 2 pr 20 ga | MS3126F10-6S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | Spare | Jumper/ Potential for Interlock |  |  | E,F
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 2, F-01 | FL02B25 | RTD_MP_1 | Cernox CX-1030-SD-HT-1.4L | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_MP_2 | DT-670A1-SD |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MP_3 | DT-670A1-SD |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_MP_4 | DT-670A1-SD |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, F-02 | FL02B25 | RTD_50K_LeadThermIntcp | DT-670A1-SD | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_5K_LeadThermIntcp | DT-670A1-SD |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_2K_LeadThermIntcp_1 | DT-670A1-SD |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_2K_LeadThermIntcp_2 | DT-670A1-SD |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 2, K-01 | FL02B25 | RTD_Line_A1 | DT-670A1-CU | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_Line_B1 | DT-670A1-CU |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_Line_B2 | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_Line_C1 | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, K-02 | FL02B25 | RTD_Line_D1 | DT-670A1-CU | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_Line_E1 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_Line_F1 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_Shield_1 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, K-03 | FL02B25 | RTD_Cooldown_1 | Cernox CX-1030-CU-HT-1.4L | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_Cooldown_2 | Cernox CX-1030-CU-HT-1.4L |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_Faraday Window 1 Port | Cernox CX-1030-SD-HT |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_Faraday Window 2 Port | Cernox CX-1030-SD-HT |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, K-04 | FL02B25 | Heater in Upstream Level Can | W4C050 50V, 100 Watt Cartridge Heater | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 2, L1-01 | FL02B25 | LL_Upstream | Cryomagnetics 3D-031-033-PB-RAD-2K 12" Liquid Level Probe | 2 pr 20 ga | 16060-02-A | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | Spare | Jumper/ Potential for Interlock |  |  | E,F
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 2, L2-01 | FL02B25 | LL_Upstream | Cryomagnetics 3D-031-033-PB-RAD-2K 12" Liquid Level Probe | 2 pr 20 ga | MS3126F10-6S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | Spare | Jumper/ Potential for Interlock |  |  | E,F

## Connectors
Flange | 4 pin | 8 pin | 12 pin | 19 pin | "N" | "SMA" | Total
| Heater | Tuner limit | Fast Tuner | RTDs
1A |  |  | 1 | 2 |  |  | 3
1B | 1 | 1 |  |  | 3 | 2 | 7
2A |  |  | 1 | 2 |  |  | 3
2B | 1 | 1 |  |  | 3 | 2 | 7
3A |  |  | 1 | 2 |  |  | 3
3B | 1 | 1 |  |  | 3 | 2 | 7
4A |  |  | 1 | 2 |  |  | 3
4B | 1 | 1 |  |  | 3 | 2 | 7
5A |  |  | 1 | 2 |  |  | 3
5B | 1 | 1 |  |  | 3 | 2 | 7
6A |  |  | 1 | 2 |  |  | 3
6B | 1 | 1 |  |  | 3 | 2 | 7
7A |  |  | 1 | 2 |  |  | 3
7B | 1 | 1 |  |  | 3 | 2 | 7
8A |  |  | 1 | 2 |  |  | 3
8B | 1 | 1 |  |  | 3 | 2 | 7
| Quad Heater | Quad V Taps |  |  |  |  | 0
"C" | 1 | 3 |  |  |  |  | 4
| Cooldown Heaters |  |  | Flux Gate |  |  | 0
"D" | 2 |  |  | 2 | 4 |  | 8
|  | Liquid Level |  |  |  |  | 0
"E" |  | 1 |  |  |  |  | 1
|  |  |  | RTD |  |  | 0
"F" |  |  |  | 2 |  |  | 2
| Heater |  |  | RTD |  |  | 0
"K" | 1 |  |  | 3 |  |  | 4
|  | Liquid Level |  |  |  |  | 0
"L" |  | 1 |  |  |  |  | 1
Totals | 12 | 13 | 8 | 23 | 28 | 16 | 100
for 2 CMs | 24 | 26 | 16 | 26 | 56 | 32 | 200
Function | Manufacturer | Cable Spec | # of runs | Need | Have |  |  |  | PR / Checked
Heater | Alpha 2424C | 4cond 18 ga; | 24 | 3000 | 3000' on order |  | 2717.28 | Newark | 375187
PZT | Alpha 2463C | 4pr 22 ga | 16 | 2000 | 2000' on order
Stepper tuners | Alpha 2467C | 4pr 20 ga | 16 | 2000 | 2000' on order
RTDs - 3 & flux | Alpha 6063C | 6 pr 20 | 16 |  |  |  | on hand
RTDs - 4 |  | 8pr 22 ga
Voltage taps | Alpha 1298C | 8cond 22ga |  | 6 runs - 1000' | on order |  | 919.45 | Digikey
Pressure/Liquid Level |  | 2pr 20 ga | 8 | 1500 | 1000' on order, 500 on hand |  | 1849.72 | Newark | 375168