## Cable Listing
Cable number | From | To | Device | Cable | Connector
1 | FL02B25 | CM 2, 1A-03 | RTD_HeVessel | 8 pr 22 ga | MS3126F14-19S
2 | FL02B25 | CM 2, 1B-07 | StrpHtr-HeVessel | 4 cond 18 ga | MS24266R14B4SN
3 | FL02B25 | CM 2, 2A-03 | RTD_HOM | 4 pr 22 ga | MS3126F14-19S
4 | FL02B25 | CM 2, 2B-07 | StrpHtr-HeVessel | 4 cond 18 ga | MS24266R14B4SN
5 | FL02B25 | CM 2, 3A-03 | RTD_HOM | 4 pr 22 ga | MS3126F14-19S
6 | FL02B25 | CM 2, 3B-07 | StrpHtr-HeVessel | 4 cond 18 ga | MS24266R14B4SN
7 | FL02B25 | CM 2, 4A-03 | RTD_HOM | 4 pr 22 ga | MS3126F14-19S
8 | FL02B25 | CM 2, 4B-07 | StrpHtr-HeVessel | 4 cond 18 ga | MS24266R14B4SN
9 | FL02B25 | CM 2, 5A-03 | RTD_HeVessel | 8 pr 22 ga | MS3126F14-19S
10 | FL02B25 | CM 2, 5B-07 | StrpHtr-HeVessel | 4 cond 18 ga | MS24266R14B4SN
11 | FL02B25 | CM 2, 6A-03 | RTD_HOM | 4 pr 22 ga | MS3126F14-19S
12 | FL02B25 | CM 2, 6B-07 | StrpHtr-HeVessel | 4 cond 18 ga | MS24266R14B4SN
13 | FL02B25 | CM 2, 7A-03 | RTD_HOM | 4 pr 22 ga | MS3126F14-19S
14 | FL02B25 | CM 2, 7B-07 | StrpHtr-HeVessel | 4 cond 18 ga | MS24266R14B4SN
15 | FL02B25 | CM 2, 8A-03 | RTD_HOM | 4 pr 22 ga | MS3126F14-19S
16 | FL02B25 | CM 2, 8B-07 | StrpHtr-HeVessel-8 | 4 cond 18 ga | MS24266R14B4SN
17 | FL02B25 | CM 2, C-04 | Heater Q1 & Q2 | 4 cond 18 ga | MS24266R14B4SN
18 | FL02B25 | CM 2, D-07 | CoolDownHtr-1 | 4 cond 18 ga | MS24266R14B4SN
19 | FL02B25 | CM 2, D-08 | CoolDownHtr-3 | 4 cond 18 ga | MS24266R14B4SN
20 | FL02B25 | CM 2, E1-01 | LL_Downstream | 2 pr 20 ga | 16060-02-A
21 | FL02B25 | CM 2, E1-BAR | Baratron
22 | FL02B25 | CM 2, E2-01 | LL_Downstream | 2 pr 20 ga | MS3126F10-6S
23 | FL02B25 | CM 2, F-01 | RTD_MP_1 | 8 pr 22 ga | MS3126F14-19S
24 | FL02B25 | CM 2, F-02 | RTD_50K_Lead | 8 pr 22 ga | MS3126F14-19S
25 | FL02B25 | CM 2, K-01 | RTD_Line_A1 | 8 pr 22 ga | MS3126F14-19S
26 | FL02B25 | CM 2, K-02 | RTD_Line_D1 | 8 pr 22 ga | MS3126F14-19S
27 | FL02B25 | CM 2, K-03 | RTD_Cooldown_1 | 8 pr 22 ga | MS3126F14-19S
28 | FL02B25 | CM 2, K-04 | Heater Up Level | 4 cond 18 ga | MS24266R14B4SN
29 | FL02B25 | CM 2, L1-01 | LL_Upstream | 2 pr 20 ga | 16060-02-A
30 | FL02B25
31 | FL02B25 | CM 2, L2-01 | LL_Upstream | 2 pr 20 ga | MS3126F10-6S
32 | FL02B27 | CM 1, 1A-03 | RTD_HeVes | 8 pr 22 ga | MS3126F14-19S
33 | FL02B27 | CM 1, 1B-07 | StrpHtr-HeVes | 4 cond 18 ga | MS24266R14B4SN
34 | FL02B27 | CM 1, 2A-03 | RTD HOM CPL | 4 pr 22 ga | MS3126F14-19S
35 | FL02B27 | CM 1, 2B-07 | StrpHtr-HeVes | 4 cond 18 ga | MS24266R14B4SN
36 | FL02B27 | CM 1, 3A-03 | RTD HOM CPL | 4 pr 22 ga | MS3126F14-19S
37 | FL02B27 | CM 1, 3B-07 | StrpHtr-HeVes | 4 cond 18 ga | MS24266R14B4SN
38 | FL02B27 | CM 1, 4A-03 | RTD HOM CPL | 4 pr 22 ga | MS3126F14-19S
39 | FL02B27 | CM 1, 4B-07 | StrpHtr-HeVes | 4 cond 18 ga | MS24266R14B4SN
40 | FL02B27 | CM 1, 5A-03 | RTD_HeVes | 8 pr 22 ga | MS3126F14-19S
41 | FL02B27 | CM 1, 5B-07 | StrpHtr-HeVes | 4 cond 18 ga | MS24266R14B4SN
42 | FL02B27 | CM 1, 6A-03 | RTD HOM CPL | 4 pr 22 ga | MS3126F14-19S
43 | FL02B27 | CM 1, 6B-07 | StrpHtr-HeVes | 4 cond 18 ga | MS24266R14B4SN
44 | FL02B27 | CM 1, 7A-03 | RTD HOM CPL | 4 pr 22 ga | MS3126F14-19S
45 | FL02B27 | CM 1, 7B-07 | StrpHtr-HeVes | 4 cond 18 ga | MS24266R14B4SN
46 | FL02B27 | CM 1, 8A-03 | RTD HOM CPL | 4 pr 22 ga | MS3126F14-19S
47 | FL02B27 | CM 1, 8B-07 | StrpHtr-HeVes | 4 cond 18 ga | MS24266R14B4SN
48 | FL02B27 | CM 1, C-04 | Heater Q1 & Q2 | 4 cond 18 ga | MS24266R14B4SN
49 | FL02B27 | CM 1, D-07 | CoolDownHtr-1 | 4 cond 18 ga | MS24266R14B4SN
50 | FL02B27 | CM 1, D-08 | CoolDownHtr-3 | 4 cond 18 ga | MS24266R14B4SN
51 | FL02B27 | CM 1, E1-01 | LL_Downstream | 4 cond 18 ga | 16060-02-A
52 | FL02B27 |  | Baratron
53 | FL02B27 | CM 1, E2-01 | LL_Downstream | 4 cond 18 ga | MS3126F10-6S
54 | FL02B27 | CM 1, F-01 | RTD_MP_1 | 8 pr 22 ga | MS3126F14-19S
55 | FL02B27 | CM 1, F-02 | RTD_50K_LeadThermIntcp | 8 pr 22 ga | MS3126F14-19S
56 | FL02B27 | CM 1, K-01 | RTD_Line_A1 | 8 pr 22 ga | MS3126F14-19S
57 | FL02B27 | CM 1, K-02 | RTD_Line_D1 | 8 pr 22 ga | MS3126F14-19S
58 | FL02B27 | CM 1, K-03 | RTD_Cooldown_1 | 8 pr 22 ga | MS3126F14-19S
59 | FL02B27 | CM 1, K-04 | Heater in Upstream Level Can | 4 cond 18 ga | MS24266R14B4SN
60 | FL02B27 | CM 1, L1-01 | LL_Upstream | 4 cond 18 ga | 16060-02-A
61 | FL02B27 | CM 1, L1-BAR | Baratron
62 | FL02B27 | CM 1, L2-01 | LL_Upstream | 4 cond 18 ga | MS3126F10-6S
63 | FL03B24 | CM 1, 1B-03 | C1_H1 | 1/4" Heliax
64 | FL03B24 | CM 1, 1B-04 | C1_H2 | 1/4" Heliax
65 | FL03B24 | CM 1, 2B-03 | C2_H1 | 1/4" Heliax
66 | FL03B24 | CM 1, 2B-04 | C2_H2 | 1/4" Heliax
67 | FL03B24 | CM 1, 3B-03 | C3_H1 | 1/4" Heliax
68 | FL03B24 | CM 1, 3B-04 | C3_H2 | 1/4" Heliax
69 | FL03B24 | CM 1, 4B-03 | C4_H1 | 1/4" Heliax
70 | FL03B24 | CM 1, 4B-04 | C4_H2 | 1/4" Heliax
71 | FL03B24 | CM 1, 5B-03 | C5_H1 | 1/4" Heliax
72 | FL03B24 | CM 1, 5B-04 | C5_H2 | 1/4" Heliax
73 | FL03B24 | CM 1, 6B-03 | C6_H1 | 1/4" Heliax
74 | FL03B24 | CM 1, 6B-04 | C6_H2 | 1/4" Heliax
75 | FL03B24 | CM 1, 7B-03 | C7_H1 | 1/4" Heliax
76 | FL03B24 | CM 1, 7B-04 | C7_H2 | 1/4" Heliax
77 | FL03B24 | CM 1, 8B-03 | C8_H1 | 1/4" Heliax
78 | FL03B24 | CM 1, 8B-04 | C8_H2 | 1/4" Heliax
79 | FL03B24 | CM 2, 1B-03 | C1_H1 | 1/4" Heliax
80 | FL03B24 | CM 2, 1B-04 | C1_H2 | 1/4" Heliax
81 | FL03B24 | CM 2, 2B-03 | C2_H1 | 1/4" Heliax
82 | FL03B24 | CM 2, 2B-04 | C2_H2 | 1/4" Heliax
83 | FL03B24 | CM 2, 3B-03 | C3_H1 | 1/4" Heliax
84 | FL03B24 | CM 2, 3B-04 | C3_H2 | 1/4" Heliax
85 | FL03B24 | CM 2, 4B-03 | C4_H1 | 1/4" Heliax
86 | FL03B24 | CM 2, 4B-04 | C4_H2 | 1/4" Heliax
87 | FL03B24 | CM 2, 5B-03 | C5_H1 | 1/4" Heliax
88 | FL03B24 | CM 2, 5B-04 | C5_H2 | 1/4" Heliax
89 | FL03B24 | CM 2, 6B-03 | C6_H1 | 1/4" Heliax
90 | FL03B24 | CM 2, 6B-04 | C6_H2 | 1/4" Heliax
91 | FL03B24 | CM 2, 7B-03 | C7_H1 | 1/4" Heliax
92 | FL03B24 | CM 2, 7B-04 | C7_H2 | 1/4" Heliax
93 | FL03B24 | CM 2, 8B-03 | C8_H1 | 1/4" Heliax
94 | FL03B24 | CM 2, 8B-04 | C8_H2 | 1/4" Heliax
95 | FL03B25 | CM 1, 1A-01 | FastTuner-1-1 | 4 pr 22 ga | MS3126F14-12S
96 | FL03B25 | CM 1, 1A-02 | RTD_Step-1 | 6 pr 20 ga | MS3126F14-19S
97 | FL03B25 | CM 1, 1B-01 | SlowTuner-1 | 4 pr 20 ga | MS3126F16-8S
98 | FL03B25 | CM 1, 1B-02 | C1_Pt (PTrans) | 3/8" Heliax | N 0°
99 | FL03B25 | CM 1, 1B-05 | Cav_1_e-Cy | 3/8" Heliax | SMA 00
100 | FL03B25 | CM 1, 1B-06 | Cav_1_e-WG | 3/8" Heliax | SMA 00
101 | FL03B25 | CM 1, Cav 1-F | Cavity 1 CPLR FOR | 3/8" Heliax | N 0°
102 | FL03B25 | CM 1, Cav 1-R | Cavity 1 CPLR REV | 3/8" Heliax | N 0°
103 | FL03B25 | CM 1, 2A-01 | FastTuner-2-1 | 4 pr 22 ga | MS3126F14-12S
104 | FL03B25 | CM 1, 2A-02 | RTD_Step-2 | 6 pr 20 ga | MS3126F14-19S
105 | FL03B25 | CM 1, 2B-01 | SlowTuner-2 | 4 pr 20 ga | MS3126F16-8S
106 | FL03B25 | CM 1, 2B-02 | C2_Pt (PTrans) | 3/8" Heliax | N 0°
107 | FL03B25 | CM 1, 2B-05 | Cav_2_e-Cy | 3/8" Heliax | SMA 00
108 | FL03B25 | CM 1, 2B-06 | Cav_2_e-WG | 3/8" Heliax | SMA 00
109 | FL03B25 | CM 1, Cav 2-F | Cavity 2 CPLR FOR | 3/8" Heliax | N 0°
110 | FL03B25 | CM 1, Cav 2-R | Cavity 2 CPLR REV | 3/8" Heliax | N 0°
111 | FL03B25 | CM 1, 3A-01 | FastTuner-3-1 | 4 pr 22 ga | MS3126F14-12S
112 | FL03B25 | CM 1, 3A-02 | RTD_Step-3 | 6 pr 20 ga | MS3126F14-19S
113 | FL03B25 | CM 1, 3B-01 | SlowTuner-3 | 4 pr 20 ga | MS3126F16-8S
114 | FL03B25 | CM 1, 3B-02 | C3_Pt (PTrans) | 3/8" Heliax | N 0°
115 | FL03B25 | CM 1, 3B-05 | Cav_3_e-Cy | 3/8" Heliax | SMA 00
116 | FL03B25 | CM 1, 3B-06 | Cav_3_e-WG | 3/8" Heliax | SMA 00
117 | FL03B25 | CM 1, Cav 3-F | Cavity 3 CPLR FOR | 3/8" Heliax | N 0°
118 | FL03B25 | CM 1, Cav 3-R | Cavity 3 CPLR REV | 3/8" Heliax | N 0°
119 | FL03B25 | CM 1, 4A-01 | FastTuner-4-1 | 4 pr 22 ga | MS3126F14-12S
120 | FL03B25 | CM 1, 4A-02 | RTD_Step-4 | 6 pr 20 ga | MS3126F14-19S
121 | FL03B25 | CM 1, 4B-01 | SlowTuner-4 | 4 pr 20 ga | MS3126F16-8S
122 | FL03B25 | CM 1, 4B-02 | C4_Pt (PTrans) | 3/8" Heliax | N 0°
123 | FL03B25 | CM 1, 4B-05 | Cav_4_e-Cy | 3/8" Heliax | SMA 00
124 | FL03B25 | CM 1, 4B-06 | Cav_4_e-WG | 3/8" Heliax | SMA 00
125 | FL03B25 | CM 1, Cav 4-F | Cavity 4 CPLR FOR | 3/8" Heliax | N 0°
126 | FL03B25 | CM 1, Cav 4-R | Cavity 4 CPLR REV | 3/8" Heliax | N 0°
127 | FL03B34 | CM 1, 5A-01 | FastTuner-5-1 | 4 pr 22 ga | MS3126F14-12S
128 | FL03B34 | CM 1, 5A-02 | RTD_Step-5 | 6 pr 20 ga | MS3126F14-19S
129 | FL03B34 | CM 1, 5B-01 | SlowTuner-5 | 4 pr 20 ga | MS3126F16-8S
130 | FL03B34 | CM 1, 5B-02 | C5_Pt (PTrans) | 3/8" Heliax | N 0°
131 | FL03B34 | CM 1, 5B-05 | Cav_5_e-Cy | 3/8" Heliax | SMA 00
132 | FL03B34 | CM 1, 5B-06 | Cav_5_e-WG | 3/8" Heliax | SMA 00
133 | FL03B34 | CM 1, Cav 5-F | Cavity 5 CPLR FOR | 3/8" Heliax | N 0°
134 | FL03B34 | CM 1, Cav 5-R | Cavity 5 CPLR REV | 3/8" Heliax | N 0°
135 | FL03B34 | CM 1, 6A-01 | FastTuner-6-1 | 4 pr 22 ga | MS3126F14-12S
136 | FL03B34 | CM 1, 6A-02 | RTD_Step-6 | 6 pr 20 ga | MS3126F14-19S
137 | FL03B34 | CM 1, 6B-01 | SlowTuner-6 | 4 pr 20 ga | MS3126F16-8S
138 | FL03B34 | CM 1, 6B-02 | C6_Pt (PTrans) | 3/8" Heliax | N 0°
139 | FL03B34 | CM 1, 6B-05 | Cav_6_e-Cy | 3/8" Heliax | SMA 00
140 | FL03B34 | CM 1, 6B-06 | Cav_6_e-WG | 3/8" Heliax | SMA 00
141 | FL03B34 | CM 1, Cav 6-F | Cavity 6 CPLR FOR | 3/8" Heliax | N 0°
142 | FL03B34 | CM 1, Cav 6-R | Cavity 6 CPLR REV | 3/8" Heliax | N 0°
143 | FL03B34 | CM 1, 7A-01 | FastTuner-7-1 | 4 pr 22 ga | MS3126F14-12S
144 | FL03B34 | CM 1, 7A-02 | RTD_Step-7 | 6 pr 20 ga | MS3126F14-19S
145 | FL03B34 | CM 1, 7B-01 | SlowTuner-7 | 4 pr 20 ga | MS3126F16-8S
146 | FL03B34 | CM 1, 7B-02 | C7_Pt (PTrans) | 3/8" Heliax | N 0°
147 | FL03B34 | CM 1, 7B-05 | Cav_7_e-Cy | 3/8" Heliax | SMA 00
148 | FL03B34 | CM 1, 7B-06 | Cav_7_e-WG | 3/8" Heliax | SMA 00
149 | FL03B34 | CM 1, Cav 7-F | Cavity 7 CPLR FOR | 3/8" Heliax | N 0°
150 | FL03B34 | CM 1, Cav 7-R | Cavity 7 CPLR REV | 3/8" Heliax | N 0°
151 | FL03B34 | CM 1, 8A-01 | FastTuner-8-1 | 4 pr 22 ga | MS3126F14-12S
152 | FL03B34 | CM 1, 8A-02 | RTD_Step-8 | 6 pr 20 ga | MS3126F14-19S
153 | FL03B34 | CM 1, 8B-01 | SlowTuner-8 | 4 pr 20 ga | MS3126F16-8S
154 | FL03B34 | CM 1, 8B-02 | C8_Pt (PTrans) | 3/8" Heliax | N 0°
155 | FL03B34 | CM 1, 8B-05 | Cav_8_e-Cy | 3/8" Heliax | SMA 00
156 | FL03B34 | CM 1, 8B-06 | Cav_8_e-WG | 3/8" Heliax | SMA 00
157 | FL03B34 | CM 1, Cav 8-F | Cavity 8 CPLR FOR | 3/8" Heliax | N 0°
158 | FL03B34 | CM 1, Cav 8-R | Cavity 8 CPLR REV | 3/8" Heliax | N 0°
159 | FL04B25 | CM 2, 5A-01 | FastTuner-5-1 | 4 pr 22 ga | MS3126F14-12S
160 | FL04B25 | CM 2, 5A-02 | RTD_Step-5 | 6 pr 20 ga | MS3126F14-19S
161 | FL04B25 | CM 2, 5B-01 | SlowTuner-5 | 4 pr 20 ga | MS3126F16-8S
162 | FL04B25 | CM 2, 5B-02 | C5_Pt (PTrans) | 3/8" Heliax | N 0°
163 | FL04B25 | CM 2, 5B-05 | Cav_5_e-Cy | 3/8" Heliax | SMA 00
164 | FL04B25 | CM 2, 5B-06 | Cav_5_e-WG | 3/8" Heliax | SMA 00
165 | FL04B25 | CM 2, Cav 5-F | Cavity 5 CPLR FOR | 3/8" Heliax | N 0°
166 | FL04B25 | CM 2, Cav 5-R | Cavity 5 CPLR REV | 3/8" Heliax | N 0°
167 | FL04B25 | CM 2, 6A-01 | FastTuner-6-1 | 4 pr 22 ga | MS3126F14-12S
168 | FL04B25 | CM 2, 6A-02 | RTD_Step-6 | 6 pr 20 ga | MS3126F14-19S
169 | FL04B25 | CM 2, 6B-01 | SlowTuner-6 | 4 pr 20 ga | MS3126F16-8S
170 | FL04B25 | CM 2, 6B-02 | C6_Pt (PTrans) | 3/8" Heliax | N 0°
171 | FL04B25 | CM 2, 6B-05 | Cav_6_e-Cy | 3/8" Heliax | SMA 00
172 | FL04B25 | CM 2, 6B-06 | Cav_6_e-WG | 3/8" Heliax | SMA 00
173 | FL04B25 | CM 2, Cav 6-F | Cavity 6 CPLR FOR | 3/8" Heliax | N 0°
174 | FL04B25 | CM 6, Cav 6-R | Cavity 6 CPLR REV | 3/8" Heliax | N 0°
175 | FL04B25 | CM 2, 7A-01 | FastTuner-7-1 | 4 pr 22 ga | MS3126F14-12S
176 | FL04B25 | CM 2, 7A-02 | RTD_Step-7 | 6 pr 20 ga | MS3126F14-19S
177 | FL04B25 | CM 2, 7B-01 | SlowTuner-7 | 4 pr 20 ga | MS3126F16-8S
178 | FL04B25 | CM 2, 7B-02 | C7_Pt (PTrans) | 3/8" Heliax | N 0°
179 | FL04B25 | CM 2, 7B-05 | Cav_7_e-Cy | 3/8" Heliax | SMA 00
180 | FL04B25 | CM 2, 7B-06 | Cav_7_e-WG | 3/8" Heliax | SMA 00
181 | FL04B25 | CM 2, Cav 7-F | Cavity 7 CPLR FOR | 3/8" Heliax | N 0°
182 | FL04B25 | CM 2, Cav 7-R | Cavity 7 CPLR REV | 3/8" Heliax | N 0°
183 | FL04B25 | CM 2, 8A-01 | FastTuner-8-1 | 4 pr 22 ga | MS3126F14-12S
184 | FL04B25 | CM 2, 8A-02 | RTD_Step-8 | 6 pr 20 ga | MS3126F14-19S
185 | FL04B25 | CM 2, 8B-01 | SlowTuner-8 | 4 pr 20 ga | MS3126F16-8S
186 | FL04B25 | CM 2, 8B-02 | C8_Pt (PTrans) | 3/8" Heliax | N 0°
187 | FL04B25 | CM 2, 8B-05 | Cav_8_e-Cy | 3/8" Heliax | SMA 00
188 | FL04B25 | CM 2, 8B-06 | Cav_8_e-WG | 3/8" Heliax | SMA 00
189 | FL04B25 | CM 2, Cav 8-F | Cavity 8 CPLR FOR | 3/8" Heliax | N 0°
190 | FL04B25 | CM 2, Cav 8-R | Cavity 8 CPLR REV | 3/8" Heliax | N 0°
191 | FL04B34 | CM 2, 1A-01 | FastTuner-1-1 | 4 pr 22 ga | MS3126F14-12S
192 | FL04B34 | CM 2, 1A-02 | RTD_Step-1 | 6 pr 20 ga | MS3126F14-19S
193 | FL04B34 | CM 2, 1B-01 | SlowTuner-1 | 4 pr 20 ga | MS3126F16-8S
194 | FL04B34 | CM 2, 1B-02 | C1_Pt (PTrans) | 3/8" Heliax | N 0°
195 | FL04B34 | CM 2, 1B-05 | Cav_1_e-Cy | 3/8" Heliax | SMA 00
196 | FL04B34 | CM 2, 1B-06 | Cav_1_e-WG | 3/8" Heliax | SMA 00
197 | FL04B34 | CM 2, Cav 1-F | Cavity 1 CPLR FOR | 3/8" Heliax | N 0°
198 | FL04B34 | CM 2, Cav 1-R | Cavity 1 CPLR REV | 3/8" Heliax | N 0°
199 | FL04B34 | CM 2, 2A-01 | FastTuner-2-1 | 4 pr 22 ga | MS3126F14-12S
200 | FL04B34 | CM 2, 2A-02 | RTD_Step-2 | 6 pr 20 ga | MS3126F14-19S
201 | FL04B34 | CM 2, 2B-01 | SlowTuner-2 | 4 pr 20 ga | MS3126F16-8S
202 | FL04B34 | CM 2, 2B-02 | C2_Pt (PTrans) | 3/8" Heliax | N 0°
203 | FL04B34 | CM 2, 2B-05 | Cav_2_e-Cy | 3/8" Heliax | SMA 00
204 | FL04B34 | CM 2, 2B-06 | Cav_2_e-WG | 3/8" Heliax | SMA 00
205 | FL04B34 | CM 2, Cav 2-F | Cavity 2 CPLR FOR | 3/8" Heliax | N 0°
206 | FL04B34 | CM 2, Cav 2-R | Cavity 2 CPLR REV | 3/8" Heliax | N 0°
207 | FL04B34 | CM 2, 3A-01 | FastTuner-3-1 | 4 pr 22 ga | MS3126F14-12S
208 | FL04B34 | CM 2, 3A-02 | RTD_Step-3 | 6 pr 20 ga | MS3126F14-19S
209 | FL04B34 | CM 2, 3B-01 | SlowTuner-3 | 4 pr 20 ga | MS3126F16-8S
210 | FL04B34 | CM 2, 3B-02 | C3_Pt (PTrans) | 3/8" Heliax | N 0°
211 | FL04B34 | CM 2, 3B-05 | Cav_3_e-Cy | 3/8" Heliax | SMA 00
212 | FL04B34 | CM 2, 3B-06 | Cav_3_e-WG | 3/8" Heliax | SMA 00
213 | FL04B34 | CM 2, Cav 3-F | Cavity 3 CPLR FOR | 3/8" Heliax | N 0°
214 | FL04B34 | CM 2, Cav 3-R | Cavity 3 CPLR REV | 3/8" Heliax | N 0°
215 | FL04B34 | CM 2, 4A-01 | FastTuner-4-1 | 4 pr 22 ga | MS3126F14-12S
216 | FL04B34 | CM 2, 4A-02 | RTD_Step-4 | 6 pr 20 ga | MS3126F14-19S
217 | FL04B34 | CM 2, 4B-01 | SlowTuner-4 | 4 pr 20 ga | MS3126F16-8S
218 | FL04B34 | CM 2, 4B-02 | C4_Pt (PTrans) | 3/8" Heliax | N 0°
219 | FL04B34 | CM 2, 4B-05 | Cav_4_e-Cy | 3/8" Heliax | SMA 00
220 | FL04B34 | CM 2, 4B-06 | Cav_4_e-WG | 3/8" Heliax | SMA 00
221 | FL04B34 | CM 2, Cav 4-F | Cavity 4 CPLR FOR | 3/8" Heliax | N 0°
222 | FL04B34 | CM 2, Cav 4-R | Cavity 4 CPLR REV | 3/8" Heliax | N 0°
223
224
225
226
227
228
229
230
231
232
233
234
235
236
237
238
239
240
241
242
243
244
245
246
247
248
249
250
251
252
253
254
255
256
257
258
259
260
261
262
263
264
265
266
267
268
269
270
271
272
273
274
275
276
277
278
279
280
281
282
283
284
285
286
287
288
289
290
291
292
293
294
295
296
297
298
299
300

## Connectors
Flange | 4 pin | 8 pin | 12 pin | 19 pin | "N" | "SMA" | Total
| Heater | Tuner limit | Fast Tuner | RTDs
1A |  |  | 1 | 2 |  |  | 3
1B | 1 | 1 |  |  | 3 | 2 | 7
2A |  |  | 1 | 2 |  |  | 3
2B | 1 | 1 |  |  | 3 | 2 | 7
3A |  |  | 1 | 2 |  |  | 3
3B | 1 | 1 |  |  | 3 | 2 | 7
4A |  |  | 1 | 2 |  |  | 3
4B | 1 | 1 |  |  | 3 | 2 | 7
5A |  |  | 1 | 2 |  |  | 3
5B | 1 | 1 |  |  | 3 | 2 | 7
6A |  |  | 1 | 2 |  |  | 3
6B | 1 | 1 |  |  | 3 | 2 | 7
7A |  |  | 1 | 2 |  |  | 3
7B | 1 | 1 |  |  | 3 | 2 | 7
8A |  |  | 1 | 2 |  |  | 3
8B | 1 | 1 |  |  | 3 | 2 | 7
| Quad Heater | Quad V Taps |  |  |  |  | 0
"C" | 1 | 3 |  |  |  |  | 4
| Cooldown Heaters |  |  | Flux Gate |  |  | 0
"D" | 2 |  |  | 2 | 4 |  | 8
|  | Liquid Level |  |  |  |  | 0
"E" |  | 1 |  |  |  |  | 1
|  |  |  | RTD |  |  | 0
"F" |  |  |  | 2 |  |  | 2
| Heater |  |  | RTD |  |  | 0
"K" | 1 |  |  | 3 |  |  | 4
|  | Liquid Level |  |  |  |  | 0
"L" |  | 1 |  |  |  |  | 1
Totals | 12 | 13 | 8 | 23 | 28 | 16 | 100
for 2 CMs | 24 | 26 | 16 | 26 | 56 | 32 | 200
Function | Manufacturer | Cable Spec | # of runs | Need | Have |  |  |  | PR / Checked
Heater | Alpha 2424C | 4cond 18 ga; | 24 | 3000 | 3000' on order |  | 2717.28 | Newark | 375187
PZT | Alpha 2463C | 4pr 22 ga | 16 | 2000 | 2000' on order
Stepper tuners | Alpha 2467C | 4pr 20 ga | 16 | 2000 | 2000' on order
RTDs - 3 & flux | Alpha 6063C | 6 pr 20 | 16 |  |  |  | on hand
RTDs - 4 |  | 8pr 22 ga
Voltage taps | Alpha 1298C | 8cond 22ga |  | 6 runs - 1000' | on order |  | 919.45 | Digikey
Pressure/Liquid Level |  | 2pr 20 ga | 8 | 1500 | 1000' on order, 500 on hand |  | 1849.72 | Newark | 375168

## Summary
Connector | Rack Loc. | Device/Sensor Name | Device/Sensor Type | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 1, 1A-01 | FL03B25 | (Piezo Actuator) FastTuner-1-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-1-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
|  | (Piezo Actuator) FastTuner-1-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-1-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-1 | Piezo Fast Tuner, Interlock |  |  | J-K & L-M
CM 1, 1A-02 | FL03B25 | RTD_StepperMotor-1 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler1a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler1b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 1A-03 | FL02B27 | RTD_HeVessel-1 Top | Cernox CX-1030-SD-HT-1.4L | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_HeVessel-1 Bottom | Cernox CX-1030-SD-HT-1.4L |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-1-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-1-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 1B-01 | FL03B25 | (Stepper Motor) SlowTuner-1 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
|  | Tuner limit Switch-1-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-1-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 1, 1B-02 | FL03B25 | C1_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, 1B-03 | FL03B24 | C1_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 1B-04 | FL03B24 | C1_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 1B-05 | FL03B25 | Cav_1_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 1B-06 | FL03B25 | Cav_1_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 1B-07 | FL02B27 | StrpHtr-HeVessel-1 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 1, Cav 1-F | FL03B25 | Cavity 1 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, Cav 1-R | FL03B25 | Cavity 1 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
Connector |  | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 1, 2A-01 | FL03B25 | (Piezo Actuator) FastTuner-2-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-2-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
|  | (Piezo Actuator) FastTuner-2-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-2-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-2 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 1, 2A-02 | FL03B25 | RTD_StepperMotor-2 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler2a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler2b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 2A-03 | FL02B27 | Spare | N/A | 4 pr 22 ga | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-2-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-2-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 2B-01 | FL03B25 | (Stepper Motor) SlowTuner-2 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
|  | Tuner limit Switch-2-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-2-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 1, 2B-02 | FL03B25 | C2_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, 2B-03 | FL03B24 | C2_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 2B-04 | FL03B24 | C2_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 2B-05 | FL03B25 | Cav_2_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 2B-06 | FL03B25 | Cav_2_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 2B-07 | FL02B27 | StrpHtr-HeVessel-2 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 1, Cav 2-F | FL03B25 | Cavity 2 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, Cav 2-R | FL03B25 | Cavity 2 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
Connector |  | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 1, 3A-01 | FL03B25 | (Piezo Actuator) FastTuner-3-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-3-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
|  | (Piezo Actuator) FastTuner-3-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-3-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-3 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 1, 3A-02 | FL03B25 | RTD_StepperMotor-3 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler2a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler2b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 3A-03 | FL02B27 | Spare | N/A | 4 pr 22 ga | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-3-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-3-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 3B-01 | FL03B25 | (Stepper Motor) SlowTuner-3 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
|  | Tuner limit Switch-3-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-3-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 1, 3B-02 | FL03B25 | C3_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, 3B-03 | FL03B24 | C3_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 3B-04 | FL03B24 | C3_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 3B-05 | FL03B25 | Cav_3_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 3B-06 | FL03B25 | Cav_3_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 3B-07 | FL02B27 | StrpHtr-HeVessel-3 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 1, Cav 3-F | FL03B25 | Cavity 3 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, Cav 3-R | FL03B25 | Cavity 3 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
Connector |  | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 1, 4A-01 | FL03B25 | (Piezo Actuator) FastTuner-4-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-4-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
|  | (Piezo Actuator) FastTuner-4-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-4-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-4 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 1, 4A-02 | FL03B25 | RTD_StepperMotor-4 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler4a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler4b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 4A-03 | FL02B27 | Spare | N/A | 4 pr 22 ga | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-4-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-4-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 4B-01 | FL03B25 | (Stepper Motor) SlowTuner-4 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
|  | Tuner limit Switch-4-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-4-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 1, 4B-02 | FL03B25 | C4_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, 4B-03 | FL03B24 | C4_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 4B-04 | FL03B24 | C4_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 4B-05 | FL03B25 | Cav_4_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 4B-06 | FL03B25 | Cav_4_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 4B-07 | FL02B27 | StrpHtr-HeVessel-4 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 1, Cav 4-F | FL03B25 | Cavity 4 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, Cav 4-R | FL03B25 | Cavity 4 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
Connector |  | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 1, 5A-01 |  | (Piezo Actuator) FastTuner-5-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-5-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
| FL03B34 | (Piezo Actuator) FastTuner-5-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-5-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-5 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 1, 5A-02 |  | RTD_StepperMotor-5 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL03B34 | RTD_MainCoupler5a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler5b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 5A-03 |  | RTD_HeVessel-5 Top | Cernox CX-1030-SD-HT-1.4L | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL02B27 | RTD_HeVessel-5 Bottom | Cernox CX-1030-SD-HT-1.4L |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-5-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-5-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 5B-01 |  | (Stepper Motor) SlowTuner-5 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
| FL03B34 | Tuner limit Switch-5-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-5-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 1, 5B-02 | FL03B34 | C5_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, 5B-03 | FL03B24 | C5_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 5B-04 | FL03B24 | C5_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 5B-05 | FL03B34 | Cav_5_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 5B-06 | FL03B34 | Cav_5_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 5B-07 | FL02B27 | StrpHtr-HeVessel-5 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 1, Cav 5-F | FL03B34 | Cavity 5 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, Cav 5-R | FL03B34 | Cavity 5 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
Connector |  | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 1, 6A-01 |  | (Piezo Actuator) FastTuner-6-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-6-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
| FL03B34 | (Piezo Actuator) FastTuner-6-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-6-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-6 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 1, 6A-02 |  | RTD_StepperMotor-6 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL03B34 | RTD_MainCoupler6a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler6b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 6A-03 |  | Spare | N/A | 4 pr 22 ga | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL02B27 | Spare | N/A |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-6-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-6-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 6B-01 |  | (Stepper Motor) SlowTuner-6 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
| FL03B34 | Tuner limit Switch-6-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-6-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 1, 6B-02 | FL03B34 | C6_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, 6B-03 | FL03B24 | C6_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 6B-04 | FL03B24 | C6_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 6B-05 | FL03B34 | Cav_6_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 6B-06 | FL03B34 | Cav_6_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 6B-07 | FL02B27 | StrpHtr-HeVessel-6 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 1, Cav 6-F | FL03B34 | Cavity 6 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, Cav 6-R | FL03B34 | Cavity 6 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
Connector |  | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 1, 7A-01 |  | (Piezo Actuator) FastTuner-7-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-7-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
| FL03B34 | (Piezo Actuator) FastTuner-7-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-7-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-7 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 1, 7A-02 |  | RTD_StepperMotor-7 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL03B34 | RTD_MainCoupler7a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler7b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 7A-03 |  | Spare | N/A | 4 pr 22 ga | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL02B27 | Spare | N/A |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-7-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-7-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 7B-01 |  | (Stepper Motor) SlowTuner-7 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
| FL03B34 | Tuner limit Switch-7-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-7-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 1, 7B-02 | FL03B34 | C7_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, 7B-03 | FL03B24 | C7_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 7B-04 | FL03B24 | C7_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 7B-05 | FL03B34 | Cav_7_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 7B-06 | FL03B34 | Cav_7_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 7B-07 | FL02B27 | StrpHtr-HeVessel-7 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 1, Cav 7-F | FL03B34 | Cavity 7 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, Cav 7-R | FL03B34 | Cavity 7 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 1, 8A-01 |  | (Piezo Actuator) FastTuner-8-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-8-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
| FL03B34 | (Piezo Actuator) FastTuner-8-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-8-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-8 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 1, 8A-02 |  | RTD_StepperMotor-8 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL03B34 | RTD_MainCoupler8a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler8b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 8A-03 |  | Spare | N/A | 4 pr 22 ga | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL02B27 | Spare | N/A |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-8-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-8-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 8B-01 |  | (Stepper Motor) SlowTuner-8 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
| FL03B34 | Tuner limit Switch-8-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-8-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 1, 8B-02 | FL03B34 | C8_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, 8B-03 | FL03B24 | C8_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 8B-04 | FL03B24 | C8_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 8B-05 | FL03B34 | Cav_8_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 8B-06 | FL03B34 | Cav_8_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 8B-07 | FL02B27 | StrpHtr-HeVessel-8 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 1, Cav 8-F | FL03B34 | Cavity 8 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, Cav 8-R | FL03B34 | Cavity 8 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection (Warm Side) | Feed Through Connector Pin
CM 1, C-01 |  | VD VT Positive Lead (Whole_Coil-VT6) | Vertical Dipole Coil Voltage Tap (+) | 8cond 22ga Alpha 1298C | MS3126F16-8S | Pin-A - Black
|  | VD VT Negative Lead (Whole_Coil-VT5) | Vertical Dipole Coil Voltage Tap (-) |  |  | Pin-B - Brown
|  | HD VT Positive Lead (Whole_Coil-VT4) | Horizontal Dipole Coil Voltage Tap (+) |  |  | Pin-C - Red
|  | HD VT Negative Lead (Whole_Coil-VT3) | Horizontal Dipole Coil Voltage Tap (-) |  |  | Pin-D - Orange
|  | Quad VT Positive Lead (Whole_Coil-VT1) | Split Quadrupole Coil Voltage Tap (+) |  |  | Pin-E - Yellow
|  | Quad VT Negative Lead (Whole_Coil-VT2) | Split Quadrupole Coil Voltage Tap (-) |  |  | Pin-F - Green
|  | Spare | Spare |  |  | Pin-G - Blue
|  | Spare | Spare |  |  | Pin-H - White
CM 1, C-02 |  | Vertical Dipole VTRes28 (PosHalf Coil) | Vertical Dipole Coil Voltage Tap (+) | 8cond 22ga Alpha 1298C | MS3126F16-8S | Pin-A - Black
|  | Vertical Dipole VTRes29 (PosHalf Coil) | Vertical Dipole Coil Voltage Tap (-) |  |  | Pin-B - Brown
|  | Vertical Dipole VTRes30 (NegHalf Coil) | Vertical Dipole Coil Voltage Tap (+) |  |  | Pin-C - Red
|  | Vertical Dipole VTRes31 (NegHalf Coil) | Vertical Dipole Coil Voltage Tap (-) |  |  | Pin-D - Orange
|  | Horizontal Dipole VTRes22 (PosHalf Coil) | Horizontal Dipole Coil Voltage Tap (+) |  |  | Pin-E - Yellow
|  | Horizontal Dipole VTRes23 (PosHalf Coil) | Horizontal Dipole Coil Voltage Tap (-) |  |  | Pin-F - Green
|  | Horizontal Dipole VTRes24 (NegHalf Coil) | Horizontal Dipole Coil Voltage Tap (+) |  |  | Pin-G - Blue
|  | Horizontal Dipole VTRes25 (NegHalf Coil) | Horizontal Dipole Coil Voltage Tap (-) |  |  | Pin-H - White
CM 1, C-03 |  | Quadrupole VTRes16 (PosHalf Coil) | Split Quadrupole Coil Voltage Tap (+) | 8cond 22ga Alpha 1298C | MS3126F16-8S | Pin-A - Black
|  | Quadrupole VTRes17 (PosHalf Coil) | Split Quadrupole Coil Voltage Tap (-) |  |  | Pin-B - Brown
|  | Quadrupole VTRes18 (NegHalf Coil) | Split Quadrupole Coil Voltage Tap (+) |  |  | Pin-C - Red
|  | Quadrupole VTRes19 (NegHalf Coil) | Split Quadrupole Coil Voltage Tap (-) |  |  | Pin-D - Orange
|  | Spare | Spare |  |  | Pin-E - Yellow
|  | Spare | Spare |  |  | Pin-F - Green
|  | Spare | Spare |  |  | Pin-G - Blue
|  | Spare | Spare |  |  | Pin-H - White
CM 1, C-04 | FL02B27 | Heater Q1 & Q2 | Split Quadrupole and Corrector Coils Persistant Current Magnet Heaters | 4 cond 18 ga Alpha 2424C | MS24266R14B4SN | 1(+), 2(-)
|  | Heater Q3 & Q4 |  |  |  | 3(+), 4(-)
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 1, D-01 |  | BPM-Top | Beam Position Monitor |  | N 00 | Single Pin,Shield
CM 1, D-02 |  | BPM-Right | Beam Position Monitor |  | N 00 | Single Pin,Shield
CM 1, D-03 |  | BPM-Bot | Beam Position Monitor |  | N 00 | Single Pin,Shield
CM 1, D-04 |  | BPM-Left | Beam Position Monitor |  | N 00 | Single Pin,Shield
CM 1, D-05 |  | Fluxgate_Cav1 | Fluxgate | 6pr 20ga | MS3126F14-19S | A,B,C,D (S+,S-,E+,E-)
|  | Fluxgate_Cav2 | Fluxgate |  |  | E,F,G,H (S+,S-,E+,E-)
|  | Fluxgate_Cav5 | Fluxgate |  |  | J,K,L,M (S+,S-,E+,E-)
|  | spare | Spare |  |  | N,P,R,S
CM 1, D-06 |  | Fluxgate_Cav7 | Fluxgate | 6pr 20ga | MS3126F14-19S | A,B,C,D (S+,S-,E+,E-)
|  | Fluxgate_Cav8 | Fluxgate |  |  | E,F,G,H (S+,S-,E+,E-)
|  | spare | Spare |  |  | J,K,L,M
|  | spare | Spare |  |  | N,P,R,S
CM 1, D-07 | FL02B27 | CoolDownHtr-1 | Primary Cool Down Heater #1 Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) , B(-)
|  | CoolDownHtr-2 | Primary Cool Down Heater #2 Kapton Polymide Strip Heater, 46.3 Ω |  |  | C(+) , D(-)
CM 1, D-08 | FL02B27 | CoolDownHtr-3 | Primary Cool Down Heater #3 Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) , B(-)
|  | Heater in Downstream Level Can | W4C050 50V, 100 Watt Cartridge Heater |  |  | C(+) , D(-)
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 1, E1-01 | FL02B27 | LL_Downstream | Cryomagnetics 3D-031-033-PB-RAD-2K 12" Liquid Level Probe | 4 cond 18 ga; Alpha 2424C | 16060-02-A | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | Spare | Jumper/ Potential for Interlock |  |  | E,F
| FL02B27 | Baratron
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 1, E2-01 | FL02B27 | LL_Downstream | Cryomagnetics 3D-031-033-PB-RAD-2K 12" Liquid Level Probe | 4 cond 18 ga; Alpha 2424C | MS3126F10-6S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | Spare | Jumper/ Potential for Interlock |  |  | E,F
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 1, F-01 | FL02B27 | RTD_MP_1 | Cernox CX-1030-SD-HT-1.4L | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_MP_2 | DT-670A1-SD |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MP_3 | DT-670A1-SD |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_MP_4 | DT-670A1-SD |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, F-02 | FL02B27 | RTD_50K_LeadThermIntcp | DT-670A1-SD | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_5K_LeadThermIntcp | DT-670A1-SD |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_2K_LeadThermIntcp_1 | DT-670A1-SD |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_2K_LeadThermIntcp_2 | DT-670A1-SD |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 1, K-01 | FL02B27 | RTD_Line_A1 | DT-670A1-CU | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_Line_B1 | DT-670A1-CU |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_Line_B2 | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_Line_C1 | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, K-02 | FL02B27 | RTD_Line_D1 | DT-670A1-CU | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_Line_E1 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_Line_F1 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_Shield_1 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, K-03 | FL02B27 | RTD_Cooldown_1 | Cernox CX-1030-CU-HT-1.4L | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_Cooldown_2 | Cernox CX-1030-CU-HT-1.4L |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_Faraday Window 1 Port | Cernox CX-1030-SD-HT |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_Faraday Window 2 Port | Cernox CX-1030-SD-HT |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, K-04 | FL02B27 | Heater in Upstream Level Can | W4C050 50V, 100 Watt Cartridge Heater | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 1, L1-01 | FL02B27 | LL_Upstream | Cryomagnetics 3D-031-033-PB-RAD-2K 12" Liquid Level Probe | 4 cond 18 ga; Alpha 2424C | 16060-02-A | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | Spare | Jumper/ Potential for Interlock |  |  | E, F
CM 1, L1-BAR | FL02B27 | Baratron
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
L2-01 | FL02B27 | LL_Upstream | Cryomagnetics 3D-031-033-PB-RAD-2K 12" Liquid Level Probe | 4 cond 18 ga; Alpha 2424C | MS3126F10-6S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | Spare | Jumper/ Potential for Interlock |  |  | E,F
Connector | Rack | Device/Sensor Name | Device/Sensor Type | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 2, 1A-01 |  | (Piezo Actuator) FastTuner-1-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-1-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
| FL04B34 | (Piezo Actuator) FastTuner-1-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-1-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-1 | Piezo Fast Tuner, Interlock |  |  | J-K & L-M
CM 2, 1A-02 |  | RTD_StepperMotor-1 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL04B34 | RTD_MainCoupler1a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler1b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 1A-03 |  | RTD_HeVessel-1 Top | Cernox CX-1030-SD-HT-1.4L | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL02B27 | RTD_HeVessel-1 Bottom | Cernox CX-1030-SD-HT-1.4L |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-1-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-1-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 1B-01 |  | (Stepper Motor) SlowTuner-1 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
| FL04B34 | Tuner limit Switch-1-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-1-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 2, 1B-02 | FL04B34 | C1_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, 1B-03 | FL03B24 | C1_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 1B-04 | FL03B24 | C1_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 1B-05 | FL04B34 | Cav_1_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 1B-06 | FL04B34 | Cav_1_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 1B-07 | FL02B27 | StrpHtr-HeVessel-1 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 2, Cav 1-F | FL04B34 | Cavity 1 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, Cav 1-R | FL04B34 | Cavity 1 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 2, 2A-01 |  | (Piezo Actuator) FastTuner-2-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-2-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
| FL04B34 | (Piezo Actuator) FastTuner-2-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-2-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-2 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 2, 2A-02 |  | RTD_StepperMotor-2 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL04B34 | RTD_MainCoupler2a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler2b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 2A-03 |  | Spare | N/A | 4 pr 22 ga | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL02B27 | Spare | N/A |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-2-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-2-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 2B-01 |  | (Stepper Motor) SlowTuner-2 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
| FL04B34 | Tuner limit Switch-2-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-2-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 2, 2B-02 | FL04B34 | C2_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, 2B-03 | FL03B24 | C2_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 2B-04 | FL03B24 | C2_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 2B-05 | FL04B34 | Cav_2_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 2B-06 | FL04B34 | Cav_2_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 2B-07 | FL02B27 | StrpHtr-HeVessel-2 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 2, Cav 2-F | FL04B34 | Cavity 2 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, Cav 2-R | FL04B34 | Cavity 2 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 2, 3A-01 |  | (Piezo Actuator) FastTuner-3-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-3-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
| FL04B34 | (Piezo Actuator) FastTuner-3-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-3-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-3 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 2, 3A-02 |  | RTD_StepperMotor-3 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL04B34 | RTD_MainCoupler2a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler2b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 3A-03 |  | Spare | N/A | 4 pr 22 ga | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL02B27 | Spare | N/A |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-3-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-3-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 3B-01 |  | (Stepper Motor) SlowTuner-3 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
| FL04B34 | Tuner limit Switch-3-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-3-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 2, 3B-02 | FL04B34 | C3_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, 3B-03 | FL03B24 | C3_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 3B-04 | FL03B24 | C3_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 3B-05 | FL04B34 | Cav_3_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 3B-06 | FL04B34 | Cav_3_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 3B-07 | FL02B27 | StrpHtr-HeVessel-3 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 2, Cav 3-F | FL04B34 | Cavity 3 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, Cav 3-R | FL04B34 | Cavity 3 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 2, 4A-01 |  | (Piezo Actuator) FastTuner-4-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-4-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
| FL04B34 | (Piezo Actuator) FastTuner-4-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-4-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-4 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 2, 4A-02 |  | RTD_StepperMotor-4 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL04B34 | RTD_MainCoupler4a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler4b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 4A-03 |  | Spare | N/A | 4 pr 22 ga | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL02B27 | Spare | N/A |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-4-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-4-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 4B-01 |  | (Stepper Motor) SlowTuner-4 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
| FL04B34 | Tuner limit Switch-4-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-4-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 2, 4B-02 | FL04B34 | C4_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, 4B-03 | FL03B24 | C4_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 4B-04 | FL03B24 | C4_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 4B-05 | FL04B34 | Cav_4_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 4B-06 | FL04B34 | Cav_4_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 4B-07 | FL02B27 | StrpHtr-HeVessel-4 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 2, Cav 4-F | FL04B34 | Cavity 4 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, Cav 4-R | FL04B34 | Cavity 4 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 2, 5A-01 |  | (Piezo Actuator) FastTuner-5-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-5-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
| FL04B25 | (Piezo Actuator) FastTuner-5-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-5-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-5 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 2, 5A-02 |  | RTD_StepperMotor-5 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL04B25 | RTD_MainCoupler5a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler5b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 5A-03 |  | RTD_HeVessel-5 Top | Cernox CX-1030-SD-HT-1.4L | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL02B25 | RTD_HeVessel-5 Bottom | Cernox CX-1030-SD-HT-1.4L |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-5-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-5-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 5B-01 |  | (Stepper Motor) SlowTuner-5 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
| FL04B25 | Tuner limit Switch-5-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-5-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 2, 5B-02 | FL04B25 | C5_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, 5B-03 | FL03B24 | C5_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 5B-04 | FL03B24 | C5_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 5B-05 | FL04B25 | Cav_5_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 5B-06 | FL04B25 | Cav_5_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 5B-07 | FL02B25 | StrpHtr-HeVessel-5 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 2, Cav 5-F | FL04B25 | Cavity 5 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, Cav 5-R | FL04B25 | Cavity 5 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 2, 6A-01 |  | (Piezo Actuator) FastTuner-6-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-6-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
| FL04B25 | (Piezo Actuator) FastTuner-6-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-6-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-6 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 2, 6A-02 |  | RTD_StepperMotor-6 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL04B25 | RTD_MainCoupler6a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler6b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 6A-03 |  | Spare | N/A | 4 pr 22 ga | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL02B25 | Spare | N/A |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-6-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-6-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 6B-01 |  | (Stepper Motor) SlowTuner-6 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
| FL04B25 | Tuner limit Switch-6-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-6-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 2, 6B-02 | FL04B25 | C6_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, 6B-03 | FL03B24 | C6_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 6B-04 | FL03B24 | C6_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 6B-05 | FL04B25 | Cav_6_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 6B-06 | FL04B25 | Cav_6_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 6B-07 | FL02B25 | StrpHtr-HeVessel-6 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 2, Cav 6-F | FL04B25 | Cavity 6 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 6, Cav 6-R | FL04B25 | Cavity 6 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 2, 7A-01 |  | (Piezo Actuator) FastTuner-7-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-7-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
| FL04B25 | (Piezo Actuator) FastTuner-7-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-7-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-7 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 2, 7A-02 |  | RTD_StepperMotor-7 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL04B25 | RTD_MainCoupler7a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler7b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 7A-03 |  | Spare | N/A | 4 pr 22 ga | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL02B25 | Spare | N/A |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-7-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-7-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 7B-01 |  | (Stepper Motor) SlowTuner-7 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
| FL04B25 | Tuner limit Switch-7-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-7-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 2, 7B-02 | FL04B25 | C7_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, 7B-03 | FL03B24 | C7_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 7B-04 | FL03B24 | C7_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 7B-05 | FL04B25 | Cav_7_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 7B-06 | FL04B25 | Cav_7_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 7B-07 | FL02B25 | StrpHtr-HeVessel-7 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 2, Cav 7-F | FL04B25 | Cavity 7 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, Cav 7-R | FL04B25 | Cavity 7 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 2, 8A-01 |  | (Piezo Actuator) FastTuner-8-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-8-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
| FL04B25 | (Piezo Actuator) FastTuner-8-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-8-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-8 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 2, 8A-02 |  | RTD_StepperMotor-8 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL04B25 | RTD_MainCoupler8a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler8b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 8A-03 |  | Spare | N/A | 4 pr 22 ga | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL02B25 | Spare | N/A |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-8-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-8-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 8B-01 |  | (Stepper Motor) SlowTuner-8 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
| FL04B25 | Tuner limit Switch-8-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-8-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 2, 8B-02 | FL04B25 | C8_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, 8B-03 | FL03B24 | C8_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 8B-04 | FL03B24 | C8_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 8B-05 | FL04B25 | Cav_8_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 8B-06 | FL04B25 | Cav_8_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 8B-07 | FL02B25 | StrpHtr-HeVessel-8 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 2, Cav 8-F | FL04B25 | Cavity 8 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, Cav 8-R | FL04B25 | Cavity 8 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection (Warm Side) | Feed Through Connector Pin
CM 2, C-01 |  | VD VT Positive Lead (Whole_Coil-VT6) | Vertical Dipole Coil Voltage Tap (+) | 8cond 22ga Alpha 1298C | MS3126F16-8S | Pin-A - Black
|  | VD VT Negative Lead (Whole_Coil-VT5) | Vertical Dipole Coil Voltage Tap (-) |  |  | Pin-B - Brown
|  | HD VT Positive Lead (Whole_Coil-VT4) | Horizontal Dipole Coil Voltage Tap (+) |  |  | Pin-C - Red
|  | HD VT Negative Lead (Whole_Coil-VT3) | Horizontal Dipole Coil Voltage Tap (-) |  |  | Pin-D - Orange
|  | Quad VT Positive Lead (Whole_Coil-VT1) | Split Quadrupole Coil Voltage Tap (+) |  |  | Pin-E - Yellow
|  | Quad VT Negative Lead (Whole_Coil-VT2) | Split Quadrupole Coil Voltage Tap (-) |  |  | Pin-F - Green
|  | Spare | Spare |  |  | Pin-G - Blue
|  | Spare | Spare |  |  | Pin-H - White
CM 2, C-02 |  | Vertical Dipole VTRes28 (PosHalf Coil) | Vertical Dipole Coil Voltage Tap (+) | 8cond 22ga Alpha 1298C | MS3126F16-8S | Pin-A - Black
|  | Vertical Dipole VTRes29 (PosHalf Coil) | Vertical Dipole Coil Voltage Tap (-) |  |  | Pin-B - Brown
|  | Vertical Dipole VTRes30 (NegHalf Coil) | Vertical Dipole Coil Voltage Tap (+) |  |  | Pin-C - Red
|  | Vertical Dipole VTRes31 (NegHalf Coil) | Vertical Dipole Coil Voltage Tap (-) |  |  | Pin-D - Orange
|  | Horizontal Dipole VTRes22 (PosHalf Coil) | Horizontal Dipole Coil Voltage Tap (+) |  |  | Pin-E - Yellow
|  | Horizontal Dipole VTRes23 (PosHalf Coil) | Horizontal Dipole Coil Voltage Tap (-) |  |  | Pin-F - Green
|  | Horizontal Dipole VTRes24 (NegHalf Coil) | Horizontal Dipole Coil Voltage Tap (+) |  |  | Pin-G - Blue
|  | Horizontal Dipole VTRes25 (NegHalf Coil) | Horizontal Dipole Coil Voltage Tap (-) |  |  | Pin-H - White
CM 2, C-03 |  | Quadrupole VTRes16 (PosHalf Coil) | Split Quadrupole Coil Voltage Tap (+) | 8cond 22ga Alpha 1298C | MS3126F16-8S | Pin-A - Black
|  | Quadrupole VTRes17 (PosHalf Coil) | Split Quadrupole Coil Voltage Tap (-) |  |  | Pin-B - Brown
|  | Quadrupole VTRes18 (NegHalf Coil) | Split Quadrupole Coil Voltage Tap (+) |  |  | Pin-C - Red
|  | Quadrupole VTRes19 (NegHalf Coil) | Split Quadrupole Coil Voltage Tap (-) |  |  | Pin-D - Orange
|  | Spare | Spare |  |  | Pin-E - Yellow
|  | Spare | Spare |  |  | Pin-F - Green
|  | Spare | Spare |  |  | Pin-G - Blue
|  | Spare | Spare |  |  | Pin-H - White
CM 2, C-04 |  | Heater Q1 & Q2 | Split Quadrupole and Corrector Coils Persistant Current Magnet Heaters | 4 cond 18 ga Alpha 2424C | MS24266R14B4SN | 1(+), 2(-)
|  | Heater Q3 & Q4 |  |  |  | 3(+), 4(-)
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 2, D-01 |  | BPM-Top | Beam Position Monitor |  | N 00 | Single Pin,Shield
CM 2, D-02 |  | BPM-Right | Beam Position Monitor |  | N 00 | Single Pin,Shield
CM 2, D-03 |  | BPM-Bot | Beam Position Monitor |  | N 00 | Single Pin,Shield
CM 2, D-04 |  | BPM-Left | Beam Position Monitor |  | N 00 | Single Pin,Shield
CM 2, D-05 |  | Fluxgate_Cav1 | Fluxgate |  | MS3126F14-19S | A,B,C,D (S+,S-,E+,E-)
|  | Fluxgate_Cav2 | Fluxgate |  |  | E,F,G,H (S+,S-,E+,E-)
|  | Fluxgate_Cav5 | Fluxgate |  |  | J,K,L,M (S+,S-,E+,E-)
|  | spare | Spare |  |  | N,P,R,S
CM 2, D-06 |  | Fluxgate_Cav7 | Fluxgate |  | MS3126F14-19S | A,B,C,D (S+,S-,E+,E-)
|  | Fluxgate_Cav8 | Fluxgate |  |  | E,F,G,H (S+,S-,E+,E-)
|  | spare | Spare |  |  | J,K,L,M
|  | spare | Spare |  |  | N,P,R,S
CM 2, D-07 | FL02B25 | CoolDownHtr-1 | Primary Cool Down Heater #1 Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) , B(-)
|  | CoolDownHtr-2 | Primary Cool Down Heater #2 Kapton Polymide Strip Heater, 46.3 Ω |  |  | C(+) , D(-)
CM 2, D-08 | FL02B25 | CoolDownHtr-3 | Primary Cool Down Heater #3 Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) , B(-)
|  | Heater in Downstream Level Can | W4C050 50V, 100 Watt Cartridge Heater |  |  | C(+) , D(-)
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 2, E1-01 | FL02B25 | LL_Downstream | Cryomagnetics 3D-031-033-PB-RAD-2K 12" Liquid Level Probe | 2 pr 20 ga | 16060-02-A | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | Spare | Jumper/ Potential for Interlock |  |  | E,F
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 2, E2-01 | FL02B25 | LL_Downstream | Cryomagnetics 3D-031-033-PB-RAD-2K 12" Liquid Level Probe | 2 pr 20 ga | MS3126F10-6S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | Spare | Jumper/ Potential for Interlock |  |  | E,F
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 2, F-01 | FL02B25 | RTD_MP_1 | Cernox CX-1030-SD-HT-1.4L | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_MP_2 | DT-670A1-SD |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MP_3 | DT-670A1-SD |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_MP_4 | DT-670A1-SD |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, F-02 | FL02B25 | RTD_50K_LeadThermIntcp | DT-670A1-SD | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_5K_LeadThermIntcp | DT-670A1-SD |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_2K_LeadThermIntcp_1 | DT-670A1-SD |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_2K_LeadThermIntcp_2 | DT-670A1-SD |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 2, K-01 | FL02B25 | RTD_Line_A1 | DT-670A1-CU | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_Line_B1 | DT-670A1-CU |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_Line_B2 | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_Line_C1 | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, K-02 | FL02B25 | RTD_Line_D1 | DT-670A1-CU | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_Line_E1 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_Line_F1 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_Shield_1 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, K-03 | FL02B25 | RTD_Cooldown_1 | Cernox CX-1030-CU-HT-1.4L | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_Cooldown_2 | Cernox CX-1030-CU-HT-1.4L |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_Faraday Window 1 Port | Cernox CX-1030-SD-HT |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_Faraday Window 2 Port | Cernox CX-1030-SD-HT |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, K-04 | FL02B25 | Heater in Upstream Level Can | W4C050 50V, 100 Watt Cartridge Heater | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 2, L1-01 | FL02B25 | LL_Upstream | Cryomagnetics 3D-031-033-PB-RAD-2K 12" Liquid Level Probe | 2 pr 20 ga | 16060-02-A | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | Spare | Jumper/ Potential for Interlock |  |  | E,F
Connector | Rack | Device/Sensor Name | Device/Sensor Type/Serial Number | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 2, L2-01 | FL02B25 | LL_Upstream | Cryomagnetics 3D-031-033-PB-RAD-2K 12" Liquid Level Probe | 2 pr 20 ga | MS3126F10-6S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | Spare | Jumper/ Potential for Interlock |  |  | E,F

## FL02B25
Connector | Rack Loc. | Device/Sensor Name | Device/Sensor Type | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 2, 1A-03 |  | RTD_HeVessel-1 Top | Cernox CX-1030-SD-HT-1.4L | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL02B25 | RTD_HeVessel-1 Bottom | Cernox CX-1030-SD-HT-1.4L |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-1-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-1-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 1B-07 | FL02B25 | StrpHtr-HeVessel-1 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 2, 2A-03 |  | Spare | N/A | 4 pr 22 ga | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL02B25 | Spare | N/A |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-2-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-2-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 2B-07 | FL02B25 | StrpHtr-HeVessel-2 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 2, 3A-03 |  | Spare | N/A | 4 pr 22 ga | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL02B25 | Spare | N/A |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-3-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-3-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 3B-07 | FL02B25 | StrpHtr-HeVessel-3 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 2, 4A-03 |  | Spare | N/A | 4 pr 22 ga | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL02B25 | Spare | N/A |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-4-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-4-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 4B-07 | FL02B25 | StrpHtr-HeVessel-4 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 2, 5A-03 |  | RTD_HeVessel-5 Top | Cernox CX-1030-SD-HT-1.4L | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL02B25 | RTD_HeVessel-5 Bottom | Cernox CX-1030-SD-HT-1.4L |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-5-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-5-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 5B-07 | FL02B25 | StrpHtr-HeVessel-5 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 2, 6A-03 |  | Spare | N/A | 4 pr 22 ga | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL02B25 | Spare | N/A |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-6-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-6-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 6B-07 | FL02B25 | StrpHtr-HeVessel-6 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 2, 7A-03 |  | Spare | N/A | 4 pr 22 ga | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL02B25 | Spare | N/A |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-7-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-7-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 7B-07 | FL02B25 | StrpHtr-HeVessel-7 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 2, 8A-03 |  | Spare | N/A | 4 pr 22 ga | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL02B25 | Spare | N/A |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-8-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-8-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 8B-07 | FL02B25 | StrpHtr-HeVessel-8 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 2, D-07 | FL02B25 | CoolDownHtr-1 | Primary Cool Down Heater #1 Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) , B(-)
|  | CoolDownHtr-2 | Primary Cool Down Heater #2 Kapton Polymide Strip Heater, 46.3 Ω |  |  | C(+) , D(-)
CM 2, D-08 | FL02B25 | CoolDownHtr-3 | Primary Cool Down Heater #3 Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) , B(-)
|  | Heater in Downstream Level Can | W4C050 50V, 100 Watt Cartridge Heater |  |  | C(+) , D(-)
CM 2, E1-01 | FL02B25 | LL_Downstream | Cryomagnetics 3D-031-033-PB-RAD-2K 12" Liquid Level Probe | 2 pr 20 ga | 16060-02-A | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | Spare | Jumper/ Potential for Interlock |  |  | E,F
CM 2, E2-01 | FL02B25 | LL_Downstream | Cryomagnetics 3D-031-033-PB-RAD-2K 12" Liquid Level Probe | 2 pr 20 ga | MS3126F10-6S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | Spare | Jumper/ Potential for Interlock |  |  | E,F
CM 2, F-01 | FL02B25 | RTD_MP_1 | Cernox CX-1030-SD-HT-1.4L | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_MP_2 | DT-670A1-SD |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MP_3 | DT-670A1-SD |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_MP_4 | DT-670A1-SD |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, F-02 | FL02B25 | RTD_50K_LeadThermIntcp | DT-670A1-SD | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_5K_LeadThermIntcp | DT-670A1-SD |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_2K_LeadThermIntcp_1 | DT-670A1-SD |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_2K_LeadThermIntcp_2 | DT-670A1-SD |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, K-01 | FL02B25 | RTD_Line_A1 | DT-670A1-CU | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_Line_B1 | DT-670A1-CU |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_Line_B2 | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_Line_C1 | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, K-02 | FL02B25 | RTD_Line_D1 | DT-670A1-CU | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_Line_E1 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_Line_F1 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_Shield_1 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, K-03 | FL02B25 | RTD_Cooldown_1 | Cernox CX-1030-CU-HT-1.4L | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_Cooldown_2 | Cernox CX-1030-CU-HT-1.4L |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_Faraday Window 1 Port | Cernox CX-1030-SD-HT |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_Faraday Window 2 Port | Cernox CX-1030-SD-HT |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, K-04 | FL02B25 | Heater in Upstream Level Can | W4C050 50V, 100 Watt Cartridge Heater | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 2, L1-01 | FL02B25 | LL_Upstream | Cryomagnetics 3D-031-033-PB-RAD-2K 12" Liquid Level Probe | 2 pr 20 ga | 16060-02-A | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | Spare | Jumper/ Potential for Interlock |  |  | E,F
CM 2, L2-01 | FL02B25 | LL_Upstream | Cryomagnetics 3D-031-033-PB-RAD-2K 12" Liquid Level Probe | 2 pr 20 ga | MS3126F10-6S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | Spare | Jumper/ Potential for Interlock |  |  | E,F

## Sheet11
Connector | Rack Loc. | Device/Sensor Name | Device/Sensor Type | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 2, 1A-01 | FL04B34 | (Piezo Actuator) FastTuner-1-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
CM 2, 1A-02 | FL04B34 | RTD_StepperMotor-1 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
CM 2, 1B-01 | FL04B34 | (Stepper Motor) SlowTuner-1 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
CM 2, 1B-02 | FL04B34 | C1_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, 1B-05 | FL04B34 | Cav_1_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 1B-06 | FL04B34 | Cav_1_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, Cav 1-F | FL04B34 | Cavity 1 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, Cav 1-R | FL04B34 | Cavity 1 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, 2A-01 | FL04B34 | (Piezo Actuator) FastTuner-2-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
CM 2, 2A-02 | FL04B34 | RTD_StepperMotor-2 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
CM 2, 2B-01 | FL04B34 | (Stepper Motor) SlowTuner-2 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
CM 2, 2B-02 | FL04B34 | C2_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, 2B-05 | FL04B34 | Cav_2_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 2B-06 | FL04B34 | Cav_2_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, Cav 2-F | FL04B34 | Cavity 2 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, Cav 2-R | FL04B34 | Cavity 2 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, 3A-01 | FL04B34 | (Piezo Actuator) FastTuner-3-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
CM 2, 3A-02 | FL04B34 | RTD_StepperMotor-3 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
CM 2, 3B-01 | FL04B34 | (Stepper Motor) SlowTuner-3 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
CM 2, 3B-02 | FL04B34 | C3_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, 3B-05 | FL04B34 | Cav_3_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 3B-06 | FL04B34 | Cav_3_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, Cav 3-F | FL04B34 | Cavity 3 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, Cav 3-R | FL04B34 | Cavity 3 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, 4A-01 | FL04B34 | (Piezo Actuator) FastTuner-4-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
CM 2, 4A-02 | FL04B34 | RTD_StepperMotor-4 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
CM 2, 4B-01 | FL04B34 | (Stepper Motor) SlowTuner-4 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
CM 2, 4B-02 | FL04B34 | C4_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, 4B-05 | FL04B34 | Cav_4_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 4B-06 | FL04B34 | Cav_4_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, Cav 4-F | FL04B34 | Cavity 4 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, Cav 4-R | FL04B34 | Cavity 4 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield

## Sheet10
CM 1, 1A-03
CM 1, 1B-07
CM 1, 2A-03
CM 1, 2B-07
CM 1, 3A-03
CM 1, 3B-07
CM 1, 4A-03
CM 1, 4B-07
CM 1, 5A-03
CM 1, 5B-07
CM 1, 6A-03
CM 1, 6B-07
CM 1, 7A-03
CM 1, 7B-07
CM 1, 8A-03
CM 1, 8B-07
CM 1, C-04
CM 1, D-07
CM 1, D-08
CM 1, E1-01
CM 1, E2-01
CM 1, F-01
CM 1, F-02
CM 1, K-01
CM 1, K-02
CM 1, K-03
CM 1, K-04
CM 1, L1-01
CM 1, L1-BAR
CVM 1, L2-01

## FL02B27
Connector | Rack Loc. | Device/Sensor Name | Device/Sensor Type | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 1, 1A-03 | FL02B27 | RTD_HeVessel-1 Top | Cernox CX-1030-SD-HT-1.4L | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_HeVessel-1 Bottom | Cernox CX-1030-SD-HT-1.4L |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-1-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-1-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 1B-07 | FL02B27 | StrpHtr-HeVessel-1 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 1, 2A-03 | FL02B27 | Spare | N/A | 4 pr 22 ga | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-2-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-2-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 2B-07 | FL02B27 | StrpHtr-HeVessel-2 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 1, 3A-03 | FL02B27 | Spare | N/A | 4 pr 22 ga | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-3-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-3-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 3B-07 | FL02B27 | StrpHtr-HeVessel-3 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 1, 4A-03 | FL02B27 | Spare | N/A | 4 pr 22 ga | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-4-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-4-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 4B-07 | FL02B27 | StrpHtr-HeVessel-4 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 1, 5A-03 |  | RTD_HeVessel-5 Top | Cernox CX-1030-SD-HT-1.4L | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL02B27 | RTD_HeVessel-5 Bottom | Cernox CX-1030-SD-HT-1.4L |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-5-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-5-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 5B-07 | FL02B27 | StrpHtr-HeVessel-5 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 1, 6A-03 |  | Spare | N/A | 4 pr 22 ga | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL02B27 | Spare | N/A |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-6-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-6-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 6B-07 | FL02B27 | StrpHtr-HeVessel-6 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 1, 7A-03 |  | Spare | N/A | 4 pr 22 ga | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL02B27 | Spare | N/A |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-7-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-7-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 7B-07 | FL02B27 | StrpHtr-HeVessel-7 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 1, 8A-03 |  | Spare | N/A | 4 pr 22 ga | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL02B27 | Spare | N/A |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Coupler_Side-8-2 Copper Clamp | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_HOM_Tuner_Side-8-4 Copper Clamp | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 8B-07 | FL02B27 | StrpHtr-HeVessel-8 | HK5347R46.3L12AU Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 1, C-04 | FL02B27 | Heater Q1 & Q2 | Split Quadrupole and Corrector Coils Persistant Current Magnet Heaters | 4 cond 18 ga Alpha 2424C | MS24266R14B4SN | 1(+), 2(-)
|  | Heater Q3 & Q4 |  |  |  | 3(+), 4(-)
CM 1, D-07 | FL02B27 | CoolDownHtr-1 | Primary Cool Down Heater #1 Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) , B(-)
|  | CoolDownHtr-2 | Primary Cool Down Heater #2 Kapton Polymide Strip Heater, 46.3 Ω |  |  | C(+) , D(-)
CM 1, D-08 | FL02B27 | CoolDownHtr-3 | Primary Cool Down Heater #3 Kapton Polymide Strip Heater, 46.3 Ω | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) , B(-)
|  | Heater in Downstream Level Can | W4C050 50V, 100 Watt Cartridge Heater |  |  | C(+) , D(-)
CM 1, E1-01 | FL02B27 | LL_Downstream | Cryomagnetics 3D-031-033-PB-RAD-2K 12" Liquid Level Probe | 4 cond 18 ga; Alpha 2424C | 16060-02-A | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | Spare | Jumper/ Potential for Interlock |  |  | E,F
| FL02B27 | Baratron
CM 1, E2-01 | FL02B27 | LL_Downstream | Cryomagnetics 3D-031-033-PB-RAD-2K 12" Liquid Level Probe | 4 cond 18 ga; Alpha 2424C | MS3126F10-6S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | Spare | Jumper/ Potential for Interlock |  |  | E,F
CM 1, F-01 | FL02B27 | RTD_MP_1 | Cernox CX-1030-SD-HT-1.4L | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_MP_2 | DT-670A1-SD |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MP_3 | DT-670A1-SD |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_MP_4 | DT-670A1-SD |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, F-02 | FL02B27 | RTD_50K_LeadThermIntcp | DT-670A1-SD | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_5K_LeadThermIntcp | DT-670A1-SD |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_2K_LeadThermIntcp_1 | DT-670A1-SD |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_2K_LeadThermIntcp_2 | DT-670A1-SD |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, K-01 | FL02B27 | RTD_Line_A1 | DT-670A1-CU | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_Line_B1 | DT-670A1-CU |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_Line_B2 | DT-670A1-CU |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_Line_C1 | DT-670A1-CU |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, K-02 | FL02B27 | RTD_Line_D1 | DT-670A1-CU | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_Line_E1 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_Line_F1 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_Shield_1 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, K-03 | FL02B27 | RTD_Cooldown_1 | Cernox CX-1030-CU-HT-1.4L | 8 pr 22 ga Alpha 6379S | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_Cooldown_2 | Cernox CX-1030-CU-HT-1.4L |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_Faraday Window 1 Port | Cernox CX-1030-SD-HT |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | RTD_Faraday Window 2 Port | Cernox CX-1030-SD-HT |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, K-04 | FL02B27 | Heater in Upstream Level Can | W4C050 50V, 100 Watt Cartridge Heater | 4 cond 18 ga; Alpha 2424C | MS24266R14B4SN | A(+) jumper to C(+) , B(-) jumper to D(-)
CM 1, L1-01 | FL02B27 | LL_Upstream | Cryomagnetics 3D-031-033-PB-RAD-2K 12" Liquid Level Probe | 4 cond 18 ga; Alpha 2424C | 16060-02-A | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | Spare | Jumper/ Potential for Interlock |  |  | E, F
CM 1, L1-BAR | FL02B27 | Baratron
CM 1, L2-01 | FL02B27 | LL_Upstream | Cryomagnetics 3D-031-033-PB-RAD-2K 12" Liquid Level Probe | 4 cond 18 ga; Alpha 2424C | MS3126F10-6S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | Spare | Jumper/ Potential for Interlock |  |  | E,F

## FL03B24
Connector | Rack Loc. | Device/Sensor Name | Device/Sensor Type | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 1, 1B-03 | FL03B24 | C1_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 1B-04 | FL03B24 | C1_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 2B-03 | FL03B24 | C2_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 2B-04 | FL03B24 | C2_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 3B-03 | FL03B24 | C3_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 3B-04 | FL03B24 | C3_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 4B-03 | FL03B24 | C4_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 4B-04 | FL03B24 | C4_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 5B-03 | FL03B24 | C5_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 5B-04 | FL03B24 | C5_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 6B-03 | FL03B24 | C6_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 6B-04 | FL03B24 | C6_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 7B-03 | FL03B24 | C7_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 7B-04 | FL03B24 | C7_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 8B-03 | FL03B24 | C8_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 1, 8B-04 | FL03B24 | C8_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 1B-03 | FL03B24 | C1_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 1B-04 | FL03B24 | C1_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 2B-03 | FL03B24 | C2_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 2B-04 | FL03B24 | C2_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 3B-03 | FL03B24 | C3_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 3B-04 | FL03B24 | C3_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 4B-03 | FL03B24 | C4_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 4B-04 | FL03B24 | C4_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 5B-03 | FL03B24 | C5_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 5B-04 | FL03B24 | C5_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 6B-03 | FL03B24 | C6_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 6B-04 | FL03B24 | C6_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 7B-03 | FL03B24 | C7_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 7B-04 | FL03B24 | C7_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 8B-03 | FL03B24 | C8_H1 (PTrans_HOM-1) | HOM 1 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield
CM 2, 8B-04 | FL03B24 | C8_H2 (PTrans_HOM-2) | HOM 2 -- Transmitted Power | 1/4" Heliax | N 0° | Single Pin,Shield

## FL03B25
Connector | Rack Loc. | Device/Sensor Name | Device/Sensor Type | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 1, 1A-01 | FL03B25 | (Piezo Actuator) FastTuner-1-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-1-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
|  | (Piezo Actuator) FastTuner-1-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-1-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-1 | Piezo Fast Tuner, Interlock |  |  | J-K & L-M
CM 1, 1A-02 | FL03B25 | RTD_StepperMotor-1 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler1a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler1b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 1B-01 | FL03B25 | (Stepper Motor) SlowTuner-1 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
|  | Tuner limit Switch-1-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-1-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 1, 1B-02 | FL03B25 | C1_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, 1B-05 | FL03B25 | Cav_1_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 1B-06 | FL03B25 | Cav_1_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, Cav 1-F | FL03B25 | Cavity 1 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, Cav 1-R | FL03B25 | Cavity 1 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, 2A-01 | FL03B25 | (Piezo Actuator) FastTuner-2-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-2-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
|  | (Piezo Actuator) FastTuner-2-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-2-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-2 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 1, 2A-02 | FL03B25 | RTD_StepperMotor-2 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler2a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler2b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 2B-01 | FL03B25 | (Stepper Motor) SlowTuner-2 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
|  | Tuner limit Switch-2-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-2-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 1, 2B-02 | FL03B25 | C2_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, 2B-05 | FL03B25 | Cav_2_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 2B-06 | FL03B25 | Cav_2_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, Cav 2-F | FL03B25 | Cavity 2 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, Cav 2-R | FL03B25 | Cavity 2 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, 3A-01 | FL03B25 | (Piezo Actuator) FastTuner-3-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-3-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
|  | (Piezo Actuator) FastTuner-3-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-3-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-3 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 1, 3A-02 | FL03B25 | RTD_StepperMotor-3 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler2a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler2b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 3B-01 | FL03B25 | (Stepper Motor) SlowTuner-3 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
|  | Tuner limit Switch-3-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-3-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 1, 3B-02 | FL03B25 | C3_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, 3B-05 | FL03B25 | Cav_3_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 3B-06 | FL03B25 | Cav_3_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, Cav 3-F | FL03B25 | Cavity 3 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, Cav 3-R | FL03B25 | Cavity 3 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, 4A-01 | FL03B25 | (Piezo Actuator) FastTuner-4-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-4-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
|  | (Piezo Actuator) FastTuner-4-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-4-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-4 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 1, 4A-02 | FL03B25 | RTD_StepperMotor-4 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler4a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler4b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 4B-01 | FL03B25 | (Stepper Motor) SlowTuner-4 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
|  | Tuner limit Switch-4-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-4-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 1, 4B-02 | FL03B25 | C4_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, 4B-05 | FL03B25 | Cav_4_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 4B-06 | FL03B25 | Cav_4_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, Cav 4-F | FL03B25 | Cavity 4 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, Cav 4-R | FL03B25 | Cavity 4 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield

## FL03B34
Connector | Rack Loc. | Device/Sensor Name | Device/Sensor Type | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 1, 5A-01 |  | (Piezo Actuator) FastTuner-5-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-5-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
| FL03B34 | (Piezo Actuator) FastTuner-5-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-5-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-5 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 1, 5A-02 |  | RTD_StepperMotor-5 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL03B34 | RTD_MainCoupler5a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler5b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 5B-01 |  | (Stepper Motor) SlowTuner-5 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
| FL03B34 | Tuner limit Switch-5-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-5-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 1, 5B-02 | FL03B34 | C5_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, 5B-05 | FL03B34 | Cav_5_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 5B-06 | FL03B34 | Cav_5_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, Cav 5-F | FL03B34 | Cavity 5 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, Cav 5-R | FL03B34 | Cavity 5 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, 6A-01 |  | (Piezo Actuator) FastTuner-6-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-6-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
| FL03B34 | (Piezo Actuator) FastTuner-6-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-6-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-6 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 1, 6A-02 |  | RTD_StepperMotor-6 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL03B34 | RTD_MainCoupler6a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler6b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 6B-01 |  | (Stepper Motor) SlowTuner-6 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
| FL03B34 | Tuner limit Switch-6-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-6-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 1, 6B-02 | FL03B34 | C6_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, 6B-05 | FL03B34 | Cav_6_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 6B-06 | FL03B34 | Cav_6_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, Cav 6-F | FL03B34 | Cavity 6 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, Cav 6-R | FL03B34 | Cavity 6 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, 7A-01 |  | (Piezo Actuator) FastTuner-7-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-7-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
| FL03B34 | (Piezo Actuator) FastTuner-7-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-7-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-7 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 1, 7A-02 |  | RTD_StepperMotor-7 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL03B34 | RTD_MainCoupler7a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler7b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 7B-01 |  | (Stepper Motor) SlowTuner-7 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
| FL03B34 | Tuner limit Switch-7-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-7-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 1, 7B-02 | FL03B34 | C7_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, 7B-05 | FL03B34 | Cav_7_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 7B-06 | FL03B34 | Cav_7_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, Cav 7-F | FL03B34 | Cavity 7 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, Cav 7-R | FL03B34 | Cavity 7 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, 8A-01 |  | (Piezo Actuator) FastTuner-8-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-8-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
| FL03B34 | (Piezo Actuator) FastTuner-8-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-8-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-8 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 1, 8A-02 |  | RTD_StepperMotor-8 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL03B34 | RTD_MainCoupler8a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler8b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 1, 8B-01 |  | (Stepper Motor) SlowTuner-8 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
| FL03B34 | Tuner limit Switch-8-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-8-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 1, 8B-02 | FL03B34 | C8_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, 8B-05 | FL03B34 | Cav_8_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, 8B-06 | FL03B34 | Cav_8_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 1, Cav 8-F | FL03B34 | Cavity 8 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 1, Cav 8-R | FL03B34 | Cavity 8 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield

## FL04B25
Connector | Rack Loc. | Device/Sensor Name | Device/Sensor Type | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 2, 5A-01 |  | (Piezo Actuator) FastTuner-5-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-5-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
| FL04B25 | (Piezo Actuator) FastTuner-5-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-5-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-5 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 2, 5A-02 |  | RTD_StepperMotor-5 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL04B25 | RTD_MainCoupler5a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler5b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 5B-01 |  | (Stepper Motor) SlowTuner-5 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
| FL04B25 | Tuner limit Switch-5-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-5-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 2, 5B-02 | FL04B25 | C5_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, 5B-05 | FL04B25 | Cav_5_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 5B-06 | FL04B25 | Cav_5_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, Cav 5-F | FL04B25 | Cavity 5 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, Cav 5-R | FL04B25 | Cavity 5 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, 6A-01 |  | (Piezo Actuator) FastTuner-6-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-6-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
| FL04B25 | (Piezo Actuator) FastTuner-6-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-6-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-6 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 2, 6A-02 |  | RTD_StepperMotor-6 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL04B25 | RTD_MainCoupler6a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler6b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 6B-01 |  | (Stepper Motor) SlowTuner-6 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
| FL04B25 | Tuner limit Switch-6-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-6-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 2, 6B-02 | FL04B25 | C6_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, 6B-05 | FL04B25 | Cav_6_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 6B-06 | FL04B25 | Cav_6_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, Cav 6-F | FL04B25 | Cavity 6 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 6, Cav 6-R | FL04B25 | Cavity 6 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, 7A-01 |  | (Piezo Actuator) FastTuner-7-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-7-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
| FL04B25 | (Piezo Actuator) FastTuner-7-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-7-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-7 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 2, 7A-02 |  | RTD_StepperMotor-7 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL04B25 | RTD_MainCoupler7a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler7b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 7B-01 |  | (Stepper Motor) SlowTuner-7 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
| FL04B25 | Tuner limit Switch-7-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-7-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 2, 7B-02 | FL04B25 | C7_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, 7B-05 | FL04B25 | Cav_7_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 7B-06 | FL04B25 | Cav_7_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, Cav 7-F | FL04B25 | Cavity 7 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, Cav 7-R | FL04B25 | Cavity 7 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, 8A-01 |  | (Piezo Actuator) FastTuner-8-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-8-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
| FL04B25 | (Piezo Actuator) FastTuner-8-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-8-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-8 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 2, 8A-02 |  | RTD_StepperMotor-8 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL04B25 | RTD_MainCoupler8a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler8b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 8B-01 |  | (Stepper Motor) SlowTuner-8 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
| FL04B25 | Tuner limit Switch-8-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-8-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 2, 8B-02 | FL04B25 | C8_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, 8B-05 | FL04B25 | Cav_8_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 8B-06 | FL04B25 | Cav_8_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, Cav 8-F | FL04B25 | Cavity 8 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, Cav 8-R | FL04B25 | Cavity 8 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield

## FL04B34
Connector | Rack Loc. | Device/Sensor Name | Device/Sensor Type | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 2, 1A-01 |  | (Piezo Actuator) FastTuner-1-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-1-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
| FL04B34 | (Piezo Actuator) FastTuner-1-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-1-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-1 | Piezo Fast Tuner, Interlock |  |  | J-K & L-M
CM 2, 1A-02 |  | RTD_StepperMotor-1 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL04B34 | RTD_MainCoupler1a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler1b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 1B-01 |  | (Stepper Motor) SlowTuner-1 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
| FL04B34 | Tuner limit Switch-1-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-1-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 2, 1B-02 | FL04B34 | C1_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, 1B-05 | FL04B34 | Cav_1_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 1B-06 | FL04B34 | Cav_1_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, Cav 1-F | FL04B34 | Cavity 1 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, Cav 1-R | FL04B34 | Cavity 1 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, 2A-01 |  | (Piezo Actuator) FastTuner-2-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-2-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
| FL04B34 | (Piezo Actuator) FastTuner-2-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-2-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-2 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 2, 2A-02 |  | RTD_StepperMotor-2 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL04B34 | RTD_MainCoupler2a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler2b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 2B-01 |  | (Stepper Motor) SlowTuner-2 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
| FL04B34 | Tuner limit Switch-2-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-2-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 2, 2B-02 | FL04B34 | C2_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, 2B-05 | FL04B34 | Cav_2_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 2B-06 | FL04B34 | Cav_2_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, Cav 2-F | FL04B34 | Cavity 2 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, Cav 2-R | FL04B34 | Cavity 2 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, 3A-01 |  | (Piezo Actuator) FastTuner-3-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-3-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
| FL04B34 | (Piezo Actuator) FastTuner-3-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-3-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-3 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 2, 3A-02 |  | RTD_StepperMotor-3 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL04B34 | RTD_MainCoupler2a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler2b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 3B-01 |  | (Stepper Motor) SlowTuner-3 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
| FL04B34 | Tuner limit Switch-3-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-3-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 2, 3B-02 | FL04B34 | C3_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, 3B-05 | FL04B34 | Cav_3_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 3B-06 | FL04B34 | Cav_3_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, Cav 3-F | FL04B34 | Cavity 3 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, Cav 3-R | FL04B34 | Cavity 3 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, 4A-01 |  | (Piezo Actuator) FastTuner-4-1 | Piezo Fast Tuner, Part# P-844K075 | 4 pr 22 ga, Alpha 2463C | MS3126F14-12S | A,B (V+,V-)
|  | (Piezo Actuator) FastTuner-4-2 | Piezo Fast Tuner, continued |  |  | C,D (V+,V-)
| FL04B34 | (Piezo Actuator) FastTuner-4-3 | Piezo Fast Tuner, Part# P-844K075 |  |  | E,F (V+,V-)
|  | (Piezo Actuator) FastTuner-4-4 | Piezo Fast Tuner, continued |  |  | G,H (V+,V-)
|  | (Piezo Actuators) Interlock-4 | Piezo Fast Tuner, Interlock |  |  | L,M
CM 2, 4A-02 |  | RTD_StepperMotor-4 | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) | 6 pr 20 ga Alpha 6036C | MS3126F14-19S | A,B,C,D (I+R,I-B,V+G,V-Y)
| FL04B34 | RTD_MainCoupler4a-12o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | E,F,G,H (I+R,I-B,V+G,V-Y)
|  | RTD_MainCoupler4b-6o'clock | PT-103 (Platinum 100 Ohm, European Standard DIN/IEC 60751) |  |  | J,K,L,M (I+R,I-B,V+G,V-Y)
|  | Spare | N/A |  |  | N,P,R,S (I+R,I-B,V+G,V-Y)
CM 2, 4B-01 |  | (Stepper Motor) SlowTuner-4 | Phytron Stepper Motor, Model #: VSS 52.200.2.5-4LP-5M-UHVC | 4 pr 20 ga Alpha #2467 | MS3126F16-8S | A,B,C,D
| FL04B34 | Tuner limit Switch-4-1 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | E,F
|  | Tuner limit Switch-4-2 | Stepper Motor Tuner limit Switches Part #: S15-276 |  |  | G,H
CM 2, 4B-02 | FL04B34 | C4_Pt (PTrans) | Cavity Pickup -- Transmitted Power (Field Probe) | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, 4B-05 | FL04B34 | Cav_4_e-Cy (EPickup_Cavity) | 70KCy e- 1.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, 4B-06 | FL04B34 | Cav_4_e-WG (EPickup_Waveguide) | 70KWg e- 2.1 (Electron Pickup) | 3/8" Heliax | SMA 00 | Single Pin,Shield
CM 2, Cav 4-F | FL04B34 | Cavity 4 coupler forward power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield
CM 2, Cav 4-R | FL04B34 | Cavity 4 coupler reflected power | WR650 Waveguide coupler | 3/8" Heliax | N 0° | Single Pin,Shield

## Sheet2
Connector | Rack Loc. | Device/Sensor Name | Device/Sensor Type | Cable | Feed Through Mating Connection(AIR) | Feed Through Connector Pin
CM 1, C-01 |  | VD VT Positive Lead (Whole_Coil-VT6) | Vertical Dipole Coil Voltage Tap (+) | 8cond 22ga Alpha 1298C | MS3126F16-8S | Pin-A - Black
|  | VD VT Negative Lead (Whole_Coil-VT5) | Vertical Dipole Coil Voltage Tap (-) |  |  | Pin-B - Brown
|  | HD VT Positive Lead (Whole_Coil-VT4) | Horizontal Dipole Coil Voltage Tap (+) |  |  | Pin-C - Red
|  | HD VT Negative Lead (Whole_Coil-VT3) | Horizontal Dipole Coil Voltage Tap (-) |  |  | Pin-D - Orange
|  | Quad VT Positive Lead (Whole_Coil-VT1) | Split Quadrupole Coil Voltage Tap (+) |  |  | Pin-E - Yellow
|  | Quad VT Negative Lead (Whole_Coil-VT2) | Split Quadrupole Coil Voltage Tap (-) |  |  | Pin-F - Green
|  | Spare | Spare |  |  | Pin-G - Blue
|  | Spare | Spare |  |  | Pin-H - White
CM 1, C-02 |  | Vertical Dipole VTRes28 (PosHalf Coil) | Vertical Dipole Coil Voltage Tap (+) | 8cond 22ga Alpha 1298C | MS3126F16-8S | Pin-A - Black
|  | Vertical Dipole VTRes29 (PosHalf Coil) | Vertical Dipole Coil Voltage Tap (-) |  |  | Pin-B - Brown
|  | Vertical Dipole VTRes30 (NegHalf Coil) | Vertical Dipole Coil Voltage Tap (+) |  |  | Pin-C - Red
|  | Vertical Dipole VTRes31 (NegHalf Coil) | Vertical Dipole Coil Voltage Tap (-) |  |  | Pin-D - Orange
|  | Horizontal Dipole VTRes22 (PosHalf Coil) | Horizontal Dipole Coil Voltage Tap (+) |  |  | Pin-E - Yellow
|  | Horizontal Dipole VTRes23 (PosHalf Coil) | Horizontal Dipole Coil Voltage Tap (-) |  |  | Pin-F - Green
|  | Horizontal Dipole VTRes24 (NegHalf Coil) | Horizontal Dipole Coil Voltage Tap (+) |  |  | Pin-G - Blue
|  | Horizontal Dipole VTRes25 (NegHalf Coil) | Horizontal Dipole Coil Voltage Tap (-) |  |  | Pin-H - White
CM 1, C-03 |  | Quadrupole VTRes16 (PosHalf Coil) | Split Quadrupole Coil Voltage Tap (+) | 8cond 22ga Alpha 1298C | MS3126F16-8S | Pin-A - Black
|  | Quadrupole VTRes17 (PosHalf Coil) | Split Quadrupole Coil Voltage Tap (-) |  |  | Pin-B - Brown
|  | Quadrupole VTRes18 (NegHalf Coil) | Split Quadrupole Coil Voltage Tap (+) |  |  | Pin-C - Red
|  | Quadrupole VTRes19 (NegHalf Coil) | Split Quadrupole Coil Voltage Tap (-) |  |  | Pin-D - Orange
|  | Spare | Spare |  |  | Pin-E - Yellow
|  | Spare | Spare |  |  | Pin-F - Green
|  | Spare | Spare |  |  | Pin-G - Blue
|  | Spare | Spare |  |  | Pin-H - White
CM 1, D-01 |  | BPM-Top | Beam Position Monitor |  | N 00 | Single Pin,Shield
CM 1, D-02 |  | BPM-Right | Beam Position Monitor |  | N 00 | Single Pin,Shield
CM 1, D-03 |  | BPM-Bot | Beam Position Monitor |  | N 00 | Single Pin,Shield
CM 1, D-04 |  | BPM-Left | Beam Position Monitor |  | N 00 | Single Pin,Shield
CM 1, D-05 |  | Fluxgate_Cav1 | Fluxgate | 6pr 20ga | MS3126F14-19S | A,B,C,D (S+,S-,E+,E-)
|  | Fluxgate_Cav2 | Fluxgate |  |  | E,F,G,H (S+,S-,E+,E-)
|  | Fluxgate_Cav5 | Fluxgate |  |  | J,K,L,M (S+,S-,E+,E-)
|  | spare | Spare |  |  | N,P,R,S
CM 1, D-06 |  | Fluxgate_Cav7 | Fluxgate | 6pr 20ga | MS3126F14-19S | A,B,C,D (S+,S-,E+,E-)
|  | Fluxgate_Cav8 | Fluxgate |  |  | E,F,G,H (S+,S-,E+,E-)
|  | spare | Spare |  |  | J,K,L,M
|  | spare | Spare |  |  | N,P,R,S
CM 2, C-01 |  | VD VT Positive Lead (Whole_Coil-VT6) | Vertical Dipole Coil Voltage Tap (+) | 8cond 22ga Alpha 1298C | MS3126F16-8S | Pin-A - Black
|  | VD VT Negative Lead (Whole_Coil-VT5) | Vertical Dipole Coil Voltage Tap (-) |  |  | Pin-B - Brown
|  | HD VT Positive Lead (Whole_Coil-VT4) | Horizontal Dipole Coil Voltage Tap (+) |  |  | Pin-C - Red
|  | HD VT Negative Lead (Whole_Coil-VT3) | Horizontal Dipole Coil Voltage Tap (-) |  |  | Pin-D - Orange
|  | Quad VT Positive Lead (Whole_Coil-VT1) | Split Quadrupole Coil Voltage Tap (+) |  |  | Pin-E - Yellow
|  | Quad VT Negative Lead (Whole_Coil-VT2) | Split Quadrupole Coil Voltage Tap (-) |  |  | Pin-F - Green
|  | Spare | Spare |  |  | Pin-G - Blue
|  | Spare | Spare |  |  | Pin-H - White
CM 2, C-02 |  | Vertical Dipole VTRes28 (PosHalf Coil) | Vertical Dipole Coil Voltage Tap (+) | 8cond 22ga Alpha 1298C | MS3126F16-8S | Pin-A - Black
|  | Vertical Dipole VTRes29 (PosHalf Coil) | Vertical Dipole Coil Voltage Tap (-) |  |  | Pin-B - Brown
|  | Vertical Dipole VTRes30 (NegHalf Coil) | Vertical Dipole Coil Voltage Tap (+) |  |  | Pin-C - Red
|  | Vertical Dipole VTRes31 (NegHalf Coil) | Vertical Dipole Coil Voltage Tap (-) |  |  | Pin-D - Orange
|  | Horizontal Dipole VTRes22 (PosHalf Coil) | Horizontal Dipole Coil Voltage Tap (+) |  |  | Pin-E - Yellow
|  | Horizontal Dipole VTRes23 (PosHalf Coil) | Horizontal Dipole Coil Voltage Tap (-) |  |  | Pin-F - Green
|  | Horizontal Dipole VTRes24 (NegHalf Coil) | Horizontal Dipole Coil Voltage Tap (+) |  |  | Pin-G - Blue
|  | Horizontal Dipole VTRes25 (NegHalf Coil) | Horizontal Dipole Coil Voltage Tap (-) |  |  | Pin-H - White
CM 2, C-03 |  | Quadrupole VTRes16 (PosHalf Coil) | Split Quadrupole Coil Voltage Tap (+) | 8cond 22ga Alpha 1298C | MS3126F16-8S | Pin-A - Black
|  | Quadrupole VTRes17 (PosHalf Coil) | Split Quadrupole Coil Voltage Tap (-) |  |  | Pin-B - Brown
|  | Quadrupole VTRes18 (NegHalf Coil) | Split Quadrupole Coil Voltage Tap (+) |  |  | Pin-C - Red
|  | Quadrupole VTRes19 (NegHalf Coil) | Split Quadrupole Coil Voltage Tap (-) |  |  | Pin-D - Orange
|  | Spare | Spare |  |  | Pin-E - Yellow
|  | Spare | Spare |  |  | Pin-F - Green
|  | Spare | Spare |  |  | Pin-G - Blue
|  | Spare | Spare |  |  | Pin-H - White
CM 2, C-04 |  | Heater Q1 & Q2 | Split Quadrupole and Corrector Coils Persistant Current Magnet Heaters | 4 cond 18 ga Alpha 2424C | MS24266R14B4SN | 1(+), 2(-)
|  | Heater Q3 & Q4 |  |  |  | 3(+), 4(-)
CM 2, D-01 |  | BPM-Top | Beam Position Monitor |  | N 00 | Single Pin,Shield
CM 2, D-02 |  | BPM-Right | Beam Position Monitor |  | N 00 | Single Pin,Shield
CM 2, D-03 |  | BPM-Bot | Beam Position Monitor |  | N 00 | Single Pin,Shield
CM 2, D-04 |  | BPM-Left | Beam Position Monitor |  | N 00 | Single Pin,Shield
CM 2, D-05 |  | Fluxgate_Cav1 | Fluxgate |  | MS3126F14-19S | A,B,C,D (S+,S-,E+,E-)
|  | Fluxgate_Cav2 | Fluxgate |  |  | E,F,G,H (S+,S-,E+,E-)
|  | Fluxgate_Cav5 | Fluxgate |  |  | J,K,L,M (S+,S-,E+,E-)
|  | spare | Spare |  |  | N,P,R,S
CM 2, D-06 |  | Fluxgate_Cav7 | Fluxgate |  | MS3126F14-19S | A,B,C,D (S+,S-,E+,E-)
|  | Fluxgate_Cav8 | Fluxgate |  |  | E,F,G,H (S+,S-,E+,E-)
|  | spare | Spare |  |  | J,K,L,M
|  | spare | Spare |  |  | N,P,R,S

## Color Codes
Cable
8 pair 22 gauge