We thank the reviewer for their helpful comments to improve the manuscript.

Please find below a reproduction of the reviewer’s comments (bolded) followed by our responses and corresponding corrections made to the re-submitted manuscript.

Kind regards,

1. The authors should consider including a sentence that states the systematic uncertainty of 0.6% in the abstract. The reader should not have to wait until Section 11 to see that number.

- The overall uncertainty we quote (not the systematic uncertainty) is 0.61%.  This includes the dominant uncertainty of 0.5% from the calculated value of the Sherman function; 0.24% from systematic uncertainties arising from our energy cut on the E counter pulse height spectrum (0.1%), the polarization of the laser beam incident on the photocathode (0.1%), and uncertainties in the beam energy and scattering angle (0.2%); and a 0.25% uncertainty from our fit to the measured asymmetry as a function of foil thickness.  The following sentence attempting to summarize this is added to the Abstract;

  - “The overall uncertainty of our beam polarization measurement, arising from the uncertainty in the value of the scattering asymmetry at zero foil thickness as determined by our fits to the measured asymmetries versus scattering foil thicknesses; the estimated systematic effects; and the (dominant) uncertainty from the calculation of the theoretical Sherman function, is 0.61%.”

2. You write in the abstract that "A simultaneous high-precision measurement of the beam polarization with a different polarimeter..." Can you be more specific and name that other polarimeter?

- We updated the corresponding sentence in the Abstract;

  - “A simultaneous high-precision measurement of the beam polarization with a different polarimeter, AESOP (Accurate Electron Spin Optical Polarimeter) under development at the University of Nebraska, clearly possible at the CEBAF accelerator, will allow a high precision comparison of our measured asymmetries with theoretical calculations of the Mott analyzing power.”

- We added a recent reference (T-2) on the development of AESOP, which itself includes further references;

  - K. Trantham, K. Foreman, and T. Gay, Demonstration of vacuum strain effects on a light-collection lens used in optical polarimetry, Applied Optics 59(9) 2715-2724.

3. In the same spirit of being more precise, can you provide a number for the energy resolution of typical detectors on page 3, where you write "particularly when the energy resolution of typical detectors is included."

Since each scattering is primarily elastic or quasi-elastic, the scattered electron energy is not a very useful discriminant, compared with the 1-10 percent level energy resolution of corresponding scintillating materials typically employed with MeV energy electron detection.

- The variety of detectors employed in Mott scattering experiments is broad, and so there is a range of energy resolution, just as there is a range in energies.  Considering both this comment and one made later #7 we discuss the energy resolution of our “E” detectors in greater detail in Section 7.   However, we have added the following sentence here for better clarity,

  - “Since each scattering is primarily elastic or quasi-elastic, the scattered electron energy is not a very useful discriminant, compared with the percent level energy resolution of commercial scintillating materials often employed for detection of MeV energy electrons (see Section 7).”

4. How large is the contribution from inelastic scattering in the target foil which "makes a negligible contribution" on page 3?

- Inelastic scattering is about 0.01% at 1 MeV, and is realistically expected to decrease with increasing beam energy.  This is discussed in Section 3, on the Calculation of the Theoretical Sherman Function, and in greater detail in the paper (ref. R-3, by X.  Roca-Maza); please see response to comment #5, below.

5. On page 4 you write "Given the dependence of the leading order radiative corrections on energy, this result provides strong circumstantial support that the net effect of these corrections largely cancels, as theoretically anticipated." Can you provide a reference for the theoretical anticipation?

- The calculation of the theoretical Sherman Function is discussed in Section 3, and in greater detail in the papers R-1 and R-3 by X.  Roca-Maza.  Probably the strongest direct evidence that the two first order correction terms have opposite signs is the calculation of the similar diagrams for the case of radiative recombination.  Though the self-energy term is only partially calculated for this case, it is indeed of the opposite sign to the vacuum polarization term (ref. S-10). A sentence was added at this point in the text to indicate our discussion on this point is addressed later in Section 3 and with associated references;

  - “This calculation, and a discussion of the many small effects that must be considered, are thoroughly discussed in Section 3 (R-1, R-3).”

6. Figure 1 has legends that state "calc. for P = ", while the captions states, those calculations are fits. What is true? Please clarify.

- Steigerwald generated values for the scattering asymmetry at zero foil thickness by fitting the measured scattering asymmetries versus target foil thickness at each of the three beam energies with his semi-empirical expression.  The ratio of this zero foil thickness asymmetry to the theoretical Sherman function then gives the beam polarization.  A comment is added to the figure caption;

  - “The fits to the three data sets (measured scattering asymmetry versus target foil thickness) are based on the semi-empirical model developed by M. Steigerwald (S-1).  The fit intercept at zero foil thickness, along with the theoretically calculated Sherman function, then determines the beam polarization.”

7. It is not clear why Figure 2, which appears in Ref T-1, is copied over to your manuscript. Is it really necessary to reprint that figure?

- We believe that showing both experimental results, done at two laboratories, over two different energy ranges, and with two very different polarimeters, each of which did not incorporate any corrections for radiative effects, is useful to support the hypothesis that the net effect of QED and bremsstrahlung effects is very small.  Hopefully, our results will encourage theorists to do the next generation calculations to bound more narrowly the net effect of all these corrections.  If the reviewer strongly believes our paper is stronger or better without this figure, we can remove it and change the text accordingly.

8. Can you use the same (strong) font in Figures 3 & 4, as you use in Figure 5? The labels on the are hard to read in Figures 3 & 4.

- Corrected.  Figures 3 & 4 have been remade with a stronger font, like Figure 5.

9. It appears that you need to add a "-" sign on page 18, 4 lines above Figure 11. Should it state "A "good" elastic scattering event has been determined to lie between -0.5sigma to +2sigma".

- Yes, thank you, that error corrected in text.

10. You use a Gaussian fit on the spectrum shown in Figure 12. This is clearly not accurate, since the left side of the peak is skewed. The spectrum may be better represented by a Gaussian convolved with an Erfc. How would that improve your results?

- Indeed, this is the case.  In fact, many months of analysis are associated with our final analysis. This work is documented in reference M-3, a lengthy 78 page technical note which is available to the reader.   However, the point is well taken, so we have modified and significantly lengthened the discussion in sub-section “Loop #2 – Compute helicity correlated asymmetries” with the following explanatory text, which we believe addresses your comment.  Notably, while there are “good” Mott events outside of the region we use to compute the asymmetry, we contend that our selection of events from the energy spectrum does not bias the scattering asymmetry we calculate using only these events.

  - “The pulse height spectrum in the E detector spectra (of Figure 12) is not Gaussian over the full range of the peak.  This is primarily because there are mechanisms that generate real or apparent energy loss, but none that generate energy gain.  So, for example, imperfect light collection from the scintillator, bremsstrahlung or Compton scattering leading to undetected photon energy, or electron (or positron) escape from the scintillator may all contribute to peak broadening on the low energy side of the peak.  While Geant4 simulations of the detector package were performed to validate these mechanisms, we have not attempted to precisely model the full energy spectrum for the purpose of defining the analysis. Instead extensive examination of the energy spectra with various functional forms (e.g. Gaussian, Lorentzian) led to the use of energy cuts between -0.5 and +2.0.  We further examined these cuts by systematically shrinking or enlarging them in 10% steps up to 30% and noting the effect these changes had on the uncertainty in the asymmetry.  In all cases, at the statistical expense of eliminating events, our choice of cuts led to the smallest uncertainty on the asymmetry, and do not bias the scattering asymmetry.  A systematic uncertainty of 0.1% is assigned to the energy cut.  Thorough details of the analysis study are described in JLab TN 17-025 (M-3).”

11. On page 22, you write "We measured the scattering asymmetry as a function of beam position on the 1 um and 0.225 um". Do you actually mean 1 um or 1.0 um or 1.00 um or 1.000 um? What is the precision on that number?

- Yes, you bring up a good point.  We took the strategy throughout the text of referring to each target foil by its “nominal thickness”, meaning the value as provided by the vendor, Lebow.  However, for our experimental analysis and theoretical models we use the physical value as measured by Field Emission Scanning Electron Microscopy.  The technique, as well as the resulting physical thickness and respective precision of all the targets studies are summarized in Appendix 1.   We first refer to Appendix 1 on page 14 when describing the polarimeter construction.   However, on page 28 where we begin using the target thicknesses in the analysis this sentence has been added,

  - “Although we refer to the foils by their respective nominal thickness as determined by the vendor we use in our experimental analysis and theoretical models the values we determined by the FESEM method (summarized in Appendix 1).”

12. In Figures 14 & 16 you show a dotted band over the entire x-range, while in Figure 15, only over a limited range. You do not state in the captions, what it means, and why it only covers a portion of the graph in Figure 15. I could guess, but I should not have to. Please fix.

- Corrected.

13. On page 31, 3rd line, there is problem with a symbol that is displayed as a square.  Please fix.

- Corrected.  The font of the epsilon symbol () was inadvertently changed.

14. You often use the symbol "~" in the text probably meaning to represent "approximately". However, this is confusing and incorrect. The mathematical symbol "~" means "similar" and not "approximate". Please check and fix. You also use the symbol "<" outside of a mathematical formula. This should be avoided in formal text if possible, and replaced with appropriate text.

[Editorial Office remark: Please replace "$\sim$" with "$\approx$" to denote numerical proximity (to distinguish from functional similarity).]

- We submitted our manuscript as a Word document.   All cases of “~” have been corrected with the appropriate symbol “”. Instances where “<” or “>” have been corrected as well.  We will work with the editor for final submission, whether it be the re-submitted Word document, or in latex.