Ideas for NPS Refurbishment prior to next experiment

Not priority ranked, nor absolutely necessary

- Crystals

  - Bleach with blue lights?

  - Measure transparency?

- PMTs

  - Test ~10 at high anode current (e.g. 10 μA until death)

  - Replace?

- PMT Base Redesign?

  - No LV regulators, replace with LC filter.

  - Find certified rad-hard components for pre-amps

  - Replace signal and HV connectors with pigtails?

  - “True” DC monitoring?

  - Heating system for Base when HV off, to stabilize calo temperature.

- Calo-mounted distribution boards

  - Redesign/replace for proper HV isolation

    - Replace boards with guided cable bundles?

  - Revise connections to PMT bases

  - Any mechanical suggestions (expand lateral movement beyond size of Calo)

  - Low Voltage distribution

  - HV connectors?

- Optical Fibers

  - LED system / driver

  - Coupling to crystals

  - Put microLEDs directly on crystal

- Cabling

  - Build/buy a moving cable tray (stacked vertically) to allow easier motion on platform

- Sweep Magnet

  - Reconsider placement, rotation with respect to crystal axis

- Calorimeter Mechanics

  - Any way to simplify repairs?

    - Expanded lateral movement of distribution boards

  - Improve temperature sensor reliability, radiation hardness

  - Cooling

  - Drain cooling lines before long term storage