## C100-CAV-NCR
Status Report for C100-CAV-NCR on 10-Jun-20
TRAV_SEQ_NUM | SN | SNTYPE | DRAWINGNUMBER | PARTDESCRIPTION | CONNECTIVITYCOMMENT | TRAVELERID | TRAVELERREVISION | TRAVSEQNUMBER | DESCRIBECOMMENT | DESCRIBESTAFF | DESCRIBEDATE | DESCRIBEDOCS | DESCRIPTIONSUPPORTFILES | NOTIFICATIONCOMMENT | RESPONDINGENGINEER | RESPONDINGDATE | DISPOSITIONCOMMENT | DISPOSITION | DISPOSITIONDOCS | REWORKCOMMENT | REWORKSTAFF | REWORKDATE | REWORKDOCS | VERIFYCOMMENT | VERIFYSTAFF | VERIFYDATE | VERIFYDOCS | FINALENGINEER | FINALDATE | FINALCOMMENT | FINALDISPOSITION | FINALDOCS | SERIALID | SRF_CVP_USER | NCR_ISSUE_DATE | DESCRIBEDOCUMENTS | ENGINEERTOEMAIL | NCR_RESPONSE_DATE | NCR_DISPOSITION | DOCUMENTATIONCOMMENT | NCR_DOCUMENTS | VENDORID
8 | 2 |  | CRM12070500000 | Cold Tuner |  | C100-CAV-INSP-CTUN2 | R1 | 1 | See inspection traveler for details. | CARPENTE | 2010-03-19 |  |  |  | JOSEPHM |  | Issues addressed with vendor & resolved | [15 empty] | 2 | CARPENTE | 2010-03-19 |  | josephm |  | Use As Is | Issues addressed with vendor & resolved
9 | F063951/01 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R3 | 20 | O-ring groove width is out of tolerance. | carpente | 2010-04-19 |  |  |  | KWILSON |  |  | Use As Is | [14 empty] | F063951/01 | carpente | 2010-04-19 |  | kwilson |  | Use As Is |  |  | Calorstat
10 | F063951/06 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R3 | 27 | Surface finish, bolt hole location, and o-ring groove width are all out of toler | GEORGED | 2010-04-19 |  |  |  | KWILSON |  | Dimensions ok. Please repair scratches as Aaron recommends. | Modify | [14 empty] | F063951/06 | GEORGED | 2010-04-19 |  | kwilson |  | Modify |  |  | Calorstat
11 | F061359/08 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R2 | 17 | O-ring groove width is out of tolerance. Scratches on both flanges must be fixed | GEORGED | 2010-04-19 |  |  |  | KWILSON |  | Dimensions ok. Please repair scratches as Aaron recommends. | Modify | [14 empty] | F061359/08 | GEORGED | 2010-04-19 |  | kwilson |  | Modify |  |  | Calorstat
12 | F063951/03 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R3 | 22 | Bellows has severe damage. | GEORGED | 2010-04-20 |  |  |  | KWILSON |  | Save to use for destructive testing. | Scrap | [14 empty] | F063951/03 | GEORGED | 2010-04-20 |  | kwilson |  | Scrap |  |  | Calorstat
13 | F063952-03 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R3 | 19 | Burrs on large flange. O-ring groove depth and o-ring groove width are out of to | georged | 2010-04-27 |  |  |  | KWILSON |  | Please repair. | Modify | [14 empty] | F063952-03 | georged | 2010-04-27 |  |  |  | Modify |  |  | Calorstat
14 | F063951-07 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R3 | 24 | O-ring groove depth and o-ring groove width are out of tolerance. Scratches on b | GEORGED | 2010-04-27 |  |  |  | KWILSON | 2010-09-15 | O-ring groove ok. | Use As Is | [14 empty] | F063951-07 | GEORGED | 2010-04-27 |  | kwilson | 2010-09-15 | Use As Is |  |  | Calorstat
15 | F063951-04 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R3 | 29 | O-ring groove width is out of tolerance. Scratches need repair on both flanges. | GEORGED | 2010-04-27 |  |  |  | KWILSON |  | Dimensions ok. Please repair scratches as Aaron recommends. | Modify | [14 empty] | F063951-04 | GEORGED | 2010-04-27 |  |  |  | Modify |  |  | Calorstat
16 | F063951-02 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R3 | 28 | O-ring groove depth and surface finish out of tolerance | carpente | 2011-01-12 |  |  |  | KWILSON |  |  | Use As Is | [14 empty] | F063951-02 | carpente | 2011-01-12 |  |  |  | Use As Is |  |  | Calorstat
18 | F063952-02 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R3 | 21 | Burrs and rough edges on knife edge and o-ring groove. O-ring groove width is ou | GEORGED | 2010-05-19 |  |  |  | KWILSON | 2010-09-15 | Repair burrs on knife edge. O-ring groove ok. | Modify | [14 empty] | F063952-02 | GEORGED | 2010-05-19 |  |  | 2010-09-15 | Modify |  |  | Calorstat
19 | F063952-01 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R3 | 26 | Two bolt holes diameter oversized on inner pattern of flange. Scratches on small | GEORGED | 2010-05-20 |  |  |  | KWILSON |  |  | Return To Vendor | [14 empty] | F063952-01 | GEORGED | 2010-05-20 |  |  |  | Return To Vendor |  |  | Calorstat
20 | F061359-12 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R2 | 13 | O-ring groove depth is out of tolerance. Nick on o-ring groove. Scratches on sm | carpente | 2010-05-27 |  |  |  | KWILSON |  | Repair scratches | Modify | [14 empty] | F061359-12 | carpente | 2010-05-27 |  |  |  | Modify |  |  | Calorstat
21 | F061359-12 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R2 | 13 | O-ring groove depth is out of tolerance. Dents inside waveguide and scratches a | georged | 2010-05-28 |  |  |  | KWILSON |  |  | Use As Is | [14 empty] | F061359-12 | georged | 2010-05-28 |  |  |  | Use As Is |  |  | Calorstat
22 | F061359-13 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R2 | 12 | Scratches on fpc flange. Copper flaking and dents inside waveguide. Discoloratio | GEORGED | 2010-05-28 |  |  |  | KWILSON |  | Damage to plating. Update: Part not usable. Consider scrap. Please do not throw away. Might be useful for testing. | Scrap | [14 empty] | F061359-13 | GEORGED | 2010-05-28 |  |  |  | Scrap |  |  | Calorstat
24 | CRM1207050000_025 |  | CRM1207050000 | C100 COLD TUNER |  | C100-CAV-INSP-CTUN | R1 | 27 | NOT CLEAN. NOT CONFORMING TO THE TRAVELER | WILCOX | 2010-06-09 |  |  |  | ABURRILL | 2011-09-01 | Returned to Vendor, gave instructions on proper cleaning technique. | Return To Vendor | [14 empty] | CRM1207050000_025 | WILCOX | 2010-06-09 |  |  | 2011-09-01 | Return To Vendor |  |  | ENCODEMA
25 | CRM1207050000_24 |  | CRM1207050000 | C100 COLD TUNER |  | C100-CAV-INSP-CTUN | R1 | 26 | NOT CLEAN. NOT CONFORMING TO THE TRAVELER | WILCOX | 2010-06-09 |  |  |  | ABURRILL | 2011-09-01 | Returned to Vendor, gave instructions on proper cleaning technique. | Return To Vendor | [14 empty] | CRM1207050000_24 | WILCOX | 2010-06-09 |  |  | 2011-09-01 | Return To Vendor |  |  | ENCODEMA
26 | CRM1207050000_023 |  | CRM1207050000 | C100 COLD TUNER |  | C100-CAV-INSP-CTUN | R1 | 25 | NOT CLEAN. NOT CONFORMING TO THE TRAVELER | WILCOX | 2010-06-09 |  |  |  | ABURRILL | 2011-09-01 | Returned to Vendor, gave instructions on proper cleaning technique. | Return To Vendor | [14 empty] | CRM1207050000_023 | WILCOX | 2010-06-09 |  |  | 2011-09-01 | Return To Vendor |  |  | ENCODEMA
27 | CRM1207050000_022 |  | CRM1207050000 | C100 COLD TUNER |  | C100-CAV-INSP-CTUN | R1 | 24 | NOT CLEAN. NOT CONFORMING TO THE TRAVELER | WILCOX | 2010-06-09 |  |  |  | ABURRILL | 2011-09-01 | Returned to Vendor, gave instructions on proper cleaning technique. | Return To Vendor | [14 empty] | CRM1207050000_022 | WILCOX | 2010-06-09 |  |  | 2011-09-01 | Return To Vendor |  |  | ENCODEMA
28 | CRM1207050000_029 |  | CRM1207050000 | C100 COLD TUNER |  | C100-CAV-INSP-CTUN | R1 | 23 | NOT CLEAN. NOT CONFORMING TO THE TRAVELER | WILCOX | 2010-06-09 |  |  |  | ABURRILL | 2011-09-01 | Returned to Vendor, gave instructions on proper cleaning technique. | Return To Vendor | [14 empty] | CRM1207050000_029 | WILCOX | 2010-06-09 |  |  | 2011-09-01 | Return To Vendor |  |  | ENCODEMA
29 | CRM1207050000_028 |  | CRM1207050000 | C100 COLD TUNER |  | C100-CAV-INSP-CTUN | R1 | 22 | NOT CLEAN. NOT CONFORMING TO THE TRAVELER | WILCOX | 2010-06-09 |  |  |  | ABURRILL |  | Returned to Vendor, gave instructions on proper cleaning technique. | Return To Vendor | [14 empty] | CRM1207050000_028 | WILCOX | 2010-06-09 |  |  |  | Return To Vendor |  |  | ENCODEMA
30 | CRM1207050000_026 |  | CRM1207050000 | C100 COLD TUNER |  | C100-CAV-INSP-CTUN | R1 | 28 | NOT CLEAN. NOT CONFORMING TO THE TRAVELER. | WILCOX | 2010-06-09 |  |  |  | ABURRILL | 2011-09-01 | Returned to Vendor, gave instructions on proper cleaning technique. | Return To Vendor | [14 empty] | CRM1207050000_026 | WILCOX | 2010-06-09 |  |  | 2011-09-01 | Return To Vendor |  |  | ENCODEMA
31 | CRM1207050000_030 |  | CRM1207050000 | C100 COLD TUNER |  | C100-CAV-INSP-CTUN | R1 | 29 | NOT CLEAN. NOT CONFORMING TO THE TRAVELER | WILCOX | 2010-06-09 |  |  |  | ABURRILL | 2011-09-01 | Returned to Vendor, gave instructions on proper cleaning technique. | Return To Vendor | [14 empty] | CRM1207050000_030 | WILCOX | 2010-06-09 |  |  | 2011-09-01 | Return To Vendor |  |  | ENCODEMA
32 | CRM1207050000_031 |  | CRM1207050000 | C100 COLD TUNER |  | C100-CAV-INSP-CTUN | R1 | 30 | NOT CLEAN. NOT CONFORMING TO THE TRAVELER | WILCOX | 2010-06-09 |  |  |  | ABURRILL | 2011-09-01 | Returned to Vendor, gave instructions on proper cleaning technique. | Return To Vendor | [14 empty] | CRM1207050000_031 | WILCOX | 2010-06-09 |  |  | 2011-09-01 | Return To Vendor |  |  | ENCODEMA
33 | CRM1207050000_032 |  | CRM1207050000 | C100 COLD TUNER |  | C100-CAV-INSP-CTUN | R1 | 31 | NOT CLEAN. NOT CONFORMING TO THE TRAVELER | WILCOX | 2010-06-09 |  |  |  | ABURRILL | 2011-09-01 | Returned to Vendor, gave instructions on proper cleaning technique. | Return To Vendor | [14 empty] | CRM1207050000_032 | WILCOX | 2010-06-09 |  |  | 2011-09-01 | Return To Vendor |  |  | ENCODEMA
34 | CRM1207050000_033 |  | CRM1207050000 | C100 COLD TUNER |  | C100-CAV-INSP-CTUN | R1 | 32 | NOT CLEAN. NOT CONFORMING TO THE TRAVELER | WILCOX | 2010-06-09 |  |  |  | ABURRILL |  | Returned to Vendor, gave instructions on proper cleaning technique. | Return To Vendor | [14 empty] | CRM1207050000_033 | WILCOX | 2010-06-09 |  |  |  | Return To Vendor |  |  | ENCODEMA
35 | CRM1207050000_034 |  | CRM1207050000 | C100 COLD TUNER |  | C100-CAV-INSP-CTUN | R1 | 33 | NOT CLEAN. NOT CONFORMING TO THE TRAVELER | WILCOX | 2010-06-09 |  |  |  | ABURRILL | 2011-09-01 | Returned to Vendor, gave instructions on proper cleaning technique. | Return To Vendor | [14 empty] | CRM1207050000_034 | WILCOX | 2010-06-09 |  |  | 2011-09-01 | Return To Vendor |  |  | ENCODEMA
36 | CRM1207050000_035 |  | CRM1207050000 | C100 COLD TUNER |  | C100-CAV-INSP-CTUN | R1 | 34 | NOT CLEAN. NOT CONFORMING TO THE TRAVELER | WILCOX | 2010-06-09 |  |  |  | ABURRILL | 2011-09-01 | Returned to Vendor, gave instructions on proper cleaning technique. | Return To Vendor | [14 empty] | CRM1207050000_035 | WILCOX | 2010-06-09 |  |  | 2011-09-01 | Return To Vendor |  |  | ENCODEMA
37 | CRM1207050000_036 |  | CRM1207050000 | C100 COLD TUNER |  | C100-CAV-INSP-CTUN | R1 | 35 | NOT CLEAN. NOT CONFORMING TO THE TRAVELER | WILCOX | 2010-06-09 |  |  |  | ABURRILL | 2011-09-01 | Returned to Vendor, gave instructions on proper cleaning technique. | Return To Vendor | [14 empty] | CRM1207050000_036 | WILCOX | 2010-06-09 |  |  | 2011-09-01 | Return To Vendor |  |  | ENCODEMA
38 | CRM1207050000_037 |  | CRM1207050000 | C100 COLD TUNER |  | C100-CAV-INSP-CTUN | R1 | 36 | NOT CLEAN. NOT CONFORMING TO THE TRAVELER | WILCOX | 2010-06-09 |  |  |  | ABURRILL | 2011-09-01 | Returned to Vendor, gave instructions on proper cleaning technique. | Return To Vendor | [14 empty] | CRM1207050000_037 | WILCOX | 2010-06-09 |  |  | 2011-09-01 | Return To Vendor |  |  | ENCODEMA
39 | CRM1207050000_038 |  | CRM1207050000 | C100 COLD TUNER |  | C100-CAV-INSP-CTUN | R1 | 37 | NOT CLEAN. NOT CONFORMING WITH THE TRAVELER | WILCOX | 2010-06-09 |  |  |  | ABURRILL | 2011-09-01 | Returned to Vendor, gave instructions on proper cleaning technique. | Return To Vendor | [14 empty] | CRM1207050000_038 | WILCOX | 2010-06-09 |  |  | 2011-09-01 | Return To Vendor |  |  | ENCODEMA
40 | CRM1207050000_039 |  | CRM1207050000 | C100 COLD TUNER |  | C100-CAV-INSP-CTUN | R1 | 38 | NOT CLEAN. NOT CONFORMING TO THE TRAVELER. | WILCOX | 2010-06-09 |  |  |  | ABURRILL | 2011-09-01 | Returned to Vendor, gave instructions on proper cleaning technique. | Return To Vendor | [14 empty] | CRM1207050000_039 | WILCOX | 2010-06-09 |  |  | 2011-09-01 | Return To Vendor |  |  | ENCODEMA
41 | CRM1207050000_040 |  | CRM1207050000 | C100 COLD TUNER |  | C100-CAV-INSP-CTUN | R1 | 39 | NOT CLEAN. NOT CONFORMING TO THE TRAVELER. | WILCOX | 2010-06-09 |  |  |  | ABURRILL | 2011-09-01 | Returned to Vendor, gave instructions on proper cleaning technique. | Return To Vendor | [14 empty] | CRM1207050000_040 | WILCOX | 2010-06-09 |  |  | 2011-09-01 | Return To Vendor |  |  | ENCODEMA
42 | CRM1207050000_041 |  | CRM1207050000 | C100 COLD TUNER |  | C100-CAV-INSP-CTUN | R1 | 40 | NOT CLEAN. NOT CONFORMING TO THE TRAVELER. | WILCOX | 2010-06-09 |  |  |  | ABURRILL | 2011-09-01 | Returned to Vendor, gave instructions on proper cleaning technique. | Return To Vendor | [14 empty] | CRM1207050000_041 | WILCOX | 2010-06-09 |  |  | 2011-09-01 | Return To Vendor |  |  | ENCODEMA
43 | F064155-04 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R3 | 42 | Tap didn't reach required depth. 3 bolt holes are incorrect. See traveler w/ | GEORGED | 2010-06-15 |  |  |  | KWILSON |  | Plesase retap if possible. Let me know if not. Thanks. | Modify | [14 empty] | F064155-04 | GEORGED | 2010-06-15 |  |  |  | Modify |  |  | Calorstat
101 | 1 |  | CRM-120-7080-1000 | Helium Header Parts |  | C100-CAV-INSP-HEHD | R1 | 1 | Parts are out of tolerance. See attached PDF for more information | GEORGED | 2010-06-24 |  |  |  | CHENG | 2010-10-20 | The out of tolerances on lengths can be absorbed by deforming the bellows. The l | Use As Is | [14 empty] | 1 | GEORGED | 2010-06-24 | CA211743-BE1B-545C-09813DD5D3696653 |  | 2010-10-20 | Use As Is |  | 6B9B35AA-BBAB-02BD-40345CB8B8C96B97
103 | 0008-2 |  | CRM12070420000 | Square Corner Head Left Half |  | C100-CAV-INSP-IMAG | R1 | 1 | Dimension out of tolerance. Meas. 1.905, Nominal is 2.00 | GEORGED | 2010-08-23 |  |  |  | CHENG | 2010-10-19 | The vendor is notified of this nonconformance. We can use the part. | Use As Is | [14 empty] | 0008-2 | GEORGED | 2010-08-23 |  |  | 2010-10-19 | Use As Is | The vendor proposed using a block of 2" height to verify this dimension. See pic | A4082985-FB09-8B40-CAF0BC36C730A757
104 | 0008-4 |  | CRM12070420008 | Square Corner Head Left Half |  | C100-CAV-INSP-IMAG | R1 | 2 | Dimension out of tolerance. Nominal is 2.00, meas. 1.95 | GEORGED | 2010-08-23 |  |  |  | CHENG | 2010-10-19 | The vendor is notified of this nonconformance. We can use the part as is. | Use As Is | [14 empty] | 0008-4 | GEORGED | 2010-08-23 |  |  | 2010-10-19 | Use As Is | The vendor proposed using a block of 2" height to verify this dimension in futur | C5A13C77-F27D-AAB9-0FD965B930097969
105 | 47789-003 |  | CRM-120-7075-0024 | C-100 Window adapter |  | C100-CAV-INSP-ADPT | R1 | 3 | Braze overflow in gap on back side of adapter. | CARPENTE | 2010-09-01 |  |  |  | HKPARK | 2010-09-01 | It is failed part but we'll proceed to RF test to see how the braze material bri | Use As Is | [14 empty] | 47789-003 | CARPENTE | 2010-09-01 |  |  | 2010-09-01 | Use As Is
106 | 47726-0015 |  | CRM-120-7014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R2 | 7 | Center connector pin off-center. Antenna length 1.459/1.469 out of tolerance. | [5 empty] | HKPARK | 2010-09-23 | If the center pin fits the gauge, center the pin. Mark the tolerance value from | Scrap | [14 empty] | 47726-0015 |  |  |  |  | 2010-09-23 | Scrap
107 | F063134-01 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R4 | 43 | Scratches on small flange.Large flange has burrs near knife edge. Nicks on knife | georged | 2010-09-02 |  |  |  | KWILSON |  | Repair scratches and burrs and nicks on knife edge. | Modify | [14 empty] | F063134-01 | georged | 2010-09-02 |  | carpente |  | Modify
108 | 47726-0016 |  | CRM-120-7014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R2 | 8 | Center connector pin off-center. Antenna length 1.459/1.469 out of tolerance. M | CARPENTE |  |  |  |  | HKPARK | 2010-11-10 | Use with R100-5 which can be matched with this probe tip length. | Use As Is | [14 empty] | 47726-0016 | CARPENTE |  |  |  | 2010-11-10 | Use As Is
109 | 47726-0017 |  | CRM-120-7014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R2 | 9 | Center connector pin off-center. Antenna length 1.459/1.469 out of tolerance. | CARPENTE |  |  |  |  | HKPARK | 2010-09-23 | If the pin fits in the gauge, center the pin carefully. If not, do not use. The | Use As Is | [14 empty] | 47726-0017 | CARPENTE |  |  |  | 2010-09-23 | Use As Is
110 | 47726-0018 |  | CRM-120-7014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R2 | 10 | Antenna length 1.459/1.469 out of tolerance. | CARPENTE |  |  |  |  | HKPARK | 2010-11-03 | Probe tip length tolerance +/-.010 is acceptable. | Use As Is | [14 empty] | 47726-0018 | CARPENTE |  |  |  | 2010-11-03 | Use As Is
111 | 47726-0019 |  | CRM-120-7014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R2 | 11 | Center connector pin off-center. Antenna length 1.459/1.469 out of tolerance. | CARPENTE |  |  |  |  | HKPARK | 2010-09-23 | If the pin fits in the gauge, center the pin carefully. If not, do not use. The | Use As Is | [14 empty] | 47726-0019 | CARPENTE |  |  |  | 2010-09-23 | Use As Is
112 | F063134-03 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R4 | 44 | Nicks on knife edge | GEORGED | 2010-09-02 |  |  |  | KWILSON |  |  | Return To Vendor | [14 empty] | F063134-03 | GEORGED | 2010-09-02 |  | carpente |  | Return To Vendor
113 | 47726-0020 |  | CRM-120-7014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R2 | 12 | Center connector pin slightly off-center. | CARPENTE |  |  |  |  | HKPARK | 2010-11-03 | Before use carefully center the pin. | Use As Is | [14 empty] | 47726-0020 | CARPENTE |  |  |  | 2010-11-03 | Use As Is
114 | 47726-0021 |  | CRM-120-7014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R2 | 13 | Center connector pin off-center. Antenna length 1.459/1.469 out of tolerance. | CARPENTE |  |  |  |  | HKPARK | 2010-11-10 | Use with R100-5 which can be matched with this probe tip length. | Use As Is | [14 empty] | 47726-0021 | CARPENTE |  |  |  | 2010-11-10 | Use As Is
115 | 47726-0022 |  | CRM-120-7014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R2 | 14 | Center connector pin fairly centered. Antenna length 1.459/1.469 out of toleran | CARPENTE |  |  |  |  | HKPARK | 2010-09-23 | If the pin fits in the gauge, center the pin carefully. If not, do not use. The | Use As Is | [14 empty] | 47726-0022 | CARPENTE |  |  |  | 2010-09-23 | Use As Is
116 | 47726-0023 |  | CRM-120-7014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R2 | 15 | Center connector pin off-center. Antenna length 1.459/1.469 out of tolerance. | CARPENTE |  |  |  |  | HKPARK | 2010-11-03 | The probe tip length is short. Tag the probe tip length. Use only when it can be | Use As Is | [14 empty] | 47726-0023 | CARPENTE |  |  |  | 2010-11-03 | Use As Is
117 | 47726-0024 |  | CRM-120-7014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R2 | 16 | Center connector pin off-center. | CARPENTE |  |  |  |  | HKPARK |  | If the center pin fits the gauge, center the pin carefully. | Use As Is | [14 empty] | 47726-0024 | CARPENTE |  |  |  |  | Use As Is
118 | 47726-0025 |  | CRM-120-7014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R2 | 17 | Center connector pin off-center. Antenna length 1.459/1.469 out of tolerance. | CARPENTE |  |  |  |  | HKPARK | 2010-11-03 | 1.464 +/- .010 is acceptable. Before use center the pin carefully. | Use As Is | [14 empty] | 47726-0025 | CARPENTE |  |  |  | 2010-11-03 | Use As Is
119 | 47726-0026 |  | CRM-120-7014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R2 | 20 | Center connector pin off-center. Antenna length 1.459/1.469 out of tolerance. | CARPENTE |  |  |  |  | HKPARK | 2010-11-03 | Probe tip length .001 shorter than min. Tag the part with the measured length. U | Use As Is | [14 empty] | 47726-0026 | CARPENTE |  |  |  | 2010-11-03 | Use As Is
120 | 47726-0027 |  | CRM-120-7014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R2 | 18 | Center connector pin off-center. Antenna length 1.459/1.469 out of tolerance. | CARPENTE |  |  |  |  | HKPARK | 2010-09-23 | If the pin fits in the gauge, center the pin carefully. If not, do not use. The | Use As Is | [14 empty] | 47726-0027 | CARPENTE |  |  |  | 2010-09-23 | Use As Is
121 | 47726-0028 |  | CRM-120-7014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R2 | 19 | Center connector pin off-center. Antenna length 1.459/1.469 out of tolerance. | CARPENTE |  |  |  |  | HKPARK | 2010-09-23 | If the pin fits in the gauge, center the pin carefully. If not, do not use. The | Use As Is | [14 empty] | 47726-0028 | CARPENTE |  |  |  | 2010-09-23 | Use As Is
122 | 47726-0029 |  | CRM-120-7014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R2 | 21 | Center connector pin off-center. Antenna length 1.459/1.469 out of tolerance. | CARPENTE |  |  |  |  | HKPARK | 2010-09-23 | If the pin fits in the gauge, center the pin carefully. If not, do not use. The | Use As Is | [14 empty] | 47726-0029 | CARPENTE |  |  |  | 2010-09-23 | Use As Is
123 | 47726-0031 |  | CRM-120-7014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R2 | 22 | Center connector pin off-center. Antenna length 1.459/1.469 out of tolerance. | CARPENTE |  |  |  |  | HKPARK | 2010-09-23 | If the pin fits in the gauge, center the pin carefully. If not, do not use. The | Use As Is | [14 empty] | 47726-0031 | CARPENTE |  |  |  | 2010-09-23 | Use As Is
124 | 47726-0032 |  | CRM-120-7014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R2 | 23 | Center connector pin off-center. Antenna length 1.459/1.469 out of tolerance. | CARPENTE |  |  |  |  | HKPARK | 2010-09-29 |  | Scrap | [14 empty] | 47726-0032 | CARPENTE |  |  |  | 2010-09-29 | Scrap
125 | 47726-0033 |  | CRM-120-7014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R2 | 24 | Center connector pin off-center. | CARPENTE |  |  |  |  | HKPARK |  | If the center pin fits the gauge, center the pin carefully. | Use As Is | [14 empty] | 47726-0033 | CARPENTE |  |  |  |  | Use As Is
126 | 47726-0034 |  | CRM-120-7014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R2 | 25 | Center connector pin very slightly off-center. Antenna length 1.459/1.469 out o | CARPENTE |  |  |  |  | HKPARK | 2010-11-03 | Probe tip length too short. Quarantine. | Scrap | [14 empty] | 47726-0034 | CARPENTE |  |  |  | 2010-11-03 | Scrap
127 | 47726-0035 |  | CRM-120-7014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R2 | 26 | Center connector pin off-center. Antenna length 1.459/1.469 out of tolerance. | CARPENTE |  |  |  |  | HKPARK | 2010-09-23 | If the pin fits in the gauge, center the pin carefully. If not, do not use. The | Use As Is | [14 empty] | 47726-0035 | CARPENTE |  |  |  | 2010-09-23 | Use As Is
128 | 47726-0036 |  | CRM-120-7014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R2 | 27 | Center connector pin off-center. Antenna length 1.459/1.469 out of tolerance. | CARPENTE |  |  |  |  | HKPARK | 2010-11-03 | Probe tip too short. Do not use. Qrantine. | Scrap | [14 empty] | 47726-0036 | CARPENTE |  |  |  | 2010-11-03 | Scrap
129 | 47726-0037 |  | CRM-120-7014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R2 | 28 | Center connector pin off-center. Antenna length 1.459/1.469 out of tolerance. | OGLE |  |  |  |  | HKPARK | 2010-11-03 | Probe length too short. Qurantine. | Scrap | [14 empty] | 47726-0037 | OGLE |  |  |  | 2010-11-03 | Scrap
130 | 47726-0038 |  | CRM-120-7014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R2 | 29 | Center connector pin very slightly off-center. Antenna length 1.459/1.469 out o | OGLE |  |  |  |  | HKPARK | 2010-09-23 | If the pin fits in the gauge, center the pin carefully. If not, do not use. The | Use As Is | [14 empty] | 47726-0038 | OGLE |  |  |  | 2010-09-23 | Use As Is
131 | 47726-0039 |  | CRM-120-7014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R2 | 30 | Center connector pin off-center. Antenna length 1.459/1.469 out of tolerance. | OGLE |  |  |  |  | HKPARK |  | If the pin fits in the gauge, center the pin carefully. If not, do not use. The | Use As Is | [14 empty] | 47726-0039 | OGLE |  |  |  |  | Use As Is
132 | 47726-0041 |  | CRM-120-7014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R2 | 32 | Center connector pin off-center. | OGLE |  |  |  |  | HKPARK | 2010-09-23 | If the center pin fits the gauge, center the pin carefully. | Use As Is | [14 empty] | 47726-0041 | OGLE |  |  |  | 2010-09-23 | Use As Is
133 | F063134-05 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R4 | 45 | Waveguide flange has a nick on the knife edge | GEORGED | 2010-09-02 |  |  |  | KWILSON |  |  | Return To Vendor | [14 empty] | F063134-05 | GEORGED | 2010-09-02 |  | carpente |  | Return To Vendor
134 | F063134-06 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R4 | 46 | Nick on knife edge and damage in o-ring groove | GEORGED | 2010-09-02 |  |  |  | KWILSON |  |  | Return To Vendor | [14 empty] | F063134-06 | GEORGED | 2010-09-02 |  | carpente |  | Return To Vendor
135 | F063134-07 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R4 | 47 | Nick on knife edge | GEORGED | 2010-09-02 |  |  |  | KWILSON |  |  | Use As Is | [14 empty] | F063134-07 | GEORGED | 2010-09-02 |  | carpente |  | Use As Is | WG was subsequently reworked, knife edge was repaired. See traveler: C100-CAV-R
136 | F064155-02 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R3 | 40 | Leak check #3 failed, scratches on small flange and surface finish out of tolera | georged | 2010-09-14 |  |  |  | KWILSON |  | Return to vendor due to failed leak check. | Return To Vendor | [14 empty] | F064155-02 | georged | 2010-09-14 |  | carpente |  | Return To Vendor
138 | F061584-02 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R3 | 39 | Surface finish out of tolerance, scratches on small flange | GEORGED | 2010-09-15 |  |  |  | KWILSON |  |  | Use As Is | [14 empty] | F061584-02 | GEORGED | 2010-09-15 |  | carpente |  | Use As Is
139 | F063950-01 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R3 | 38 | Surface finish and o-ring depth out of tolerance. Scratches on small flg | georged | 2010-09-15 |  |  |  | KWILSON |  | Surface finish and o-ring groove ok. Repair scratches. | Modify | [14 empty] | F063950-01 | georged | 2010-09-15 |  | carpente |  | Modify
140 | F064155-03 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R3 | 37 | Scratches on small flg, discoloration and stains inside waveguide | georged | 2010-09-15 |  |  |  | KWILSON |  | Repair scratches. | Modify | [14 empty] | F064155-03 | georged | 2010-09-15 |  | carpente |  | Modify
141 | F064155-01 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R3 | 36 | O-ring depth out of tolerance, repairable lite scratches on small flg. Very smal | carpente | 2010-12-03 |  |  |  | KWILSON |  | O-ring groove ok. Repair scratches. Clean up plating as required. | Modify | [14 empty] | F064155-01 | carpente | 2010-12-03 |  |  |  | Modify
142 | F061584-03 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R3 | 35 | Surface finish and o-ring groove depth are out of tolerance, and repairable sc | georged | 2010-09-15 |  |  |  | KWILSON |  | Repair scratches. Which probably results in better surface finish too. Travele | Modify | [14 empty] | F061584-03 | georged | 2010-09-15 |  |  |  | Modify
143 | F063950-02 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R3 | 30 | Surface finish out of tolerance, measured 20 m inch | GEORGED | 2010-09-16 |  |  |  | KWILSON |  |  | Use As Is | [14 empty] | F063950-02 | GEORGED | 2010-09-16 | 5B501C47-C312-C73A-F1423E8327727FEE |  |  | Use As Is
144 | C100-RI-003 |  | CRM1207013-0001 | C100 Cavity |  | C100-CAV-INSP | R2 | 4 | please see attached file with "NCR" postfix. | CARPENTE | 2010-09-22 |  |  |  | MARHAUSE |  |  | Use As Is | [14 empty] | C100-RI-003 | CARPENTE | 2010-09-22 | 3A6D24E1-9C8C-DD36-089CF9A49EDE0B30 |  |  | Use As Is
145 | C100-RI-004 |  | CRM1207013-0001 | C100 Cavity |  | C100-CAV-INSP | R2 | 5 | NCR created in error. | CARPENTE | 2010-09-22 |  |  |  | CARPENTE |  | NCR created in error. | Use As Is | [14 empty] | C100-RI-004 | CARPENTE | 2010-09-22 | 3A78CB88-9789-B9A0-9D94D94ED386E19B |  |  | Use As Is
146 | C100-RI-006 |  | CRM1207013-0001 | C100 Cavity |  | C100-CAV-INSP | R2 | 8 | Please see attached file with "NCR" postfix. | CARPENTE | 2010-09-23 |  |  |  | MARHAUSE |  |  | Remeasure | [14 empty] | C100-RI-006 | CARPENTE | 2010-09-23 | 3F2DA43F-C99C-E92E-CB4D4480D38445E1 |  |  | Remeasure | Cavity re-measured, out of tolerance issues corrected. BC 09/16/11
147 | C100-RI-008 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R2 | 10 | VOIDS ON FLANGE FACE:BOTH FPC AND HOM ENDS | [5 empty] | ABURRILL | 2011-09-19 | Monitor cavity RF test for signs of virtual leak or excessive field emission, us | Use As Is | [14 empty] | C100-RI-008 |  |  |  |  | 2011-09-19 | Use As Is
148 | F067763-14 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R4 | 49 | 15 of 16 outside pattern tapped holes not deep enough. | GEORGED | 2010-09-27 |  |  |  | CARPENTE | 2011-01-11 | holes were re-checked with fixture. All holes now ok. | Use As Is | [14 empty] | F067763-14 | GEORGED | 2010-09-27 |  |  | 2011-01-11 | Use As Is
149 | F067763-19 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R4 | 50 | 16 of 16 tapped holes not deep enough on outside pattern. O-ring groove depth o | GEORGED | 2010-09-27 |  |  |  | KWILSON |  | Retap holes please. | Modify | [14 empty] | F067763-19 | GEORGED | 2010-09-27 |  | carpente |  | Modify
150 | F067763-20 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R4 | 52 | 15 of 16 tapped holes not deep enough on outside pattern | GEORGED | 2010-09-27 |  |  |  | KWILSON |  | Retap holes please | Modify | [14 empty] | F067763-20 | GEORGED | 2010-09-27 |  | carpente |  | Modify
151 | C100-RI-010 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R2 | 11 | BURR PRESENT ON FPC FACE AT HOLE EDGE | OGLE | 2010-09-27 |  |  |  | MARHAUSE | 2011-03-29 |  | Use As Is | [14 empty] | C100-RI-010 | OGLE | 2010-09-27 |  |  | 2011-03-29 | Use As Is
152 | F067763-07 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R4 | 54 | Depth fixture used, depth of one hole bad. O-ring groove depth out of tolerance | georged | 2010-09-28 |  |  |  | KWILSON |  | O-ring groove ok. Surface finish ok. Repair scratches. Retap hole with issue. | Modify | [14 empty] | F067763-07 | georged | 2010-09-28 |  |  |  | Modify
153 | F067763-01 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R4 | 56 | O-ring groove depth out of tolerance 0.001, No Go gauge screws in completely 1 t | GEORGED | 2010-09-28 |  |  |  | KWILSON |  |  | Return To Vendor | [14 empty] | F067763-01 | GEORGED | 2010-09-28 |  |  |  | Return To Vendor
154 | 47726-0043 |  | CRM12070140020 | HOM FEED THRU PROBE |  | C100-CAV-INSP-HMFT | R3 | 34 | LENGTH 1.459/1.469 CHECKS 1.454; OUT OF TOL .005 | OGLE | 2010-09-28 |  |  |  | HKPARK | 2010-09-29 | The tolerance is still in reasonable range. By matching during the cavity assemb | Use As Is | [14 empty] | 47726-0043 | OGLE | 2010-09-28 |  |  | 2010-09-29 | Use As Is
155 | 47726-0044 |  | CRM12070140020 | HOM FEED THRU PROBE |  | C100-CAV-INSP-HMFT | R3 | 35 | CONNECTOR CENTER PIN FAILS FUNCTIONAL GAGE; LENGTH 1.459/1.469 CHECKS 1.454; OUT | OGLE | 2010-09-28 |  |  |  | HKPARK | 2010-12-02 | Use if the pin can be centered and connected to the connector. The tip length o | Use As Is | [14 empty] | 47726-0044 | OGLE | 2010-09-28 |  |  | 2010-12-02 | Use As Is
156 | 47726-0045 |  | CRM12070140020 | HOM FEED THRU PROBE |  | C100-CAV-INSP-HMFT | R3 | 36 | LENGTH 1.459/1.469 CHECKS 1.457; OUT OF TOL .002 | OGLE | 2010-09-28 |  |  |  | HKPARK | 2010-09-29 | The tolerance is still in reasonable range. By matching during the cavity assemb | Use As Is | [14 empty] | 47726-0045 | OGLE | 2010-09-28 |  |  | 2010-09-29 | Use As Is
157 | 47726-0047 |  | CRM12070140020 | HOM FEED THRU PROBE |  | C100-CAV-INSP-HMFT | R3 | 38 | LENGTH 1.459/1.469 CHECKS 1.454 OUT OF TOL .005 | OGLE | 2010-09-28 |  |  |  | HKPARK | 2010-09-29 | The tolerance is still in reasonable range. By matching during the cavity assemb | Use As Is | [14 empty] | 47726-0047 | OGLE | 2010-09-28 |  |  | 2010-09-29 | Use As Is
158 | 47726-0048 |  | CRM12070140020 | HOM FEED THRU PROBE |  | C100-CAV-INSP-HMFT | R3 | 39 | LENGTH 1.459/1.469 CHECKS 1.453; OUT OF TOL .006 | OGLE | 2010-09-28 |  |  |  | HKPARK | 2010-12-02 | Use ONLY with matching HOM can. | Use As Is | [14 empty] | 47726-0048 | OGLE | 2010-09-28 |  |  | 2010-12-02 | Use As Is
159 | 47726-0050 |  | CRM12070140020 | HOM FEED THRU PROBE |  | C100-CAV-INSP-HMFT | R3 | 41 | LENGTH 1.459/1.469 CHECKS 1.455; OUT OF TOL. .004 | OGLE | 2010-09-28 |  |  |  | HKPARK | 2010-09-29 | The tolerance is still in reasonable range. By matching during the cavity assemb | Use As Is | [14 empty] | 47726-0050 | OGLE | 2010-09-28 |  |  | 2010-09-29 | Use As Is
160 | 47726-0051 |  | CRM12070140020 | HOM FEED THRU PROBE |  | C100-CAV-INSP-HMFT | R3 | 42 | LENGTH 1.459/1.469 CHECKS 1.456; OUT OF TOL..003 | OGLE | 2010-09-28 |  |  |  | HKPARK | 2010-09-29 | The tolerance is still in reasonable range. By matching during the cavity assemb | Use As Is | [14 empty] | 47726-0051 | OGLE | 2010-09-28 |  |  | 2010-09-29 | Use As Is
161 | 47726-0049 |  | CRM12070140020 | HOM FEED THRU PROBE |  | C100-CAV-INSP-HMFT | R3 | 40 | LENGTH 1.459/1.469 CHECKS 1.454; OUT OF TOL .005 | OGLE | 2010-09-28 |  |  |  | HKPARK | 2010-09-29 | The tolerance is still in reasonable range. By matching during the cavity assemb | Use As Is | [14 empty] | 47726-0049 | OGLE | 2010-09-28 |  |  | 2010-09-29 | Use As Is
162 | 47726-0042 |  | CRM120-7014-0020 | hom feed thru |  | C100-CAV-INSP-HMFT | R3 | 33 | LENGTH 1.459/1.469 CHECKS 1.457;OUT OF TOL. .002 | OGLE | 2010-09-28 |  |  |  | HKPARK | 2010-09-29 | The tolerance is still in reasonable range. By matching during the cavity assemb | Use As Is | [14 empty] | 47726-0042 | OGLE | 2010-09-28 |  |  | 2010-09-29 | Use As Is
163 | F067763-18 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R4 | 58 | Repairable Light scratches on small flange, will need repairing. O-ring groove | georged | 2010-09-29 |  |  |  | KWILSON |  | Repair scratches, p-ring groove ok | Modify | [14 empty] | F067763-18 | georged | 2010-09-29 |  |  |  | Modify
164 | F067764-04 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R4 | 59 | 16 of 16 outer pattern threaded holes not deep enough | GEORGED | 2010-09-29 |  |  |  | CARPENTE | 2011-01-04 | Bolt hole depth fixture was used after initial inspection. Holes passed. | Use As Is | [14 empty] | F067764-04 | GEORGED | 2010-09-29 |  |  | 2011-01-04 | Use As Is
165 | F067763-15 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R4 | 60 | Repairable scratches on small flange will need lapping | carpente | 2010-12-03 |  |  |  | KWILSON |  | Repair scratches | Modify | [14 empty] | F067763-15 | carpente | 2010-12-03 |  |  |  | Modify
166 | F067763-06 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R4 | 61 | 16 of 16 outer pattern holes not deep enough | GEORGED | 2010-09-29 |  |  |  | KWILSON |  | Retap holes please | Modify | [14 empty] | F067763-06 | GEORGED | 2010-09-29 |  |  |  | Modify
167 | F067764-06 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R4 | 62 | Light scratches on small flange | georged | 2010-11-19 |  |  |  | KWILSON |  | Repair scratches. | Modify | [14 empty] | F067764-06 | georged | 2010-11-19 |  |  |  | Modify
168 | F067764-12 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R4 | 63 | NoGo side of gauge screws completely down into two bolt holes | GEORGED | 2010-10-19 |  |  |  | KWILSON |  |  | Return To Vendor | [14 empty] | F067764-12 | GEORGED | 2010-10-19 |  |  |  | Return To Vendor
169 | F067763-11 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R4 | 64 | 1 of 16 outer pattern holes not deep enough | GEORGED | 2010-09-29 |  |  |  | KWILSON |  | Please retap. | Modify | [14 empty] | F067763-11 | GEORGED | 2010-09-29 |  |  |  | Modify
170 | C100-RI-011 |  | CRM1207013-0001 | C100 Cavity |  | C100-CAV-INSP | R2 | 13 | Overall length and FPC end flange true position are out of tolerance | GEORGED | 2010-10-12 |  |  |  | MARHAUSE | 2010-10-15 | On cavity RI011 (as for RI012) the final length is 2 mm smaller than expected. T | Use As Is | [14 empty] | C100-RI-011 | GEORGED | 2010-10-12 |  |  | 2010-10-15 | Use As Is | On cavity RI011 (as for RI012) the final length is 2 mm smaller than expected. T
171 | C100-RI-012 |  | CRM120-7014-0020 | CAVITY |  | C100-CAV-INSP | R2 | 14 | END FLANGE TRUE POSITION FPC END; REQUIRED .060 RFS;CKS. .111 OUT OF TOL .051 | OGLE | 2010-10-12 |  |  |  | MARHAUSE | 2010-10-15 | On cavity RI012 (as for RI011) the final length is 2 mm smaller than expected. T | Use As Is | [14 empty] | C100-RI-012 | OGLE | 2010-10-12 |  |  | 2010-10-15 | Use As Is | On cavity RI012 (as for RI011) the final length is 2 mm smaller than expected. T
172 | FLG0037-021 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R1 | 11 | FLANGE THICKNES 1.360 CKS. 1.366 OUT/TOL .001 | OGLE | 2010-10-14 |  |  |  | HKPARK | 2010-10-15 | Out tol value not affecting assembly | Use As Is | [14 empty] | FLG0037-021 | OGLE | 2010-10-14 |  |  | 2010-10-15 | Use As Is
173 | FLG0037-022 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R1 | 12 | FLANGE THICKNESS 1.360 CKS 1.368 OUT/TOL. .003 | OGLE | 2010-10-14 |  |  |  | HKPARK | 2010-10-15 | Out tol range not significant and critical to the performance | Use As Is | [14 empty] | FLG0037-022 | OGLE | 2010-10-14 |  |  | 2010-10-15 | Use As Is
174 | FLG0037-023 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R1 | 13 | FLANGE THICKNESS 1.360 CKS. 1.367 OUT/TOL .002 | OGLE | 2010-10-14 |  |  |  | HKPARK | 2010-10-15 | Out tol range not significant or critical to the performance | Use As Is | [14 empty] | FLG0037-023 | OGLE | 2010-10-14 |  |  | 2010-10-15 | Use As Is
175 | FLG0037-024 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R1 | 14 | FLANGE THICKNESS 1.360 CKS. 1.367 OUT/TOL .002 | OGLE | 2010-10-15 |  |  |  | HKPARK | 2010-10-15 | Out tol range not affecting the performance | Use As Is | [14 empty] | FLG0037-024 | OGLE | 2010-10-15 |  |  | 2010-10-15 | Use As Is
176 | FLG0037-026 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R1 | 16 | FLANGE THICKNESS 1.360 CKS. 1.366 OUT/TOL .001 | OGLE | 2010-10-15 |  |  |  | HKPARK | 2010-11-24 | Out tol level acceptable. | Use As Is | [14 empty] | FLG0037-026 | OGLE | 2010-10-15 |  |  | 2010-11-24 | Use As Is
177 | FLG0037-027 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R1 | 17 | FLANGE THICKNESS 1.360 CKS. 1.366 OUT/TOL .001 CIRC 17 X 3.435 CKS 3.441 OUT/TO | OGLE | 2010-10-15 |  |  |  | HKPARK | 2010-10-19 | Out of tolerance range acceptable | Use As Is | [14 empty] | FLG0037-027 | OGLE | 2010-10-15 |  |  | 2010-10-19 | Use As Is
178 | FLG0037-028 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R1 | 18 | NICK ON KNIFE EDGE. FLANGE THICKNESS 1.360 CKS. 1.368 OUT/TOL .003 CIRC. 18 Z | OGLE | 2010-10-15 |  |  |  | HKPARK | 2010-12-09 | Tag and store. They should be returned to the vendors for the replacement. | Scrap | [14 empty] | FLG0037-028 | OGLE | 2010-10-15 |  |  | 2010-12-09 | Scrap
179 | FLG0037-029 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R1 | 19 | NICK ON KNIFE EDGE FLANGE THICKNESS 1.360 CKS. 1.367 OUT/TOL .002 | OGLE | 2010-10-15 |  |  |  | HKPARK | 2010-12-09 | Tag and store. It will be notified to the vendor for the replacement or credit. | Scrap | [14 empty] | FLG0037-029 | OGLE | 2010-10-15 |  |  | 2010-12-09 | Scrap
180 | C100-RI-007 |  | INSP-CRM1207013-0001 | C100 Cavity |  | C100-CAV-INSP | R2 | 9 | See attached file | CARPENTE | 2010-10-18 |  |  |  | MARHAUSE | 2011-03-29 | polish surfaces adequately. | Modify | [14 empty] | C100-RI-007 | CARPENTE | 2010-10-18 | C00BDC91-F8DF-8370-0D94DA87D62A372F |  | 2011-03-29 | Modify | Issued resolved in C100-CAV-NCR, seq. # 251 BC 9/16/11
181 | MPF0004 |  | CRM-120-7075-0024 A | C100 Brazement Adapter |  | C100-CAV-INSP-ADPT | R2 | 5 | Brazement failed intial RF test. Hot spot at one end of ceramic with temperatur |  | 2010-10-19 |  |  |  | HKPARK | 2010-10-19 | Do not use. The material analysis is planned for this adapter. | Scrap | [14 empty] | MPF0004 |  | 2010-10-19 |  |  | 2010-10-19 | Scrap
182 | F067763-01 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R4 | 181 | Depth fixure used, depth bad on one inner hole. NoGo gauge screws completely dow | georged | 2010-10-22 |  |  |  | CARPENTE | 2011-09-16 | RTV as per C100-CAV-NCR, seq. # 153 and 448 BC 09/16/11 | Return To Vendor | [14 empty] | F067763-01 | georged | 2010-10-22 |  | kwilson | 2011-09-16 | Return To Vendor
183 | F064155-04 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R4 | 183 | Depth fixture used, one hole failed. Depth is bad | georged | 2010-10-22 |  |  |  | KWILSON |  | Please retap. | Modify | [14 empty] | F064155-04 | georged | 2010-10-22 |  | kwilson |  | Modify
184 | 12070750027-10/25/10 |  | CRM1207075-0027 | Window Adapter Detail |  |  |  |  | See NCR Report attachment for detail dimensional non-conforming. | ogle | 2010-10-25 |  |  | Inspection results: Window Adapter Detail. 13 of 20 failed for dimensional insp | HKPARK | 2010-10-27 | Use 21, 22, 24, 25, 28, 31, 32, 34, 35 39, and 40 DO NOT USE 29 and 36. | Use As Is | [14 empty] | 12070750027-10/25/10 | ogle | 2010-10-25 | E6594E47-EF32-0D62-D2234D745FE40F2E | hkpark |  | Use As Is
185 | C100-RI-015 |  | CRM1207013-0001 | C100 Cavity |  | C100-CAV-INSP | R2 | 15 |  | carpente | 2010-10-28 |  |  |  | MARHAUSE |  |  | Use As Is | [14 empty] | C100-RI-015 | carpente | 2010-10-28 | F2B5F484-ED2E-7361-3D735E4D8D4B7A14 |  |  | Use As Is
186 | C100-RI-017 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R2 | 16 | See attachment for dimensional non-conforming. | ogle | 2010-10-29 |  |  |  | MARHAUSE |  |  | Use As Is | [14 empty] | C100-RI-017 | ogle | 2010-10-29 | F9786CAE-A645-5854-6CF28FA9B803B61F |  |  | Use As Is
187 | ISI-0001 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | C100-CAV-INSP-ADPT | R2 | 6 | SEE ATTACHMENT FOR NON CONFORMING REPORT | ogle | 2010-11-01 |  |  |  | HKPARK | 2010-11-02 | Do Not use but proceed to the leak test and RF test. | Scrap | [14 empty] | ISI-0001 | ogle | 2010-11-01 | 09FCCDB0-000A-B7DF-3B5B2AE8FD232ABE |  | 2010-11-02 | Scrap
188 | ISI-0002 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | C100-CAV-INSP-ADPT | R2 | 7 | SEE ATTACHMENT FOR NON CONFORMING REPORT. | ogle | 2010-11-01 |  |  |  | HKPARK | 2010-11-02 | Do Not use but proceed to the leak test and RF test. | Scrap | [14 empty] | ISI-0002 | ogle | 2010-11-01 | 0A06A1C2-A5E7-522B-988205F2071CAFD2 |  | 2010-11-02 | Scrap
189 | ISI-0003 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | C100-CAV-INSP-ADPT | R2 | 8 | SEE ATTACHMENT FOR NON CONFORMING REPORT DETAIL | ogle | 2010-11-01 |  |  |  | HKPARK | 2010-11-02 | Do Not use but proceed to the leak test and RF test. | Scrap | [14 empty] | ISI-0003 | ogle | 2010-11-01 | 0A0F93E8-E3AE-D75F-AFD10E71935696F9 |  | 2010-11-02 | Scrap
190 | F067763-20 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R4 | 52 | Copper flaking inside waveguide. | georged | 2010-11-04 |  |  |  | KWILSON |  |  | Return To Vendor | [14 empty] | F067763-20 | georged | 2010-11-04 |  | carpente |  | Return To Vendor
191 | C100-RI-014 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R2 | 18 | RAISED MATERIAL ON END FLANGE FACE (FPC END) SCRATCHES ON FIELD PROBE FACE. SE | ogle | 2010-11-04 |  |  |  | MARHAUSE | 2011-03-29 | the blemishes detected require appropriate polishing on the effected surfaces wi | Modify | [14 empty] | C100-RI-014 | ogle | 2010-11-04 | 19CAF8AF-D3CD-90B1-EEF5748FC6450E60 |  | 2011-03-29 | Modify
192 | C100-RI-018 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R2 | 19 | SEE ATTACHMENT FOR DIMENSIONAL NON- CONFORMING DIMENSIONS. | ogle | 2010-11-05 |  |  |  | MARHAUSE | 2011-03-29 |  | Use As Is | [14 empty] | C100-RI-018 | ogle | 2010-11-05 | 1F32ED31-CAA7-467D-9B5D2578D51BECB3 |  | 2011-03-29 | Use As Is
193 | C100-RI-019 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R2 | 20 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | ogle | 2010-11-08 |  |  |  | MARHAUSE | 2011-03-29 |  | Use As Is | [14 empty] | C100-RI-019 | ogle | 2010-11-08 | 2E33DA3F-B10A-F91C-E54E67AA35525344 |  | 2011-03-29 | Use As Is
194 | ISI-0001 |  | CRM120-7014-0020 | hom feed thru |  | C100-CAV-INSP-HMFT | R3 | 44 | DEBRIS INTERNAL CONNECTOR END ON CERAMIC SURFACE. TOOL CHATTER ON SEAL SURFACE. | ogle | 2010-11-10 |  |  |  | HKPARK | 2010-11-11 | Use if the debris can be removed. The effect of chatter on seal surface should b | Use As Is | [14 empty] | ISI-0001 | ogle | 2010-11-10 | 3907D9E1-B46C-3D7A-F761118FBDC4F8A2 |  | 2010-11-11 | Use As Is
195 | ISI-0002 |  | CRM120-7014-0020 | hom feed thru |  | C100-CAV-INSP-HMFT | R3 | 45 | DEBRIS FOUND INTERNAL/CONNECTOR END NEAR CERAMIC SURFACE.TOOL CHATTER ON SEAL SU | ogle | 2010-11-10 |  |  |  | HKPARK | 2010-11-11 | Remove the debris. The probe tip length is acceptable. | Use As Is | [14 empty] | ISI-0002 | ogle | 2010-11-10 | 390CD6E3-0734-E9DA-8975B1FB2D8431F6 |  | 2010-11-11 | Use As Is
196 | ISI-0003 |  | CRM120-7014-0020 | hom feed thru |  | C100-CAV-INSP-HMFT | R3 | 46 | DEBRIS FOUND INTERNAL/CONNECTOR END NEAR CERAMIC SURFACE.TOOL CHATTER ON SEAL SU | ogle | 2010-11-10 |  |  |  | HKPARK | 2010-11-11 | Remove the debris. Proceed to the leak check. Leak check will test the effect of | Use As Is | [14 empty] | ISI-0003 | ogle | 2010-11-10 |  |  | 2010-11-11 | Use As Is
197 | F067763-04 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R4 | 53 | Leak check redone, it passed. scratches on small flange, nick on knife edge, o-r | georged | 2010-11-11 |  |  |  | KWILSON |  | O-ring groove ok. Repair scratches on flange and nick on knife edge. | Modify | [14 empty] | F067763-04 | georged | 2010-11-11 |  |  |  | Modify
198 | C100-RI-005 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP2 | R2 | 3 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | ogle | 2010-11-11 |  |  |  | MARHAUSE |  |  | Use As Is | [14 empty] | C100-RI-005 | ogle | 2010-11-11 | 3DAAE393-0241-A2DC-81EE1EB8D9B626EC |  |  | Use As Is
199 | C100-RI-006 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP2 | R2 | 4 | SEE ATTACHMENT FOR DIMENSIONAL CON-CONFORMING | ogle | 2010-11-11 |  |  |  | MARHAUSE | 2010-11-12 |  | Use As Is | [14 empty] | C100-RI-006 | ogle | 2010-11-11 | 3DB92B73-0445-1620-CF2046AD8D6A91B1 |  | 2010-11-12 | Use As Is
200 | C100-RI-020 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 22 | LIGHT SCRATCH ON FACE. DIS-COLORATION SPOT ON FACE. SEE ATTACHMENT FOR DIMENSIO | ogle | 2010-11-12 |  |  |  | MARHAUSE | 2010-11-18 | Lap FPC flange for remedy. | Use As Is | [14 empty] | C100-RI-020 | ogle | 2010-11-12 | 42D7B68F-97EA-AE14-DA65AFD5DEC3CBD8 |  | 2010-11-18 | Use As Is
201 | C100-RI-021 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 23 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. DAMAGE FOUND ON DATUM DIAMETER O | ogle | 2010-11-12 |  |  |  | MARHAUSE | 2010-11-18 |  | Use As Is | [14 empty] | C100-RI-021 | ogle | 2010-11-12 | 436EE39F-E0CE-8F68-31E7206AF1C927B3 |  | 2010-11-18 | Use As Is
202 | C100-RI-004 |  | CRM1207013-0001 | C100 Cavity |  | C100-CAV-INSP2 | R3 | 6 | The NCR report contains all out of tolerance measurements, the dish diameter mea | carpente | 2010-11-15 |  |  |  | MARHAUSE | 2010-11-15 |  | Use As Is | [14 empty] | C100-RI-004 | carpente | 2010-11-15 | 4FB62870-079C-A6EA-2DD86F2EC1E17851 |  | 2010-11-15 | Use As Is
203 | F061359-02 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R2 | 15 | O-ring groove width out of tolerance 0.001 scratches on small flange, repair in | georged | 2010-11-15 |  |  |  | KWILSON |  | Please repair scratches on flange. | Modify | [14 empty] | F061359-02 | georged | 2010-11-15 |  |  |  | Modify
204 | F067763-08 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R4 | 48 | O-ring groove depth out of tolerance 0.003, bad scratches on small flange | georged | 2010-11-15 |  |  |  | KWILSON |  | Repair scratches, o-ring groove OK | Modify | [14 empty] | F067763-08 | georged | 2010-11-15 |  |  |  | Modify
205 | 123 |  | vrm123 | xxx |  |  |  |  | 123 | carpente | 2010-11-16 |  |  |  | CARPENTE | 2010-11-16 |  | Use As Is | [14 empty] | 123 | carpente | 2010-11-16 |  | carpente | 2010-11-16 | Use As Is
206 | C100-RI-022 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 24 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | ogle | 2010-11-16 |  |  |  | MARHAUSE | 2010-11-18 |  | Use As Is | [14 empty] | C100-RI-022 | ogle | 2010-11-16 | 575916F1-0F56-813D-97BE64DC05A5D626 |  | 2010-11-18 | Use As Is
207 | F063950_02 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R3 | 30 | small scratch on face of small flange. | carpente | 2010-11-17 |  |  |  | KWILSON |  |  | Modify | [14 empty] | F063950_02 | carpente | 2010-11-17 |  |  |  | Modify
208 | C100-RI-023 |  | CRM1207013-0001 | C100 Cavity |  | C100-CAV-INSP | R3 | 25 | See attached dimensional info. | georged | 2010-11-17 |  |  |  | MARHAUSE | 2010-11-18 |  | Use As Is | [14 empty] | C100-RI-023 | georged | 2010-11-17 | 5B7AEAC9-C231-D63A-F3C3576B2A416669 |  | 2010-11-18 | Use As Is
209 | F058824-02 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R1 | 3 | Scratches on small flange. Copper on face on small flange is flaking off. Nicks | georged | 2010-11-17 |  |  |  | KWILSON |  | Repair scratches. Lightly polish flaking copper. | Modify | [14 empty] | F058824-02 | georged | 2010-11-17 |  |  |  | Modify
210 | F067763-16 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R4 | 55 | Nick on knife edge | georged | 2010-11-17 |  |  |  | KWILSON |  | Return to vendor due to nick on knife edge. | Return To Vendor | [14 empty] | F067763-16 | georged | 2010-11-17 |  |  |  | Return To Vendor
211 | F058824_01 |  | CRM-120-7070-0000 | C100 waveguide |  | C100-CAV-INSP-WGD | R1 | 2 | Plating flaking off of small flange. Looks bad. see pics in step 19 in traveler. | georged | 2010-11-17 |  |  |  | KWILSON |  | Lightly polish off flaking plating with scotchbrite. | Modify | [14 empty] | F058824_01 | georged | 2010-11-17 |  |  |  | Modify
212 | C100-RI-025 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 26 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING. | ogle | 2010-11-17 |  |  |  | MARHAUSE | 2010-11-18 |  | Use As Is | [14 empty] | C100-RI-025 | ogle | 2010-11-17 | 5CAF22BA-EABC-6BD4-266050CBF3B0D039 |  | 2010-11-18 | Use As Is
213 | C100-RI-026 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 27 | CONSIDERABLE SCRATCHES AND NICKS ON THE FPC FLANGE FACE. SEE ATTACHMENT FOR DIME | ogle | 2010-11-17 |  |  |  | MARHAUSE | 2010-11-18 | The scratches and nicks on the FPC flange are too significant to accept the cavi | Return To Vendor | [14 empty] | C100-RI-026 | ogle | 2010-11-17 | 5D2F8D03-9B14-926B-EB421F97CB17B621 |  | 2010-11-18 | Return To Vendor
214 | F067764-02 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R4 | 57 | Scratches on small flange. Nicks on knife edge | georged | 2010-11-18 |  |  |  | KWILSON |  | Repair knife edge and scratches. | Modify | [14 empty] | F067764-02 | georged | 2010-11-18 |  |  |  | Modify
215 | F061359-09 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R2 | 9 | Alot of scratches on small flange, will definitely need lapping to repair | georged | 2010-11-18 |  |  |  | KWILSON |  | Repair scratches. | Modify | [14 empty] | F061359-09 | georged | 2010-11-18 |  |  |  | Modify
216 | F067763-05 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R4 | 51 | Scratches on small flange, will need lapping | georged | 2010-11-18 |  |  |  | KWILSON |  | Repair scratches. | Modify | [14 empty] | F067763-05 | georged | 2010-11-18 |  |  |  | Modify
217 | F063950-04 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R3 | 31 | Nick on knife edge, not terrible and is repairable. Scratches on small flange | georged | 2010-11-18 |  |  |  | KWILSON |  | Repair knife edge and scratches. | Modify | [14 empty] | F063950-04 | georged | 2010-11-18 |  |  |  | Modify
218 | F061359-10 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R2 | 10 | Surface finish out of tolerance, measured 18 m inch. Light scratches present on | georged | 2010-11-18 |  |  |  | KWILSON |  | Repair scratches. | Modify | [14 empty] | F061359-10 | georged | 2010-11-18 |  |  |  | Modify
219 | F063951-05 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R3 | 25 | Light scratches on small flange | georged | 2010-11-18 |  |  |  | KWILSON |  | Repair scratches. | Modify | [14 empty] | F063951-05 | georged | 2010-11-18 |  |  |  | Modify
220 | C100-RI-028 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 28 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING | ogle | 2010-11-18 |  |  |  | MARHAUSE | 2010-11-18 |  | Use As Is | [14 empty] | C100-RI-028 | ogle | 2010-11-18 | 61AAB4FD-9170-3CA7-A4058BE324555EEF |  | 2010-11-18 | Use As Is
221 | C100-RI-027 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 29 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING. | ogle | 2010-11-18 |  |  |  | MARHAUSE | 2010-11-19 |  | Use As Is | [14 empty] | C100-RI-027 | ogle | 2010-11-18 | 625B6D68-0D8C-8691-E4E62BFF307BE3CA |  | 2010-11-19 | Use As Is
222 | C100-RI-003 |  | CRM1207013-0001 | C100 Cavity |  | C100-CAV-INSP2 | R3 | 8 |  | carpente | 2010-11-19 |  |  |  | MARHAUSE | 2010-11-19 |  | Use As Is | [14 empty] | C100-RI-003 | carpente | 2010-11-19 | 656DA5B6-B6FA-9C7A-F3ADB48D38F099A6 |  | 2010-11-19 | Use As Is
223 | ISI-0002 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | C100-CAV-INSP-ADPT | R2 | 7 | Drury - RF Test shows hot spot at one end of ceramic (~300 C) which is ~160 C ho |  | 2010-11-19 |  |  | This brazement failed the RF Test. | HKPARK | 2010-11-19 |  | Scrap | [14 empty] | ISI-0002 |  | 2010-11-19 |  |  | 2010-11-19 | Scrap
224 | F061584-01 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R3 | 32 | Scratches on Cu inside waveguide. Scratches on small flange. Nick on knife edg | georged | 2010-11-19 |  |  |  | KWILSON |  | Please repair scratches on small flange as required and nick on knife edge if po | Modify | [14 empty] | F061584-01 | georged | 2010-11-19 |  |  |  | Modify
225 | 22 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | C100-CAV-INSP-ADPT | R2 | 10 | BRAZE RESIDUE IN GAP. SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING. | ogle | 2010-11-19 |  |  |  | HKPARK | 2010-11-22 | Thinner ceramic does not affect RF performance. Continue to the next step | Use As Is | [14 empty] | 22 | ogle | 2010-11-19 | 674EF982-ABA6-26B6-65EEA6054334C5C9 |  | 2010-11-22 | Use As Is
226 | 27 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | C100-CAV-INSP-ADPT | R2 | 13 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING | ogle | 2010-11-19 |  |  |  | HKPARK | 2010-11-22 | Thinner ceramic does not decrease RF performance. Continue the next step. | Use As Is | [14 empty] | 27 | ogle | 2010-11-19 | 675CE5A7-0233-5F3E-38A81321E9E51B37 |  | 2010-11-22 | Use As Is
227 | [5 empty] | C100-CAV-INSP-ADPT | R2 | 15 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING. | ogle | 2010-11-19 |  |  |  | HKPARK | 2010-11-22 | Thinner ceramic does not decrease RF performance. Continue the next step. | Use As Is | [15 empty] | ogle | 2010-11-19 | 6765E27A-0C1D-4EA3-739C0BE228FF6AF0 |  | 2010-11-22 | Use As Is
228 | 39 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | C100-CAV-INSP-ADPT | R2 | 18 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING | ogle | 2010-11-19 |  |  |  | HKPARK | 2010-11-22 | Thinner ceramic does not decrease RF performance. Continue the next step. | Use As Is | [14 empty] | 39 | ogle | 2010-11-19 | 677212F7-D807-FF1F-E055FF225CE94D52 |  | 2010-11-22 | Use As Is
229 | F061359-04 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R2 | 7 | Blisters on copper, scratches on small flange will need lapping, dents inside wa | georged | 2010-11-22 |  |  |  | KWILSON |  | Repair scratches. | Modify | [14 empty] | F061359-04 | georged | 2010-11-22 |  |  |  | Modify
230 | F061359-07 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R2 | 8 | O-ring groove depth out of tolerance 0.001 , Copper flaking, dents and scratches | georged | 2010-11-22 |  |  |  | KWILSON |  |  | Use As Is | [14 empty] | F061359-07 | georged | 2010-11-22 |  |  |  | Use As Is
231 | R100_RI_008 |  |  |  | Cavity was field emission loaded at 12 MV/m during vertical test. | C100-CAV-VTRF | R2 | 11 | Cavity was field emission loaded at 12 MV/m during vertical test. |  | 2010-11-22 |  |  |  | MARHAUSE | 2010-11-22 | cavity requires HPR and needs to be remeasured. No chemistry first. | Remeasure | [14 empty] | R100_RI_008 |  | 2010-11-22 |  |  | 2010-11-22 | Remeasure
232 | C100-RI-011 |  | CRM1207013-0001 | C100 Cavity |  | C100-CAV-INSP2 | R3 | 9 | See attached report | carpente | 2010-11-22 |  |  |  | MARHAUSE |  |  | Use As Is | [14 empty] | C100-RI-011 | carpente | 2010-11-22 | 75421334-DAC8-F944-559BF3D12C9FC625 |  |  | Use As Is
233 | C100-RI-029 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 30 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING | ogle | 2010-11-22 |  |  |  | MARHAUSE | 2010-11-22 |  | Use As Is | [14 empty] | C100-RI-029 | ogle | 2010-11-22 | 7611E22E-FE60-38F2-ADC29E59F58B76A1 |  | 2010-11-22 | Use As Is
234 | C100-RI-030 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 31 | HOM PROBE B: LIGHT STAINS ON FACE;SMALL SPOT STAIN IN SEAL GROOVE. SEE ATTACHME | ogle | 2010-11-22 |  |  |  | MARHAUSE | 2010-11-23 |  | Use As Is | [14 empty] | C100-RI-030 | ogle | 2010-11-22 | 76BD41A6-994C-FCDB-3B829AD55415309E |  | 2010-11-23 | Use As Is
235 | C100-RI-031 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 32 | LIGHT SCRATCHES ON FLANGE FACE. SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING. | ogle | 2010-11-23 |  |  |  | MARHAUSE | 2010-11-23 |  | Use As Is | [14 empty] | C100-RI-031 | ogle | 2010-11-23 | 7B7F6374-91EE-9D93-2C2CAFFBF9BD8F95 |  | 2010-11-23 | Use As Is
236 | C100-RI-032 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 33 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING. | ogle | 2010-11-23 |  |  |  | MARHAUSE | 2010-11-23 |  | Use As Is | [14 empty] | C100-RI-032 | ogle | 2010-11-23 | 7BFEF181-FD1C-CE3B-0F122B77D2BCADE8 |  | 2010-11-23 | Use As Is
237 | C100-RI-033 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 34 | LIGHT SCRATCHES ON FPC FACE. SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING. | ogle | 2010-11-24 |  |  |  | MARHAUSE | 2010-11-25 |  | Use As Is | [14 empty] | C100-RI-033 | ogle | 2010-11-24 | 8095B0D7-AE8B-DE52-F015F87FABB501A2 |  | 2010-11-25 | Use As Is
238 | C100-RI-034 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 35 | LIGHT SCRATCHES ON FPC FACE. SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING. | ogle | 2010-11-24 |  |  |  | MARHAUSE | 2010-11-25 |  | Use As Is | [14 empty] | C100-RI-034 | ogle | 2010-11-24 | 81263A5F-CAC1-E580-309C93FE0DA6A195 |  | 2010-11-25 | Use As Is
239 | C100-RI-035 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 36 | SCRATCHES ON FLANGE FACE; SOME WITHIN SEAL AREA. FIELD PROBE:SURFACE DEFECT ON | ogle | 2010-11-29 |  |  |  | MARHAUSE | 2010-11-30 |  | Use As Is | [14 empty] | C100-RI-035 | ogle | 2010-11-29 | 9A2981D7-9C4E-3DD5-CBA075F688AA1EA9 |  | 2010-11-30 | Use As Is
240 | C100-RI-036 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 37 | LIGHT SCRATCHES ON FACE; NICK SPOT ON FACE.SEE ATTACHMENT FOR DIMENSIONAL NON CO | ogle | 2010-11-29 |  |  |  | MARHAUSE | 2010-11-30 |  | Use As Is | [14 empty] | C100-RI-036 | ogle | 2010-11-29 | 9AC25F2E-B2D1-CF79-687A71CEBC83EA9D |  | 2010-11-30 | Use As Is
241 | C100-RI-012 |  | CRM1207013-0001 | C100 Cavity |  | C100-CAV-INSP2 | R3 | 10 |  | carpente | 2010-11-30 |  |  |  | MARHAUSE | 2010-11-30 | Cavities RI011 and RI-012 have been knowingly built too short due to a fabricati | Use As Is | [14 empty] | C100-RI-012 | carpente | 2010-11-30 | 9DAE3618-F1F2-89A2-C28F56E8DEC63029 |  | 2010-11-30 | Use As Is
242 | C100-RI-016 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 38 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING | ogle | 2010-11-30 |  |  |  | MARHAUSE | 2010-12-01 |  | Use As Is | [14 empty] | C100-RI-016 | ogle | 2010-11-30 | 9F268F83-E57B-F199-C994FB989617994B |  | 2010-12-01 | Use As Is
243 | C100-RI-037 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 39 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING | ogle | 2010-11-30 |  |  |  | MARHAUSE | 2010-12-01 |  | Use As Is | [14 empty] | C100-RI-037 | ogle | 2010-11-30 | A01D9216-DE41-8D20-EDF14B3F7367B2D4 |  | 2010-12-01 | Use As Is
244 | C100-RI-038 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 40 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING. | ogle | 2010-12-01 |  |  |  | MARHAUSE | 2010-12-02 |  | Use As Is | [14 empty] | C100-RI-038 | ogle | 2010-12-01 | A4E97E41-E718-D687-98614F411EB19933 |  | 2010-12-02 | Use As Is
245 | C100-RI-039 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 41 | HEAVY SCRATCH IN SEAL AREA ON FACE. SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMI | ogle | 2010-12-01 |  |  |  | MARHAUSE | 2010-12-02 | JLAB does not accept cavities with heavy scratches on the FPC sealing surfaces f | Return To Vendor | [14 empty] | C100-RI-039 | ogle | 2010-12-01 | A53B618D-0201-0ADB-69014920405620E3 |  | 2010-12-02 | Return To Vendor
246 | C100-RI-040 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 42 | VOID ON FIELD PROBE FACE,OUTSIDE SEAL AREA. SEE ATTACHMENT FOR DIMENSIONAL NON C | ogle | 2010-12-02 |  |  |  | MARHAUSE | 2010-12-07 | use as is but clean void properly | Use As Is | [14 empty] | C100-RI-040 | ogle | 2010-12-02 | A9791168-C106-7120-F27A2C8179AA3B8A |  | 2010-12-07 | Use As Is
247 | C100-RI-013 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP2 | R3 | 11 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING | ogle | 2010-12-02 |  |  |  | MARHAUSE | 2010-12-07 |  | Use As Is | [14 empty] | C100-RI-013 | ogle | 2010-12-02 | A9DD1CB1-F6AB-8639-B7A3E9104E0BD246 |  | 2010-12-07 | Use As Is
248 | C100-RI-041 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 43 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING | ogle | 2010-12-02 |  |  |  | MARHAUSE | 2011-03-29 |  | Use As Is | [14 empty] | C100-RI-041 | ogle | 2010-12-02 | AA510E5E-F7D4-7545-8CEE1A11A1BFF066 |  | 2011-03-29 | Use As Is
249 | F058824-04 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R1 | 5 | Small amount of copper plating needs removal for seal clearance, dark stains nee | carpente | 2010-12-03 |  |  |  | KWILSON |  | Lightly polish to mechanically remove black stains and burr near bolt hole on sm | Modify | [14 empty] | F058824-04 | carpente | 2010-12-03 |  |  |  | Modify
250 | F058824-03 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R1 | 4 | Small amount of copper plating needs removal for seal clearance. | carpente | 2010-12-03 |  |  |  | KWILSON |  | Repair plating. Check with Brian Carpenter for procedure. | Modify | [14 empty] | F058824-03 | carpente | 2010-12-03 |  |  |  | Modify
251 | C100-RI-007 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP2 | R3 | 12 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING | ogle | 2010-12-03 |  |  |  | MARHAUSE | 2010-12-07 |  | Use As Is | [14 empty] | C100-RI-007 | ogle | 2010-12-03 | AF10F868-A6CD-0198-C955424358C3C611 |  | 2010-12-07 | Use As Is
252 | C100-RI-005 |  |  |  |  | C100-CAV-VTRF | R2 | 13 | Shorted field probe cable in dewar prevented accurate Qext measurement for FP |  | 2010-12-06 |  |  |  | MARHAUSE | 2010-12-06 | repair/replacement/re-tightening of cable or cable adapter at feedthrough or por | Modify | [14 empty] | C100-RI-005 |  | 2010-12-06 |  |  | 2010-12-06 | Modify
253 | F058824-01 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R1 | 2 | Copper is peeling up inside of the waveguide.See picture. NCR # 211 also address | carpente | 2010-12-06 |  |  |  | KWILSON |  | Do not throw out! | Scrap | [14 empty] | F058824-01 | carpente | 2010-12-06 | BD8C24C7-047B-386B-9D4F4B3C3BFE4DE0 |  |  | Scrap
254 | C100-RI-042 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 44 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING | ogle | 2010-12-06 |  |  |  | MARHAUSE | 2010-12-07 |  | Use As Is | [14 empty] | C100-RI-042 | ogle | 2010-12-06 | BE492AD3-CB6D-C324-EDD923CB66129822 |  | 2010-12-07 | Use As Is
255 | C100-RI-043 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 45 | HEAVY DAMAGE SPOT AT EDGE OF MOUNTING HOLE HOM A;SEE ATTACHMENT FOR DIMENSIONAL | ogle | 2010-12-06 |  |  |  | MARHAUSE | 2010-12-07 | The integrity of the damaged HOM A flange needs to be checked, particularly the | Modify | [14 empty] | C100-RI-043 | ogle | 2010-12-06 | BF0274F9-C202-604D-B8CF37CBD63FDCDE |  | 2010-12-07 | Modify
256 | C100-RI-044 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 46 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING. | ogle | 2010-12-07 |  |  |  | MARHAUSE | 2010-12-14 |  | Use As Is | [14 empty] | C100-RI-044 | ogle | 2010-12-07 | C347ED4F-FE52-F32A-980937BE29B35824 |  | 2010-12-14 | Use As Is
257 | C100-RI-045 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 47 | STAIN IN SEAL AREA; FIELD PROBE. SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING. | ogle | 2010-12-07 |  |  |  | MARHAUSE | 2010-12-14 |  | Use As Is | [14 empty] | C100-RI-045 | ogle | 2010-12-07 | C4497DB3-B705-1380-6FD4CFDDE8F4D5B8 |  | 2010-12-14 | Use As Is
258 | C100-RI-014 |  |  | CAVITY |  | C100-CAV-INSP2 | R3 | 13 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING | ogle | 2010-12-08 |  |  |  | MARHAUSE | 2010-12-14 |  | Use As Is | [14 empty] | C100-RI-014 | ogle | 2010-12-08 | C8360129-D5BD-DF1C-A8A7198A97DF72A1 |  | 2010-12-14 | Use As Is
259 | C100-RI-046 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 48 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING | ogle | 2010-12-08 |  |  |  | MARHAUSE | 2010-12-14 |  | Use As Is | [14 empty] | C100-RI-046 | ogle | 2010-12-08 | C8C6E92D-F7BD-B9CB-73B90114D9CAFF14 |  | 2010-12-14 | Use As Is
260 | C100-RI-047 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 49 | SCRATCHES AND NICKS ON FPC FLANGE FACE;SEE PHOTOS IN TRAVELER. SEE ATTACHMENT F | ogle | 2010-12-08 |  |  |  | MARHAUSE | 2010-12-14 | Please determine how severe the flange surface is blemished when compared to RI- | Modify | [14 empty] | C100-RI-047 | ogle | 2010-12-08 | C95CBA1C-BD03-1627-38F7729375109E6F |  | 2010-12-14 | Modify
261 | RI-PRO-04 |  | CRM1207014-0020 | HOM Feedthru |  | C100-CAV-INSP-HMFT | R3 | 48 | Failed RF test, see traveler for data. | carpente | 2010-12-09 |  |  |  | HKPARK | 2011-04-01 |  | Scrap | [14 empty] | RI-PRO-04 | carpente | 2010-12-09 |  |  | 2011-04-01 | Scrap
262 | MPF-0006 |  | CRM120-7014-0020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 54 | SCRATCH IN SEAL GROOVE. | ogle | 2010-12-10 |  |  |  | HKPARK | 2010-12-13 | Polish the seal surface and proceed with leak check | Modify | [14 empty] | MPF-0006 | ogle | 2010-12-10 |  |  |  | Use As Is | Polish done and used-10/31/11
263 | MPF-0009 |  | CRM12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 55 | CONTAMINATION ON CERAMIC INSERT INSIDE CONNECTOR; LIGHT SCRATCH ON SEAL GROOVE | ogle | 2010-12-11 |  |  |  | HKPARK | 2010-12-14 | Polish the seal surface and proceed with leak check. Contamination will be verif | Modify | [14 empty] | MPF-0009 | ogle | 2010-12-11 |  |  |  | Modify | Modified and used - HKP 11/2/11
264 | MPF-0010 |  | CRM120-7014-0020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 56 | CONTAMINATION ON CERAMIC INSERT INSIDE CONNECTOR; LIGHT SCRATCH ON SEAL GROOVE | ogle | 2010-12-11 |  |  |  | HKPARK | 2010-12-14 | Polish the seal surface and proceed with leak check. Contamination will be verif | Modify | [14 empty] | MPF-0010 | ogle | 2010-12-11 |  |  |  | Modify | Modified and used - HKP 11/2/11
265 | MPF-0012 |  | CRM120-7014-0020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 57 | CONTAMINATION ON CERAMIC INSERT INSIDE CONNECTOR; LIGHT SCRATCH ON SEAL GROOVE | ogle | 2010-12-11 |  |  |  | HKPARK | 2010-12-14 | Polish the seal surface and proceed with leak check. Contamination will be verif | Modify | [14 empty] | MPF-0012 | ogle | 2010-12-11 |  |  |  | Modify | Modified and used - HKP 11/2/11
266 | MPF-0013 |  | CRM120-7014-0020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 58 | CONTAMINATION SPOTS IN SEAL GROOVE; LIGHT SCRATCH ON SEAL GROOVE. SEE ATTACHM | ogle | 2010-12-11 |  |  |  | HKPARK | 2010-12-14 | Polish the seal surface and proceed with leak check. Contamination will be verif | Modify | [14 empty] | MPF-0013 | ogle | 2010-12-11 | D3D90FA3-E7D1-8836-EC5DE03734E4B852 |  |  | Modify | Modified and used - HKP 11/2/11
267 | MPF-0014 |  | CRM120-7014-0020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 59 | CONTAMINATION ON CERAMIC INSERT INSIDE CONNECTOR; LIGHT SCRATCH ON SEAL GROOVE | ogle | 2010-12-11 |  |  |  | HKPARK | 2010-12-14 | Polish the seal surface and proceed with leak check. Contamination will be verif | Modify | [14 empty] | MPF-0014 | ogle | 2010-12-11 | D3DB090A-C187-12F7-AC914BFAC266AB17 |  |  | Modify | Modified and used - HKP 11/2/11
268 | MPF-0015 |  | CRM120-7014-0020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 60 | CONTAMINATION ON CERAMIC INSERT INSIDE CONNECTOR; LIGHT SCRATCH ON SEAL GROOVE | ogle | 2010-12-11 |  |  |  | HKPARK | 2010-12-14 | Polish the seal surface and proceed with leak check. Contamination will be verif | Modify | [14 empty] | MPF-0015 | ogle | 2010-12-11 | D3DCE9BD-9BDD-D6B6-608FFF3D1053189E |  |  | Modify | Modified and used - HKP 11/2/11
269 | MPF-0016 |  | CRM120-7014-0020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 61 | LIGHT CONTAMINATION ON CERAMIC INSERT INSIDE CONNECTOR; LIGHT SCRATCH ON SEAL | ogle | 2010-12-11 |  |  |  | HKPARK | 2010-12-14 | Polish the seal surface and proceed with leak check. Contamination will be verif | Modify | [14 empty] | MPF-0016 | ogle | 2010-12-11 |  |  |  | Modify | Modified and used - HKP 11/2/11
270 | MPF-0018 |  | CRM120-7014-0020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 62 | LIGHT CONTAMINATION ON CERAMIC INSERT INSIDE CONNECTOR; LIGHT SCRATCH ON SEAL | ogle | 2010-12-11 |  |  |  | HKPARK | 2010-12-14 | Polish the seal surface and proceed with leak check. Contamination will be verif | Modify | [14 empty] | MPF-0018 | ogle | 2010-12-11 | D3E04AA8-E57B-8AE1-6E9535422293A7FB |  |  | Modify | Modified and used - HKP 11/2/11
271 | MPF-0020 |  | CRM120-7014-0020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 63 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING | ogle | 2010-12-11 |  |  |  | HKPARK | 2010-12-13 | The out tol value is acceptable | Use As Is | [14 empty] | MPF-0020 | ogle | 2010-12-11 | D3E1D66E-92E7-160D-8984EAB9E68949A5 |  | 2010-12-13 | Use As Is
272 | MPF-0023 |  | CRM120-7014-0020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 64 | LIGHT CONTAMINATION ON CERAMIC INSERT INSIDE CONNECTOR; | ogle | 2010-12-11 |  |  |  | HKPARK | 2010-12-14 | Loose contamination can be cleaned if it's material contamination it will be ver | Use As Is | [14 empty] | MPF-0023 | ogle | 2010-12-11 |  |  | 2010-12-14 | Use As Is
273 | MPF-0024 |  | CRM120-7014-0020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 65 | LIGHT SCRATCH ON SEAL GROOVE. SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING | ogle | 2010-12-11 |  |  |  | HKPARK | 2010-12-14 | Polish the seal surface and proceed with leak check. Out tol level is acceptable | Modify | [14 empty] | MPF-0024 | ogle | 2010-12-11 | D3E51919-0D30-E58D-6BF97F324779AE7F |  |  | Modify | Modified and used - HKP 11/2/11
274 | MPF-0025 |  | CRM120-7014-0020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 66 | LIGHT SCRATCH ON SEAL GROOVE. | ogle | 2010-12-11 |  |  |  | HKPARK | 2010-12-13 | Polish the seal surface and proceed with leak check. | Modify | [14 empty] | MPF-0025 | ogle | 2010-12-11 |  |  |  | Modify | Modified and used - HKP 11/2/11
275 | MPF-0027 |  | CRM120-7014-0020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 67 | CONTAMINATION ON CERAMIC INSERT INSIDE CONNECTOR; LIGHT SCRATCH ON SEAL GROOVE | ogle | 2010-12-11 |  |  |  | HKPARK | 2010-12-14 | Polish the seal surface and proceed with leak check. Contamination will be verif | Modify | [14 empty] | MPF-0027 | ogle | 2010-12-11 | D3E850CD-03FA-223F-BF73D2F864F6FCE8 |  |  | Modify | Modified and used - HKP 11/2/11
276 | MPF-0028 |  | CRM120-7014-0020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 68 | CONTAMINATION ON CERAMIC INSERT INSIDE CONNECTOR; SEE ATTACHMENT FOR DIMENSIONA | ogle | 2010-12-11 |  |  |  | HKPARK | 2010-12-14 | If it's loose contamination clean. if it looks material Contamination it will be | Use As Is | [14 empty] | MPF-0028 | ogle | 2010-12-11 | D3E9B003-B2BD-46C0-2DD35E350D9E6D87 |  | 2010-12-14 | Use As Is
277 | MPF-0031 |  | CRM120-7014-0020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 70 | LIGHT SCRATCH ON SEAL GROOVE. | ogle | 2010-12-11 |  |  |  | HKPARK | 2010-12-13 | Polish seal surface and proceed with leak check | Modify | [14 empty] | MPF-0031 | ogle | 2010-12-11 |  |  |  | Modify | Modified and used - HKP 11/2/11
278 | C100_RI_007 |  |  |  | Vertical RF test, step 13. | C100-CAV-VTRF | R2 | 19 | Cavity lock frequency 1496.3975 below specification of 1496.4 MHz @29 torr. |  | 2010-12-12 |  |  |  | MARHAUSE | 2010-12-14 |  | Use As Is | [14 empty] | C100_RI_007 |  | 2010-12-12 |  |  | 2010-12-14 | Use As Is
279 | C100_RI_012 |  |  |  | VTA RF Test Step 13 | C100-CAV-VTRF | R2 | 20 | Cavity lock frequency at 29 torr was 1496.2506 which is below specification of 1 |  | 2010-12-12 |  |  |  | MARHAUSE |  | It was decided to use the cavity for the first C100 string. The lower frequency | Use As Is | [14 empty] | C100_RI_012 |  | 2010-12-12 |  |  |  | Use As Is
280 | C100_RI_010 |  |  |  | page 2, step 13 | C100-CAV-VTRF | R2 | 21 | cavity lock frequency 1496.3718 is below specified limit of 1496.4MHz |  | 2010-12-14 |  |  |  | MARHAUSE | 2010-12-14 |  | Use As Is | [14 empty] | C100_RI_010 |  | 2010-12-14 |  |  | 2010-12-14 | Use As Is
281 | C100-RI-048 |  | CRM12070130001 | C100 cavity |  | C100-CAV-INSP | R3 | 50 | Please see attached PDF for NCR data | georged | 2010-12-15 |  |  |  | MARHAUSE | 2010-12-15 | Please determine how severe the flange surface is blemished when compared to RI- | Modify | [14 empty] | C100-RI-048 | georged | 2010-12-15 | EA2C67BD-A933-B422-279C8085CF0A66E2 |  | 2010-12-15 | Modify
282 | C100-RI-013 |  | CRM12070130001 | C100 cavity |  | C100-CAV-INSP2 | R3 | 15 | FPC flange location dimension out of tolerance 0.004 , measured 4.511 | georged | 2010-12-17 |  |  |  | MARHAUSE | 2010-12-17 |  | Use As Is | [14 empty] | C100-RI-013 | georged | 2010-12-17 |  |  | 2010-12-17 | Use As Is
283 | C100-RI-002 |  | CRM12070130001 | C100 Cavity |  | C100-CAV-INSP2 | R3 | 16 | Please see attached PDF | georged | 2010-12-20 |  |  |  | MARHAUSE | 2010-12-20 |  | Use As Is | [14 empty] | C100-RI-002 | georged | 2010-12-20 | 03F6C870-9757-E9AA-1853326D8CA642B1 |  | 2010-12-20 | Use As Is
284 | C100-RI-049 |  | CRM12070130001 | C100 Cavity |  | C100-CAV-INSP | R3 | 51 | Please see attached PDF | georged | 2010-12-20 |  |  |  | MARHAUSE | 2010-12-20 |  | Use As Is | [14 empty] | C100-RI-049 | georged | 2010-12-20 | 050E885E-E445-1EA8-6C0C681EFA354AF1 |  | 2010-12-20 | Use As Is
285 | C100-RI-080 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 52 | LIGHT SCRATCHES ON FPC FLANGE FACE. SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMIN | OGLE | 2010-12-22 | 59BC7D4F-9C93-8B32-32367425FE159B88 |  |  | MARHAUSE | 2011-01-04 |  | Use As Is | [14 empty] | C100-RI-080 | ogle | 2010-12-22 | 0F794E33-FCAF-0B90-4654657CEBEBBD39 |  |  | Use As Is
286 | C100-RI-081 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 53 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING. | ogle | 2011-01-03 |  |  |  | MARHAUSE | 2011-01-04 |  | Use As Is | [14 empty] | C100-RI-081 | ogle | 2011-01-03 | 4E49B722-FCD4-4DFF-299CE59FB097135F |  | 2011-01-04 | Use As Is
287 | C100-RI-082 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 54 | SCRATCH ON FPC FACE IN SEAL AREA. SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING | ogle | 2011-01-03 |  |  |  | MARHAUSE | 2011-01-04 |  | Use As Is | [14 empty] | C100-RI-082 | ogle | 2011-01-03 | 4F2CA42A-98D2-3FD1-1E4AFD13432641DA |  | 2011-01-04 | Use As Is | Use as is with FPC polishing required (as usual) but particularly beware of the | 516DEC32-E360-E36C-3E3D39EE9140AA85
288 | C100-RI-083 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 55 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING. | ogle | 2011-01-05 |  |  |  | MARHAUSE | 2011-03-29 |  | Use As Is | [14 empty] | C100-RI-083 | ogle | 2011-01-05 | 5845B5A7-B206-18C4-F6C346733AEE61D1 |  | 2011-03-29 | Use As Is
289 | FLG0037-031 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R1 | 21 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING | ogle | 2011-01-05 |  |  |  | HKPARK | 2011-01-06 | Out tol range for the thickness and bolt hole location is acceptable(usuable). | Use As Is | [14 empty] | FLG0037-031 | ogle | 2011-01-05 | 5936FE27-9FDA-AE06-3A54CF6C2DC47045 |  | 2011-01-06 | Use As Is
290 | FLG0037-032 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R1 | 22 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING | ogle | 2011-01-05 |  |  |  | HKPARK | 2011-01-06 | use as is if the adapter fits. The flange thickness is acceptable. | Use As Is | [14 empty] | FLG0037-032 | ogle | 2011-01-05 | 59439CCF-DCE4-C063-BF0BA1CF5F0EFCDB |  | 2011-01-06 | Use As Is
291 | FLG0037-033 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R1 | 23 | SEE ATTACHMENT FOR DIMENSIONAL NON- CONFORMING | ogle | 2011-01-05 |  |  |  | HKPARK | 2011-01-06 | Out tol range for the thickness and bolt hole location is acceptable(usuable). | Use As Is | [14 empty] | FLG0037-033 | ogle | 2011-01-05 | 594E9C7D-A57E-685A-31FB837410AD9BDA |  | 2011-01-06 | Use As Is
292 | FLG0037-034 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R1 | 24 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING. | ogle | 2011-01-05 |  |  |  | HKPARK | 2011-01-06 | Use as is if the adapter fits. Out tol range for the thickness and bolt hole l | Use As Is | [14 empty] | FLG0037-034 | ogle | 2011-01-05 | 59651B7F-F6E2-D2E0-09A4FF3E88FA7978 |  | 2011-01-06 | Use As Is
293 | FLG0037-035 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R1 | 25 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING | ogle | 2011-01-05 |  |  |  | HKPARK | 2011-01-06 | Out tol range for the thickness is acceptable. | Use As Is | [14 empty] | FLG0037-035 | ogle | 2011-01-05 | 596CECA2-01E6-FEA7-0512B908A001FCE6 |  | 2011-01-06 | Use As Is
294 | FLG0037-030 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R1 | 20 | DAMAGED KNIFE EDGE. SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING | ogle | 2011-01-05 |  |  |  | HKPARK | 2011-01-06 |  | Scrap | [14 empty] | FLG0037-030 | ogle | 2011-01-05 | 5977C8E2-A571-E92F-E8026C3D186A359A |  | 2011-01-06 | Scrap
295 | FLG0037-036 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R1 | 26 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING | ogle | 2011-01-06 |  |  |  | HKPARK | 2011-04-01 |  | Use As Is | [14 empty] | FLG0037-036 | ogle | 2011-01-06 | 5D78D8F1-D3C1-86F0-D372D04182165782 |  | 2011-04-01 | Use As Is
296 | C100-RI-084 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 56 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING | ogle | 2011-01-06 |  |  |  | MARHAUSE | 2011-01-07 |  | Use As Is | [14 empty] | C100-RI-084 | ogle | 2011-01-06 | 5E22979C-AA4E-E495-D1F943E3ADDBDAF5 |  | 2011-01-07 | Use As Is
297 | C100-RI-017 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP2 | R3 | 17 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING | ogle | 2011-01-06 |  |  |  | MARHAUSE | 2011-01-07 |  | Use As Is | [14 empty] | C100-RI-017 | ogle | 2011-01-06 | 5E9E800D-0E6E-8952-0BC430CDD7ED388A |  | 2011-01-07 | Use As Is
298 | F067764-17 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R5 | 90 | Copper flaking inside waveguide. Scratches on small flange.Scratches inside o-ri | georged | 2011-01-10 |  |  |  | KWILSON |  | Repair scratches. Flatness ok. | Modify | [14 empty] | F067764-17 | georged | 2011-01-10 |  |  |  | Modify
299 | F058824-03 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R1 | 4 | length out of tolerance 0.016 , o-ring groove width out of tolerance 0.003 | carpente | 2011-01-10 |  |  |  | KWILSON |  |  | Use As Is | [14 empty] | F058824-03 | carpente | 2011-01-10 |  |  |  | Use As Is
300 | C100-RI-050 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 57 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING | ogle | 2011-01-10 |  |  |  | MARHAUSE | 2011-01-11 |  | Use As Is | [14 empty] | C100-RI-050 | ogle | 2011-01-10 | 73481690-A7C2-8786-8DEA904A4BBD163E |  | 2011-01-11 | Use As Is
301 | F063951-07 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R3 | 24 | Dent found in bellows. | carpente | 2011-01-12 |  |  |  | CARPENTE | 2011-01-12 | Save as emergency backup if we run out of other waveguides. | Scrap | [14 empty] | F063951-07 | carpente | 2011-01-12 |  |  | 2011-01-12 | Scrap
302 | C100-RI-001 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP2 | R3 | 18 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING | ogle | 2011-01-12 |  |  |  | MARHAUSE | 2011-01-13 |  | Use As Is | [14 empty] | C100-RI-001 | ogle | 2011-01-12 | 7D8D2F4F-F35C-D38C-A69FB3CBD1FAC931 |  | 2011-01-13 | Use As Is
303 | C100-RI-018 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP2 | R3 | 19 | SEE ATTACHMENT FOR DIMENSIONAL NONCONFORMING | ogle | 2011-01-14 |  |  |  | MARHAUSE | 2011-01-18 |  | Use As Is | [14 empty] | C100-RI-018 | ogle | 2011-01-14 | 8702224D-F0A5-FF62-60F29CC9B05E0114 |  | 2011-01-18 | Use As Is
304 | C100-RI-051 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 58 | STAINS IN SEAL GROOVE OF FIELD PROBE AND HOM FLANGE A.(SEE PHOTOS IN TRAVELER) | ogle | 2011-01-18 |  |  |  | MARHAUSE | 2011-01-18 |  | Use As Is | [14 empty] | C100-RI-051 | ogle | 2011-01-18 | 9BFFFA2D-94B4-732D-16D3DABBCACE1F25 |  | 2011-01-18 | Use As Is | Sealing grooves in field probe, HOM A and HOM B flange show stains and should be
305 | C100_RI_008 |  |  |  |  | C100-CAV-VTRF | R2 | 26 | VTA RF test: cavity was field emission loaded at 7 MV/m. RF processing ineffect |  | 2011-01-19 |  |  |  | MARHAUSE | 2011-01-19 |  | Modify | [14 empty] | C100_RI_008 |  | 2011-01-19 |  |  | 2011-01-19 | Modify | A light chemistry (30um EP, degrease, HPR) is recommended, followed by retuning
306 | C100-RI-008 |  |  |  | Attached is the action item list for the path forward for C100-8 | C100-CAV-VTRF | R2 | 26 |  |  | 2011-01-20 |  |  |  | MARHAUSE | 2011-01-20 | Modify as described in the action item list. Recovery plan yet unknown. | Modify | [14 empty] | C100-RI-008 |  | 2011-01-20 | A3FBA127-DC23-B243-9BB9DD54B953F6BC |  | 2011-01-20 | Modify
307 | C100-RI-052 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 59 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING | ogle | 2011-01-20 |  |  |  | MARHAUSE | 2011-01-20 |  | Use As Is | [14 empty] | C100-RI-052 | ogle | 2011-01-20 | A5D60365-A134-CB0F-C77EAB424A202C40 |  | 2011-01-20 | Use As Is
308 | C100-RI-053 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 60 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING | ogle | 2011-01-20 |  |  |  | MARHAUSE | 2011-01-21 |  | Use As Is | [14 empty] | C100-RI-053 | ogle | 2011-01-20 | A6B26349-A831-5911-C580A32062CCC16E |  | 2011-01-21 | Use As Is
309 | C100_RI_014 |  |  |  |  | C100-CAV-VTRF | R3 | 30 | C100_RI_014 lock frequency was 1496.3347 MHz which is below the specified minimu |  | 2011-01-24 |  |  |  | ABURRILL | [17 empty] | C100_RI_014 |  | 2011-01-24 |  |  |  | Use As Is
310 | C100_RI_014 |  |  |  |  | C100-CAV-VTRF | R3 | 30 | C100_RI_014 lock frequency was 1496.3347 MHz which is below the specified minimu |  | 2011-01-24 |  |  |  | ABURRILL | 2011-11-17 |  | Use As Is | [14 empty] | C100_RI_014 |  | 2011-01-24 |  |  |  | Use As Is
311 | C100_RI_014 |  |  |  |  | C100-CAV-VTRF | R3 | 30 | C100-RI-014 lock frequency was 1496.3347 MHz which is below the specified minimu |  | 2011-01-24 |  |  |  | ABURRILL | [17 empty] | C100_RI_014 |  | 2011-01-24 |  |  |  | Use As Is
312 | C100_RI_014 |  |  |  |  | C100-CAV-VTRF | R3 | 30 | lock frequency out of spec low |  | 2011-01-24 |  |  |  | ABURRILL | [17 empty] | C100_RI_014 |  | 2011-01-24 |  |  |  | Use As Is
315 | ISI-0004 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | C100-CAV-INSP-ADPT | R2 | 19 | GAP BETWEEN COPPER EYELET AND CERAMIC. SEE ATTACHMENT FOR DIMENSIONAL NON CONFO | ogle | 2011-01-25 |  |  |  | HKPARK | 2011-04-01 |  | Scrap | [14 empty] | ISI-0004 | ogle | 2011-01-25 | BF4F58F5-90F7-A29E-2928A9AAE0F0F032 |  | 2011-04-01 | Scrap
316 | C100-RI-016 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP2 | R3 | 20 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING. | ogle | 2011-01-28 |  |  |  | MARHAUSE | 2011-01-28 |  | Use As Is | [14 empty] | C100-RI-016 | ogle | 2011-01-28 | CD6FF464-9F88-BF34-382CEC3088E04128 |  | 2011-01-28 | Use As Is
317 | C100-RI-013 |  |  |  | VTA High Power RF Test of cavity with Helium vessel. | C100-CAV-VTRF | R3 | 31 | cavity failed high power RF test in VTA due to field emmission loading at 15MV/m |  | 2011-01-31 |  |  |  | MARHAUSE | 2011-01-31 | High Pressure rinse cavity to recover from FE loading | Modify | [14 empty] | C100-RI-013 |  | 2011-01-31 |  | kdavis | 2011-01-31 | Modify
318 | C100-RI-019 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP2 | R3 | 21 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING. | ogle | 2011-02-03 |  |  |  | MARHAUSE | 2011-02-03 |  | Use As Is | [14 empty] | C100-RI-019 | ogle | 2011-02-03 | ECCE1F3D-DA1A-24C6-DF741E9D4A982D0E |  | 2011-02-03 | Use As Is
319 | C100_RI_002 |  |  |  |  | C100-CAV-VTRF | R3 | 36 | Cavity lock Freq at 1496.3874MHz is below specification value 1496.4MHz. |  | 2011-02-03 |  |  |  | MARHAUSE | 2011-02-08 |  | Use As Is | [14 empty] | C100_RI_002 |  | 2011-02-03 |  |  | 2011-02-08 | Use As Is
320 | C100-RI-020 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP2 | R3 | 22 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING | ogle | 2011-02-07 |  |  |  | MARHAUSE | 2011-02-08 |  | Use As Is | [14 empty] | C100-RI-020 | ogle | 2011-02-07 | 00A25D6C-EBDA-9A30-A89CED88255B7960 |  | 2011-02-08 | Use As Is
321 | C100_RI_008 |  |  |  |  | C100-CAV-VTRF | R3 | 37 | Cavity is field emission loaded at 14.4 MV/m |  | 2011-02-05 |  |  |  | MARHAUSE | 2011-02-08 | HPR could not fully recover cavity. FE still persistent. Next step is internal i | Modify | [14 empty] | C100_RI_008 |  | 2011-02-05 |  |  | 2011-02-08 | Modify
322 | F067764-15 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R5 | 95 | Surface finish out of tolerance, measured 18 m inches. Flatness out of toleranc | georged | 2011-02-07 |  |  |  | KWILSON |  |  | Use As Is | [14 empty] | F067764-15 | georged | 2011-02-07 |  |  |  | Use As Is
323 | F067764-05 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R5 | 97 | Flatness out of tolerance, measured 0.005" . Scratches on small flange need lap | georged | 2011-02-08 |  |  |  | KWILSON |  | Repair scratches. Try to correct flatness at the same time. | Modify | [14 empty] | F067764-05 | georged | 2011-02-08 |  |  |  | Modify
324 | F067764-10 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R5 | 91 | Scratches on small flange need repair. Flatness out of tolerance, meas. 0.005" | georged | 2011-02-08 |  |  |  | KWILSON |  | Repair scratches. Try to correct flatness issue at the same time. | Modify | [14 empty] | F067764-10 | georged | 2011-02-08 |  |  |  | Modify
325 | F063952-02 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-RWRK-WGD | R1 | 2 | Damage in o-ring groove. Cu missing on small flange. Pictures of damage are uplo | georged | 2011-02-09 |  |  |  | KWILSON |  |  | Use As Is | [14 empty] | F063952-02 | georged | 2011-02-09 | 0AEAAAFB-FCF8-86B1-2117F65C73DF361D |  |  | Use As Is
326 | C100-RI-021 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP2 | R3 | 23 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING. | ogle | 2011-02-10 |  |  |  | MARHAUSE | 2011-02-10 |  | Use As Is | [14 empty] | C100-RI-021 | ogle | 2011-02-10 | 100168D0-BEFA-5871-8652CAC550473D60 |  | 2011-02-10 | Use As Is
327 | [5 empty] | C100-CAV-VTRF | R3 | 39 | cavity vacuum below lambda (36Torr)is 5e-4mBarr |  | 2011-02-10 |  |  |  | MARHAUSE | 2011-02-11 | High vacuum pressure below lambda point indicates superfluid leak at sealing sur | Modify | [16 empty] | 2011-02-10 |  |  | 2011-02-11 | Modify |  | 14EA4230-BBC7-ECE5-836D41DA4061DC6C
328 | MPF-0022 |  | CRM12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 71 | BRAZE MATERIAL MISSING AROUND THE SAPPHIRE INSERT. | ogle | 2011-02-11 |  |  |  | HKPARK | 2011-02-14 | The braze bond between the sapphire insert and the sleeve is more likely robust | Use As Is | [14 empty] | MPF-0022 | ogle | 2011-02-11 | 162A9659-9D5E-DAB1-9C006FB928B47353 |  | 2011-02-14 | Use As Is
329 | MPF-0026 |  | CRM12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 72 | LIGHT CONTAMINATION ON SAPPHIRE SURFACE CONNECTOR SIDE. SAPPHIRE CRACKED AT EDG | ogle | 2011-02-11 |  |  |  | HKPARK | 2011-02-14 | Do not use. Put aside this part. | Scrap | [14 empty] | MPF-0026 | ogle | 2011-02-11 |  |  | 2011-02-14 | Scrap
330 | MPF-0032 |  | CRM12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 73 | LIGHT CONTAMINATION ON SAPPHIRE SURFACE CONNECTOR END. | ogle | 2011-02-11 |  |  |  | HKPARK | 2011-02-14 | Harmless contamination was seen before. Proceed to the next step. If the contami | Use As Is | [14 empty] | MPF-0032 | ogle | 2011-02-11 |  |  | 2011-02-14 | Use As Is
331 | MPF-0033 |  | CRM12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 74 | LIGHT CONTAMINATION ON SAPPHIRE SURFACE CONNECTOR END. | ogle | 2011-02-11 |  |  |  | HKPARK | 2011-02-14 | Harmless contamination was seen before. Proceed to the next step. If the contami | Use As Is | [14 empty] | MPF-0033 | ogle | 2011-02-11 |  |  | 2011-02-14 | Use As Is
332 | MPF-0034 |  | CRM12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 75 | LIGHT CONTAMINATION ON SAPPHIRE SURFACE CONNECTOR END. | ogle | 2011-02-11 |  |  |  | HKPARK | 2011-02-14 | Harmless contamination was seen before. Proceed to the next step. If the contami | Use As Is | [14 empty] | MPF-0034 | ogle | 2011-02-11 |  |  | 2011-02-14 | Use As Is
333 | MPF-0036 |  | CRM12070140020 | HOM FEED THRU | The out tolerance of the tip length is acceptable. The scratch will be tested at | C100-CAV-INSP-HMFT | R3 | 77 | LIGHT SCRATCH IN SEAL GROOVE. SEE TRAVELER FOR PHOTO. SEE ATTACHMENT FOR DIMENS | ogle | 2011-02-11 |  |  |  | HKPARK | 2011-02-14 | Out tolerance for the tip length is acceptable. The scratch will be tested at th | Use As Is | [14 empty] | MPF-0036 | ogle | 2011-02-11 | 1642E8F6-07FF-399D-C30D0E8A72D6AA0F |  | 2011-02-14 | Use As Is
334 | MPF-0037 |  | CRM12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 78 | LIGHT SCRATCH IN SEAL GROOVE.SEE TRAVELER FOR PHOTO. | ogle | 2011-02-11 |  |  |  | HKPARK | 2011-02-14 | Proceed to the leak check. | Use As Is | [14 empty] | MPF-0037 | ogle | 2011-02-11 |  |  | 2011-02-14 | Use As Is
335 | MPF-0039 |  | CRM12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 79 | LIGHT SCRATCH IN SEAL GROOVE. SEE TRAVELER FOR PHOTO. | ogle | 2011-02-11 |  |  |  | HKPARK | 2011-02-14 | Proceed to the leak check. | Use As Is | [14 empty] | MPF-0039 | ogle | 2011-02-11 |  |  | 2011-02-14 | Use As Is
336 | MPF-0040 |  | CRM12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 80 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING. | ogle | 2011-02-11 |  |  |  | HKPARK | 2011-02-14 | The out tolerance of the tip length is acceptable. | Use As Is | [14 empty] | MPF-0040 | ogle | 2011-02-11 | 1659E0B1-040C-41E6-00E758F7358B129A |  | 2011-02-14 | Use As Is
337 | F064155-04 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R3 | 42 | Scratches on small flange | georged | 2011-02-14 |  |  |  | KWILSON |  | Please repair. | Modify | [14 empty] | F064155-04 | georged | 2011-02-14 |  |  |  | Modify
338 | MPF-0042 |  | CRM120-7014-0020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 82 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING | ogle | 2011-02-15 |  |  |  | HKPARK | 2011-04-01 | The deviation amount OK. | Use As Is | [14 empty] | MPF-0042 | ogle | 2011-02-15 | 2B9F88F7-E858-4FBA-DA2974D0E9C745A7 |  | 2011-04-01 | Use As Is
339 | MPF-0043 |  | CRM120-7014-0020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 83 | CENTER CONNECTOR PIN VERY LOOSE BUT ATTACHED.CENTER CONNECTOR PIN DEPTH:REQUIRED | ogle | 2011-02-15 |  |  |  | HKPARK | 2011-04-01 |  | Scrap | [14 empty] | MPF-0043 | ogle | 2011-02-15 | 2BA38292-EF14-D9C1-26CAF0052C3F2045 |  | 2011-04-01 | Scrap
340 | MPF-0046 |  | CRM120-7014-0020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 86 | VERY LIGHT SCRATCH IN SEAL GROOVE. BRAZE MATERIAL ON GROOVE BOSS OUTSIDE OF SEAL | ogle | 2011-02-15 |  |  |  | HKPARK | 2011-05-13 |  | Use As Is | [14 empty] | MPF-0046 | ogle | 2011-02-15 | 2BA8CDC1-F6CB-E83B-B97F99AFA4E3D3BC |  | 2011-05-13 | Use As Is
341 | MPF-0054 |  | CRM120-7014-0020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 90 | BRAZE MATERIAL ON SEAL BOSS. LIGHT SCRATCH IN SEAL GROOVE.SEE ATTACHMENT IN TRAV | ogle | 2011-02-15 |  |  |  | HKPARK | 2011-05-13 | Used as is. | Use As Is | [14 empty] | MPF-0054 | ogle | 2011-02-15 |  |  | 2011-05-13 | Use As Is
342 | MPF-0047 |  | CRM120-7014-0020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 91 | BRAZE MATERIAL IN SEAL GROOVE.SEE PHOTO IN TRAVELER. | ogle | 2011-02-15 |  |  |  | HKPARK | 2011-05-13 | used as is. | Use As Is | [14 empty] | MPF-0047 | ogle | 2011-02-15 |  |  | 2011-05-13 | Use As Is
343 | MPF-0057 |  | CRM120-7014-0020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 92 | VERY LIGHT SCRATCH IN SEAL GROOVE. SEE PHOTO IN TRAVELER. | ogle | 2011-02-15 |  |  |  | HKPARK | [17 empty] | MPF-0057 | ogle | 2011-02-15 |  |  |  | Scrap
344 | MPF-0062 |  | CRM120-7014-0020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 94 | BRAZE MATERIAL ON SEAL BOSS. LIGHT SCRATCH IN SEAL GROOVE. SEE PHOTO IN TRAVELE | ogle | 2011-02-15 |  |  |  | HKPARK |  | Cleaned,polished and used | [15 empty] | MPF-0062 | ogle | 2011-02-15 |  |  |  | Use As Is
345 | F063951-02 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-RWRK-WGD | R2 | 14 | Flatness measured .005" | georged | 2011-02-16 |  |  |  | KWILSON |  |  | Use As Is | [14 empty] | F063951-02 | georged | 2011-02-16 |  |  |  | Use As Is
346 | [5 empty] | C100-CAV-VTRF | R3 | 40 | Cavity lock frequency was 1496.38 MHz, below specification of 1496.400 MHz. | forehand | 2011-02-16 |  |  |  | MARHAUSE | 2011-02-16 |  | Use As Is | [15 empty] | forehand | 2011-02-16 |  |  | 2011-02-16 | Use As Is
347 | F063950-01 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-RWRK-WGD | R2 | 13 | Flatness out of tolerance. Measured .006" . Fingerprints inside waveguide, milk | georged | 2011-02-16 |  |  |  | KWILSON |  | Please set aside. Flatness is questionable. Will see if this one is needed at | Scrap | [14 empty] | F063950-01 | georged | 2011-02-16 | 33CB0FDF-033E-9CCD-75CB670506DC6BB3 |  |  | Scrap
348 | C100-RI-025 |  | CRM-120-7013-0001 | C100 Cavity |  | C100-CAV-INSP2 | R3 | 26 | Please see attached PDF | georged | 2011-02-24 |  |  |  | MARHAUSE | 2011-02-28 |  | Use As Is | [14 empty] | C100-RI-025 | georged | 2011-02-24 | 58864CC2-C523-FA7A-DD7A54FB4053F7D0 |  | 2011-02-28 | Use As Is
349 |  |  |  |  | C100_RI_017, RF testing on 24-Feb-2011 | C100-CAV-VTRF | R3 | 42 | Cavity lock frequency of 1496.3790 MHz below specification of 1496.400 MHz |  | 2011-02-25 |  |  |  | MARHAUSE | 2011-02-28 |  | Use As Is | [16 empty] | 2011-02-25 |  |  | 2011-02-28 | Use As Is
350 | C100_RI_019 |  |  |  |  | C100-CAV-VTRF | R3 | 41 | Cavity lock frequency was 1496.2914; below specification of 1496.4 MHz. |  | 2011-02-28 |  |  |  | MARHAUSE |  | Lock frequency just below lower limit. We had already one cavity before at lower | Use As Is | [14 empty] | C100_RI_019 |  | 2011-02-28 |  |  |  | Use As Is
351 | [5 empty] | C100-CAV-VTRF | R3 | 44 | Cavity lock frequency 1496.2905 MHz; below specification of 1496.400 MHz |  | 2011-03-01 |  |  |  | MARHAUSE | 2011-03-04 | The lower threshold should have been 1496.3 MHz changed from 1496.4 MHz. Theref | Use As Is | [16 empty] | 2011-03-01 |  |  | 2011-03-04 | Use As Is
352 | C100_RI_021 |  |  |  |  | C100-CAV-VTRF | R3 | 44 | Cavity was field emission loaded at 7MV/m, radiation of 10mrem/h; did not achiev |  | 2011-03-01 |  |  |  | MARHAUSE | 2011-03-04 | I don't see the serial number for this cavity. Which one is it? Based on RI-008 | Modify | [14 empty] | C100_RI_021 |  | 2011-03-01 |  |  | 2011-03-04 | Modify
353 | C100-RI-027 |  | CRM1207013-0001 | C100 Cavity |  | C100-CAV-INSP2 | R3 | 27 | Please see PDF | carpente | 2011-03-02 |  |  |  | MARHAUSE | 2011-03-04 |  | Use As Is | [14 empty] | C100-RI-027 | carpente | 2011-03-02 | 777ADC82-A49F-F27E-09CF1C7DDDE60572 |  | 2011-03-04 | Use As Is
354 | OML-0007 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | C100-CAV-INSP-ADPT | R2 | 32 | LIGHT SMUDGE SPOTS ON CERAMIC SURFACE. SEE ATTACHMENT FOR DIMENSIONAL NONCONFORM | ogle | 2011-03-04 |  |  |  | HKPARK | 2011-03-07 | Proceed to the next step. RF test will result will show how the smudge affect th | Use As Is | [14 empty] | OML-0007 | ogle | 2011-03-04 | 8359B7F2-0F30-B12F-F29F45A2167DC649 |  | 2011-03-07 | Use As Is
355 | OML-0008 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | C100-CAV-INSP-ADPT | R2 | 33 | VERY LIGHT SMUDGE SPOTS ON CERAMIC SURFACE.SEE TRAVELER FOR PHOTO. | ogle | 2011-03-04 |  |  |  | HKPARK | 2011-03-07 | Proceed to the next step. RF test will result will show how the smudge affect th | Use As Is | [14 empty] | OML-0008 | ogle | 2011-03-04 |  |  | 2011-03-07 | Use As Is
356 | OML-0009 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | C100-CAV-INSP-ADPT | R2 | 34 | SEE ATTACHMENT FOR DIMENSIONAL NONCONFORMING. | ogle | 2011-03-04 |  |  |  | HKPARK | 2011-03-07 | Thinner ceramic won't affect RF and the measured deviation won't jeopardize the | Use As Is | [14 empty] | OML-0009 | ogle | 2011-03-04 | 83689AE7-099C-4AC7-307F5F12D12FCF56 |  | 2011-03-07 | Use As Is
357 | OML-0010 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | C100-CAV-INSP-ADPT | R2 | 35 | SEE ATTACHMENT FOR DIMENSIONAL NONCONFORMING. | ogle | 2011-03-04 |  |  |  | HKPARK | 2011-03-07 |  | Use As Is | [14 empty] | OML-0010 | ogle | 2011-03-04 | 836F76A3-FB7F-6336-7B1AB7AE54B64B30 |  | 2011-03-07 | Use As Is
358 | OML-0011 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | C100-CAV-INSP-ADPT | R2 | 36 | VERY LIGHT SMUDGE SPOTS ON CERAMIC SURFACE. | ogle | 2011-03-04 |  |  |  | HKPARK | 2011-03-07 | Proceed to the next step. RF test will result will show how the smudge affect th | Use As Is | [14 empty] | OML-0011 | ogle | 2011-03-04 | 83A6333B-F8D0-A44F-3F79C81D8D9A151B |  | 2011-03-07 | Use As Is
359 | OML-0012 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | C100-CAV-INSP-ADPT | R2 | 37 | VERY LIGHT SMUDGE SPOTS ON CERAMIC SURFACE.SEE ATTACHMENT FOR DIMENSIONAL NONCON | ogle | 2011-03-04 |  |  |  | HKPARK | 2011-03-07 |  | Use As Is | [14 empty] | OML-0012 | ogle | 2011-03-04 | 83ACA84B-D21C-02C6-730DD77813813AF1 |  | 2011-03-07 | Use As Is
360 | OML-0013 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | C100-CAV-INSP-ADPT | R2 | 38 | VERY LIGHT SMUDGE SPOTS ON CERAMIC SURFACE. | ogle | 2011-03-04 |  |  |  | HKPARK | 2011-03-07 | Proceed to the next step. RF test will result will show how the smudge affect th | Use As Is | [14 empty] | OML-0013 | ogle | 2011-03-04 | 83B172CB-C5D4-76E5-5C3322F205DA6E7C |  | 2011-03-07 | Use As Is
361 | OML-0014 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | C100-CAV-INSP-ADPT | R2 | 39 | VERY LIGHT SMUDGE SPOTS ON CERAMIC SURFACE. | ogle | 2011-03-04 |  |  |  | HKPARK | 2011-03-07 | Proceed to the next step. RF test will result will show how the smudge affect th | Use As Is | [14 empty] | OML-0014 | ogle | 2011-03-04 | 83B508FA-AD02-EF16-BBD3574948F5E4F5 |  | 2011-03-07 | Use As Is
362 | OML-0015 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | C100-CAV-INSP-ADPT | R2 | 40 | VERY LIGHT SMUDGE SPOTS ON CERAMIC SURFACE. | ogle | 2011-03-04 |  |  |  | HKPARK | 2011-03-07 | Proceed to the next step. RF test will result will show how the smudge affect th | Use As Is | [14 empty] | OML-0015 | ogle | 2011-03-04 | 83BFC1DE-958D-01C9-26D8FA4692401605 |  | 2011-03-07 | Use As Is
363 | OML-0017 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | C100-CAV-INSP-ADPT | R2 | 42 | VERY LIGHT SMUDGE SPOTS ON CERAMIC SURFACE. | ogle | 2011-03-04 |  |  |  | HKPARK | 2011-03-07 | Proceed to the next step. RF test will result will show how the smudge affect th | Use As Is | [14 empty] | OML-0017 | ogle | 2011-03-04 | 83C66778-0F62-340C-641CFAC330E378AD |  | 2011-03-07 | Use As Is
364 | OML-0022 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | C100-CAV-INSP-ADPT | R2 | 47 | SEE ATTACHMENT FOR DIMENSIONAL NONCONFORMING. | ogle | 2011-03-04 |  |  |  | HKPARK | 2011-03-07 |  | Use As Is | [14 empty] | OML-0022 | ogle | 2011-03-04 | 83EFD621-9D52-08EA-50CCC849FDAFF44E |  | 2011-03-07 | Use As Is
365 | OML-0023 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | C100-CAV-INSP-ADPT | R2 | 48 | SEE ATTACHMENT FOR DIMENSIONAL NONCONFORMING | ogle | 2011-03-04 |  |  |  | HKPARK | 2011-03-07 |  | Use As Is | [14 empty] | OML-0023 | ogle | 2011-03-04 | 83F54D1E-E9A3-7934-E64BBBA96463CC47 |  | 2011-03-07 | Use As Is
366 | OML-0024 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | C100-CAV-INSP-ADPT | R2 | 49 | VERY LIGHT SMUDGE SPOTS ON CERAMIC SURFACE. | ogle | 2011-03-04 |  |  |  | HKPARK | 2011-03-07 | Proceed to the next step. RF test will result will show how the smudge affect th | Use As Is | [14 empty] | OML-0024 | ogle | 2011-03-04 | 840080A2-B252-4C82-3C10F1A91550BD23 |  | 2011-03-07 | Use As Is
367 | OML-0027 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | C100-CAV-INSP-ADPT | R2 | 52 | VERY LIGHT SMUDGE SPOTS ON CERAMIC SURFACE. | ogle | 2011-03-04 |  |  |  | HKPARK | 2011-03-07 | Proceed to the next step. RF test will result will show how the smudge affect th | Use As Is | [14 empty] | OML-0027 | ogle | 2011-03-04 | 842F04EF-9581-99D9-05E601F023E55B19 |  | 2011-03-07 | Use As Is
368 | OML-0028 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | C100-CAV-INSP-ADPT | R2 | 53 | VERY LIGHT SMUDGE SPOTS ON CERAMIC SURFACE. | ogle | 2011-03-04 |  |  |  | HKPARK | 2011-03-07 | Proceed to the next step. RF test will result will show how the smudge affect th | Use As Is | [14 empty] | OML-0028 | ogle | 2011-03-04 | 8431DE24-E5C6-0F73-E6C49CA81DB1C721 |  | 2011-03-07 | Use As Is
369 | OML-0030 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | C100-CAV-INSP-ADPT | R2 | 55 | VERY LIGHT SMUDGE SPOTS ON CERAMIC SURFACE. | ogle | 2011-03-04 |  |  |  | HKPARK | 2011-03-07 | Proceed to the next step. RF test will result will show how the smudge affect th | Use As Is | [14 empty] | OML-0030 | ogle | 2011-03-04 | 8436E295-D21C-C6F1-B5CBA7C99F2172B4 |  | 2011-03-07 | Use As Is
370 | C100-RI-028 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP2 | R3 | 28 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING. | ogle | 2011-03-07 |  |  |  | MARHAUSE | 2011-03-08 |  | Use As Is | [14 empty] | C100-RI-028 | ogle | 2011-03-07 | 925F100F-B93C-EABC-E489E4AAB3246DEB |  | 2011-03-08 | Use As Is
371 | C100-RI-016 |  |  | Cavity | NCR associated with step 4. | C100-CAV-TUNE | R3 | 31 | The cavity is no longer in spec. for field flatness after having the helium vess | forehand | 2011-03-08 |  |  |  | MARHAUSE | 2011-03-29 | do not use cavity until further notice. | Modify | [14 empty] | C100-RI-016 | forehand | 2011-03-08 |  | hogan | 2011-03-29 | Modify
372 | C100-RI-027 |  | CRM120-7020-000 | C100 Helium Vessel Assembly |  | C100-CAV-INST-HELV | R1 | 24 | Helium Vessel is rotated with respect to cavity 1.5 degrees. See CMM report. | carpente | 2011-03-09 |  |  |  | MARHAUSE | 2011-03-29 | discuss if rotation is an issue with regard to alignment during horizontal assem | Modify | [14 empty] | C100-RI-027 | carpente | 2011-03-09 | 9AC06AF0-C4D2-4F48-13CAEAE090B7260D |  | 2011-03-29 | Modify
373 | [5 empty] | C100-CAV-VTRF | R3 | 47 | Cavity lock frequency of 1496.3052 MHz, below low-specification of 1496.400 MHz |  | 2011-03-10 |  |  |  | MARHAUSE | 2011-03-10 |  | Use As Is | [16 empty] | 2011-03-10 |  |  | 2011-03-10 | Use As Is
374 | C100-RI-024 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 61 | LIGHT SCRATCHES ON FPC FLANGE.SEE ATTACHMENT FOR DIMENSIONAL NONCONFORMING. | ogle | 2011-03-10 |  |  |  | MARHAUSE | 2011-03-11 | FPC sealing surface needs to be polished accordingly. | Modify | [14 empty] | C100-RI-024 | ogle | 2011-03-10 | A214962E-9FD4-B0B0-399ABBA0E6D5442D |  | 2011-03-11 | Modify
375 | F067764-10 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-RWRK-WGD | R3 | 18 | Flatness measured 0.005" Repair not possible in lapping room, no tooling to fix | georged | 2011-03-11 |  |  |  | KWILSON |  |  | Use As Is | [14 empty] | F067764-10 | georged | 2011-03-11 |  |  |  | Use As Is
376 | MPF-0039 |  | CRM120-7014-0020 | HOM Feedthru |  | C100-CAV-D3 | R1 | 67 | Center pin on HOM feedthru was noticed to be loose during a cavity assembly by C | carpente | 2011-03-11 |  |  |  | HKPARK | 2011-03-17 |  | Scrap | [14 empty] | MPF-0039 | carpente | 2011-03-11 |  | hkpark | 2011-03-17 | Scrap
377 | F064155-03 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-RWRK-WGD | R3 | 22 | Nicks on knife edge, damage repairable | georged | 2011-03-14 |  |  |  | KWILSON |  | Please repair knife edges. | Modify | [14 empty] | F064155-03 | georged | 2011-03-14 |  |  |  | Modify
378 | C100-RI-030 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP2 | R3 | 31 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | ogle | 2011-03-14 |  |  |  | MARHAUSE | 2011-03-14 |  | Use As Is | [14 empty] | C100-RI-030 | ogle | 2011-03-14 | B619732B-974D-70E0-B0F056B5EAAF4ADB |  | 2011-03-14 | Use As Is
379 | C100-RI-031 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP2 | R3 | 32 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | ogle | 2011-03-15 |  |  |  | MARHAUSE | 2011-03-15 |  | Use As Is | [14 empty] | C100-RI-031 | ogle | 2011-03-15 | BB3C7031-F8DC-9309-467D8C4FA701738D |  | 2011-03-15 | Use As Is
380 | OML-0036 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | C100-CAV-INSP-ADPT | R3 | 62 | VERY LIGHT SMUDGE SPOTS ON CERAMIC SURFACE. | ogle | 2011-03-15 |  |  |  | HKPARK | 2011-03-17 | The smudge was reported by the vendor and we tested samples to find it's OK. | Use As Is | [14 empty] | OML-0036 | ogle | 2011-03-15 | BCACCB85-BE19-2CED-ED785B77258B2D5E |  | 2011-03-17 | Use As Is
381 | OML-0037 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | C100-CAV-INSP-ADPT | R3 | 63 | VERY LIGHT SMUDGE SPOTS ON CERAMIC SURFACE. | ogle | 2011-03-15 |  |  |  | HKPARK | 2011-03-17 | The smudge was reported by the vendor and we tested samples to find it's OK. | Use As Is | [14 empty] | OML-0037 | ogle | 2011-03-15 | BCADB20C-B43E-18E9-922DE0DC9B45E06D |  | 2011-03-17 | Use As Is
382 | OML-0039 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | C100-CAV-INSP-ADPT | R3 | 65 | VERY LIGHT SMUDGE SPOTS ON CERAMIC SURFACE. | ogle | 2011-03-15 |  |  |  | HKPARK | 2011-03-17 | The smudge was reported by the vendor and we tested samples to find it's OK. | Use As Is | [14 empty] | OML-0039 | ogle | 2011-03-15 | BCAEBEFA-FA90-E485-85FCE0C263B05FBB |  | 2011-03-17 | Use As Is
383 | OML-0040 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | C100-CAV-INSP-ADPT | R3 | 66 | VERY LIGHT SMUDGE SPOTS ON CERAMIC SURFACE. | ogle | 2011-03-15 |  |  |  | HKPARK | 2011-03-17 | The smudge was reported by the vendor and we tested samples to find it's OK. | Use As Is | [14 empty] | OML-0040 | ogle | 2011-03-15 | BCAFCC50-D352-DC76-D4A501C0D01D69A3 |  | 2011-03-17 | Use As Is
384 | OML-0041 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | C100-CAV-INSP-ADPT | R3 | 67 | VERY LIGHT SMUDGE SPOTS ON CERAMIC SURFACE. | ogle | 2011-03-15 |  |  |  | HKPARK | 2011-03-17 | The smudge was reported by the vendor and we tested samples to find it's OK. | Use As Is | [14 empty] | OML-0041 | ogle | 2011-03-15 | BCB0AF18-C4A1-365E-140ACA3E1F715643 |  | 2011-03-17 | Use As Is
385 | OML-0042 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | C100-CAV-INSP-ADPT | R3 | 68 | VERY LIGHT SMUDGE SPOTS ON CERAMIC SURFACE. SEE ATTACHMENT FOR DIMENSIONAL NON- | ogle | 2011-03-15 |  |  |  | HKPARK | 2011-03-17 | Smudge was tested OK. The reported thickness range is not affecting the RF perf | Use As Is | [14 empty] | OML-0042 | ogle | 2011-03-15 | BCB1E185-00E1-1C86-63652BEE0FE821C2 |  | 2011-03-17 | Use As Is
386 | OML-0043 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | C100-CAV-INSP-ADPT | R3 | 69 | VERY LIGHT SMUDGE SPOTS ON CERAMIC SURFACE. BRAZE MATERIAL ON COPPER EYELET. BLA | ogle | 2011-03-15 |  |  |  | HKPARK | 2011-03-17 | Smudge was tested OK. The reported thickness range is not affecting the RF perf | Use As Is | [14 empty] | OML-0043 | ogle | 2011-03-15 | BCB4786A-96D9-E332-0282CB4B220E8F0E |  | 2011-03-17 | Use As Is
387 | OML-0044 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | C100-CAV-INSP-ADPT | R3 | 70 | VERY LIGHT SMUDGE SPOTS ON CERAMIC SURFACE. SEE ATTACHMENT FOR DIMENSIONAL NON-C | ogle | 2011-03-15 |  |  |  | HKPARK | 2011-03-17 | Smudge was tested OK. The reported thickness range is not affecting the RF perf | Use As Is | [14 empty] | OML-0044 | ogle | 2011-03-15 | BCB6B528-C63C-0488-22F8622DA71CAFF3 |  | 2011-03-17 | Use As Is
388 | MPF-0035 |  | CRM-120-7014-0020 | HOM Feedthru |  | C100-CAV-INSP-HMFT | R3 | 76 | Antenna completely broken off, with corresponding loose center pin. | carpente | 2011-03-16 |  |  |  | HKPARK | 2011-03-17 |  | Scrap | [14 empty] | MPF-0035 | carpente | 2011-03-16 | BFC1EFF9-005F-A2A4-45F044862F293D01 |  | 2011-03-17 | Scrap
389 | MPF-0022 |  | CRM-120-7014-0020 | HOM Feedthru |  | C100-CAV-INSP-HMFT | R3 | 71 | Pin is extremely loose. BC 03/16/11 | carpente | 2011-03-16 |  |  |  | HKPARK | 2011-03-17 |  | Scrap | [14 empty] | MPF-0022 | carpente | 2011-03-16 |  |  | 2011-03-17 | Scrap
390 | MPF-0026 |  | CRM-120-7014-0020 | HOM Feedthru |  | C100-CAV-INSP-HMFT | R3 | 72 | Pin is extremely loose. BC 03/16/11 | carpente | 2011-03-16 |  |  |  | HKPARK | 2011-04-01 |  | Scrap | [14 empty] | MPF-0026 | carpente | 2011-03-16 |  |  | 2011-04-01 | Scrap
391 | FLG0037-037 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 28 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | ogle | 2011-03-16 |  |  |  | HKPARK | 2011-03-17 |  | Use As Is | [14 empty] | FLG0037-037 | ogle | 2011-03-16 | C167C0D8-0583-5799-AA17A0C13DC3DA4B |  | 2011-03-17 | Use As Is
392 | FLG0037-038 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 29 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | ogle | 2011-03-16 |  |  |  | HKPARK | 2011-03-17 |  | Use As Is | [14 empty] | FLG0037-038 | ogle | 2011-03-16 | C169E6C1-ACC5-9705-4505D51B5351304A |  | 2011-03-17 | Use As Is
393 | FLG0037-040 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 31 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | ogle | 2011-03-16 |  |  |  | HKPARK | 2011-03-17 |  | Use As Is | [14 empty] | FLG0037-040 | ogle | 2011-03-16 | C16BD6EE-C498-C503-4C924E1A263BBDE6 |  | 2011-03-17 | Use As Is
394 | FLG0037-042 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 33 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | ogle | 2011-03-16 |  |  |  | HKPARK | 2011-03-17 |  | Use As Is | [14 empty] | FLG0037-042 | ogle | 2011-03-16 | C17DAF5E-9063-7040-51E623C5AB22A54B |  | 2011-03-17 | Use As Is
395 | FLG0037-043 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 34 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | ogle | 2011-03-16 |  |  |  | HKPARK | 2011-03-17 |  | Use As Is | [14 empty] | FLG0037-043 | ogle | 2011-03-16 | C17EEE0F-EE04-52BC-F0D5599FEDD6C31C |  | 2011-03-17 | Use As Is
396 | FLG0037-044 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 35 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | ogle | 2011-03-16 |  |  |  | HKPARK | 2011-03-17 |  | Use As Is | [14 empty] | FLG0037-044 | ogle | 2011-03-16 | C1800DFE-9B46-1A18-0B75BA739990964C |  | 2011-03-17 | Use As Is
397 | FLG0037-046 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 37 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | ogle | 2011-03-16 |  |  |  | HKPARK | 2011-03-17 |  | Use As Is | [14 empty] | FLG0037-046 | ogle | 2011-03-16 | C1819C1F-D82C-AACA-5F886176515CB749 |  | 2011-03-17 | Use As Is
398 | FLG0037-047 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 38 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | ogle | 2011-03-16 |  |  |  | HKPARK | 2011-03-17 |  | Use As Is | [14 empty] | FLG0037-047 | ogle | 2011-03-16 | C18347AA-F1D7-7178-C6CB2AA3E8DC41A5 |  | 2011-03-17 | Use As Is
399 | FLG0037-048 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 39 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | ogle | 2011-03-16 |  |  |  | HKPARK | 2011-03-17 |  | Use As Is | [14 empty] | FLG0037-048 | ogle | 2011-03-16 | C184CA02-D07E-ECD9-4EAA2D0CB0627895 |  | 2011-03-17 | Use As Is
400 | FLG0037-049 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 40 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | ogle | 2011-03-16 |  |  |  | HKPARK | 2011-03-17 | Tag it and put aside. | Return To Vendor | [14 empty] | FLG0037-049 | ogle | 2011-03-16 | C1C65C4D-BF42-85F9-9464607842794400 |  | 2011-03-17 | Return To Vendor
401 | FLG0037-050 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 41 | SAMPLE ADAPTER WILL NOT GO. SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | ogle | 2011-03-16 |  |  |  | HKPARK | 2011-03-17 | Tag it and put it aside. These will be returned to the vendor to get them replac | Return To Vendor | [14 empty] | FLG0037-050 | ogle | 2011-03-16 | C1D2EC12-F760-5C4A-C89BB4AB64CB22C3 |  | 2011-03-17 | Return To Vendor
402 | F067763-06 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R4 | 61 | Repairable scratches on small flange | georged | 2011-03-17 |  |  |  | KWILSON |  | Please repair scratches. | Modify | [14 empty] | F067763-06 | georged | 2011-03-17 |  |  |  | Modify
403 | F067763-11 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R4 | 64 | Surface finish out of tolerance, meas. 19 micro inches. Repairable scratches on | georged | 2011-03-17 |  |  |  | KWILSON |  | Please repair scratches. | Modify | [14 empty] | F067763-11 | georged | 2011-03-17 |  |  |  | Modify
404 | F067764-09 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R5 | 93 | Repairable scratches on small flange. Flatness out of tolerance. Meas. 0.005" | georged | 2011-03-17 |  |  |  | KWILSON |  | Please repair scratches. Will accept flatness. | Modify | [14 empty] | F067764-09 | georged | 2011-03-17 |  |  |  | Modify
405 | OML-0037 |  | CRM1207075-0024 | C100 Window Adapter |  | C100-CAV-INSP-ADPT | R3 | 63 | Failed leak test | carpente | 2011-03-18 |  |  |  | HKPARK | 2011-03-18 | The adapter will be replaced by the vendor. | Scrap | [14 empty] | OML-0037 | carpente | 2011-03-18 |  |  | 2011-03-18 | Scrap
406 | MPF-034 |  | CRM1207014-0020 | HOM Feedthru |  | C100-CAV-INSP-HMFT | R3 | 75 | Antenne now broken as of 3/18/11, found during additional leak test on same date | carpente | 2011-03-21 |  |  |  | HKPARK | 2011-03-21 |  | Scrap | [14 empty] | MPF-034 | carpente | 2011-03-21 |  | hogan | 2011-03-21 | Scrap
407 | F067764-22 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R5 | 92 | Scratches on small flange. Nick on knife edge. During leak check, won't seal unl | georged | 2011-03-22 |  |  |  | KWILSON |  | Repair scratches and knife edge. Bolts often torqued to 140 in-lb, so 125 is ok | Modify | [14 empty] | F067764-22 | georged | 2011-03-22 |  |  |  | Modify
408 | C100-RI-032 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP2 | R3 | 33 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | ogle | 2011-03-22 |  |  |  | MARHAUSE | 2011-03-22 |  | Use As Is | [14 empty] | C100-RI-032 | ogle | 2011-03-22 | DFD34D2F-9459-3F91-038A731E94F40646 |  | 2011-03-22 | Use As Is
409 | FLG0037-039 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 30 | DAMAGE TO KNIFE EDGE. SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | ogle | 2011-03-22 |  |  |  | HKPARK | 2011-03-23 | Put it aside. The vendor has to replace it. | Scrap | [14 empty] | FLG0037-039 | ogle | 2011-03-22 | E0A92236-F544-DCF5-3F573278020A1436 |  | 2011-03-23 | Scrap
410 | FLG0037-041 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 32 | DAMAGED KNIFE EDGE. SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | ogle | 2011-03-22 |  |  |  | HKPARK | 2011-03-23 | Put it aside. The vendor has to replace it. | Scrap | [14 empty] | FLG0037-041 | ogle | 2011-03-22 | E0AE02E0-AAEE-B6BD-DCBFA59C389AC767 |  | 2011-03-23 | Scrap
411 | FLG0037-045 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 36 | DAMAGED KNIFE EDGE. SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | ogle | 2011-03-22 |  |  |  | HKPARK | 2011-03-23 | Put it aside. The vendor should replace it. | Scrap | [14 empty] | FLG0037-045 | ogle | 2011-03-22 | E0B0859B-FE27-649D-07BB5ABBA85F8689 |  | 2011-03-23 | Scrap
412 | F067763-14 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R4 | 49 | Scratches on small flange. Black stain and chip inside waveguide. Nicks on knif | georged | 2011-03-23 |  |  |  | KWILSON |  | Please repair scratches on small flange and nick on knife edge. Black stain of | Modify | [14 empty] | F067763-14 | georged | 2011-03-23 |  |  |  | Modify
413 | C100-RI-027 |  | CRM120-7020-0000 | C100 Helium Vessel Assembly |  | C100-CAV-INST-HELV | R1 | 24 | After re-welding the tuner tabs on the cavity dishes, (2) of the tabs are mis-lo | carpente | 2011-03-24 |  |  |  | HOGAN | 2011-03-25 | The rotation of the tabs is correct. The 0.10" offset will be compensated for a | Use As Is | [14 empty] | C100-RI-027 | carpente | 2011-03-24 |  |  | 2011-03-25 | Use As Is | N/A
414 | F067764-03 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R5 | 96 | Scratches on small flange. Copper blister inside waveguide. | georged | 2011-03-24 |  |  |  | KWILSON |  | Repair scratches. I think blister is actually a plated weld bead. | Modify | [14 empty] | F067764-03 | georged | 2011-03-24 |  |  |  | Modify
415 | MPF-0028 |  | CRM120-7014-0020 | HOM Probe |  | C100-CAV-INSP-HMFT | R3 | 68 | HOM Probe failed after cavity string assembly while still in clean room. Sapphir | carpente | 2011-03-24 |  |  |  | HKPARK | 2011-03-24 |  | Scrap | [14 empty] | MPF-0028 | carpente | 2011-03-24 |  |  | 2011-03-24 | Scrap
416 | F067764-01 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R5 | 94 | Copper flaking inside the waveguide. Nick on knife edge. Scratches on small flan | georged | 2011-03-25 |  |  |  | KWILSON |  | Repair scratches and knife edge. Copper ok. | Use As Is | [14 empty] | F067764-01 | georged | 2011-03-25 |  | carpente |  | Use As Is
417 | OML-0045 |  | CRM12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 72 | BLACK RESIDUE ON COPPER EYELET SURFACE. | ogle | 2011-03-25 |  |  |  | HKPARK | 2011-04-01 | The residue is likely from the brazing tool. If it comes off use as is. | Use As Is | [14 empty] | OML-0045 | ogle | 2011-03-25 | EFC85C51-B573-705F-35CE5B723028C933 |  | 2011-04-01 | Use As Is
418 | OML-0046 |  | CRM120-70750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 73 | BLACK RESIDUE ON COPPER EYELET SURFACE. | ogle | 2011-03-25 |  |  |  | HKPARK | 2011-04-01 | The residue is likely from the brazing fixture. If it comes off use as is. | Use As Is | [14 empty] | OML-0046 | ogle | 2011-03-25 | EFCB0B9A-F867-FB59-80B974FD3EE654F8 |  | 2011-04-01 | Use As Is
419 | OML-0047 |  | CRM-12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 74 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | ogle | 2011-03-25 |  |  |  | HKPARK | 2011-04-01 | The deviation amount not critical. | Use As Is | [14 empty] | OML-0047 | ogle | 2011-03-25 | EFCDBFAD-F077-296B-55BDCA5B4360F4B1 |  | 2011-04-01 | Use As Is
420 | OML-0048 |  | CRM12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 75 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | ogle | 2011-03-25 |  |  |  | HKPARK | 2011-04-01 | The deviation amount not critical. | Use As Is | [14 empty] | OML-0048 | ogle | 2011-03-25 | EFD2F67E-F72F-DD60-1503058AA0110EB1 |  | 2011-04-01 | Use As Is
421 | OML-0049 |  | CRM12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 76 | BLACK RESIDUE ON COPPER EYELET SURFACE.BRAZE MATERIAL ON COPPER EYELET SURFACE. | ogle | 2011-03-25 |  |  |  | HKPARK | 2011-04-01 | The residue is likely from the brazing fixture. If it comes off use as is. The | Use As Is | [14 empty] | OML-0049 | ogle | 2011-03-25 | EFD642B9-FCDA-D9EA-4868990751ED1219 |  | 2011-04-01 | Use As Is
422 | OML-0050 |  | CRM12070750024 |  |  | C100-CAV-INSP-ADPT | R3 | 77 | BLACK RESIDUE ON COPPER EYELET SURFACE. | ogle | 2011-03-25 |  |  |  | HKPARK | 2011-04-01 | The residue is likely from the brazing fixture. If it comes off use as is. | Use As Is | [14 empty] | OML-0050 | ogle | 2011-03-25 | EFD8AC97-E1DC-FA25-D9F3D63ED59C4D48 |  | 2011-04-01 | Use As Is
423 | OML-0051 |  | CRM-12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 78 | BLACK RESIDUE ON COPPER EYELET SURFACE. VOIDS IN BRAZE SEAM BETWEEN COPPER EYELE | ogle | 2011-03-25 |  |  |  | HKPARK | 2011-06-30 | The residue is likely from the brazing fixture. If it comes off use as is. Thi | Return To Vendor | [14 empty] | OML-0051 | ogle | 2011-03-25 | EFE412FA-C278-919B-4FA841654B9FDD76 |  | 2011-06-30 | Return To Vendor
424 | OML-0052 |  | CRM-12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 79 | BLACK RESIDUE ON COPPER EYELET SURFACE. VOIDS IN BRAZE SEAM BETWEEN COPPER EYELE | ogle | 2011-03-25 |  |  |  | HKPARK | 2011-04-01 | The residue is likely from the brazing fixture. If it comes off use as is. Thi | Use As Is | [14 empty] | OML-0052 | ogle | 2011-03-25 | EFE6AF79-AF22-1EF4-FB5B6BD9EF0E4B53 |  | 2011-04-01 | Use As Is
425 | OML-0053 |  | CRM-12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 80 | VOIDS IN BRAZE SEAM BETWEEN COPPER EYELET AND WINDOW DETAIL. | ogle | 2011-03-25 |  |  |  | HKPARK | 2011-04-01 | This gap between adapter and eyelet can be either filled with the braze alloy or | Use As Is | [14 empty] | OML-0053 | ogle | 2011-03-25 | EFE8EA0E-F0D7-D3F5-890CCD3AC721BD9F |  | 2011-04-01 | Use As Is
426 | OML-0054 |  | CRM-12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 81 | BLACK RESIDUE ON COPPER EYELET SURFACE. | ogle | 2011-03-25 |  |  |  | HKPARK | 2011-04-01 | The residue is likely from the brazing fixture. If it comes off use as is. | Use As Is | [14 empty] | OML-0054 | ogle | 2011-03-25 | EFEAA2A5-DB9F-DEED-2F3BDBE75D09108A |  |  | Use As Is
427 | OML-0054 |  | CRM-12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 81 | BLACK RESIDUE ON COPPER EYELET SURFACE. | ogle | 2011-03-25 |  |  |  | HKPARK |  | cleaned and welded. | [15 empty] | OML-0054 | ogle | 2011-03-25 | EFF0AE95-C3BB-12BE-702F81C6B5D08AF6 |  |  | Modify
428 | OML-0055 |  | CRM-12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 82 | BLACK RESIDUE ON COPPER EYELET SURFACE. | ogle | 2011-03-25 |  |  |  | HKPARK | 2011-04-01 | The residue is likely from the brazing fixture. If it comes off use as is. | Use As Is | [14 empty] | OML-0055 | ogle | 2011-03-25 | EFF31857-9B79-7CDE-226DC18411B8EA39 |  | 2011-04-01 | Use As Is
429 | C-100-RI-026 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 62 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | ogle | 2011-03-28 |  |  |  | MARHAUSE | 2011-03-29 |  | Use As Is | [14 empty] | C-100-RI-026 | ogle | 2011-03-28 | FEAFBF66-F576-C833-29CE61EFE4275BFE |  | 2011-03-29 | Use As Is
430 | C100-RI-039 |  | CRM12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 63 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | ogle | 2011-03-29 |  |  |  | MARHAUSE | 2011-04-04 |  | Use As Is | [14 empty] | C100-RI-039 | ogle | 2011-03-29 | 047A4EB8-D122-54CF-30661DB910C2D114 |  | 2011-04-04 | Use As Is
431 | 47726-0001 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 96 | SPOT OF BRAZE RESIDUE IN SEAL GROOVE. NICK IN SEAL GROOVE. SEE ATTACHMENT FOR | ogle | 2011-03-29 |  |  |  | HKPARK | 2011-03-30 |  | Scrap | [14 empty] | 47726-0001 | ogle | 2011-03-29 | 04B34156-BFE0-D1F1-9C6A06A959A4E54B |  | 2011-03-30 | Scrap
432 | 47726-0005 |  | CRM1207014-0020 | HOM Feedthru |  | C100-CAV-INSP-HMFT | R3 | 97 | Center pin is completely broken off. | carpente | 2011-03-30 |  |  |  | HKPARK | 2011-03-30 |  | Scrap | [14 empty] | 47726-0005 | carpente | 2011-03-30 |  |  | 2011-03-30 | Scrap
433 | C100-RI-033 |  | 12070130001 | CAVITY |  | C100-CAV-INSP2 | R3 | 35 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | ogle | 2011-04-01 |  |  |  | MARHAUSE | 2011-04-04 |  | Use As Is | [14 empty] | C100-RI-033 | ogle | 2011-04-01 | 12C40642-C2B1-D3FF-22BAF556D8F93384 |  | 2011-04-04 | Use As Is
434 | C100-RI-008 |  | 12070130001 | CAVITY |  | C100-CAV-INSP2 | R3 | 30 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | ogle | 2011-04-01 |  |  |  | MARHAUSE | 2011-04-04 |  | Use As Is | [14 empty] | C100-RI-008 | ogle | 2011-04-01 | 12D0EA3F-D033-B9F8-0DC5A655057F6478 |  | 2011-04-04 | Use As Is
435 | FLG0037-054 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 45 | DAMAGE TO OUTSIDE EDGE. SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | ogle | 2011-04-01 |  |  |  | HKPARK | 2011-04-04 | The flange is functional but do not use at this time. | Scrap | [14 empty] | FLG0037-054 | ogle | 2011-04-01 | 13C04F76-FA35-72F6-66BE67B15FA352E5 |  | 2011-04-04 | Scrap
436 | FLG0037-055 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 46 |  | ogle | 2011-04-01 |  |  |  | HKPARK | 2011-04-04 | The deviation value too high for a safe connection. | Scrap | [14 empty] | FLG0037-055 | ogle | 2011-04-01 | 13D0E80B-D97E-6E1C-C97DF60FA64B4B07 |  | 2011-04-04 | Scrap
437 | FLG0037-056 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 47 | POCKET WILL NOT ACCEPT TEST ADAPTER. SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORM | ogle | 2011-04-01 |  |  |  | HKPARK | 2011-04-04 | The flange should be replaced by the vendor. Tag it and store. | Scrap | [14 empty] | FLG0037-056 | ogle | 2011-04-01 | 13D2FA55-F563-8368-EF5C6FEBA1D871D9 |  | 2011-04-04 | Scrap
438 | FLG0037-059 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 50 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | ogle | 2011-04-01 |  |  |  | HKPARK | 2011-04-04 | These thru holes are oversized. The deviation is low enough to make a connection | Use As Is | [14 empty] | FLG0037-059 | ogle | 2011-04-01 | 13D53F4B-A073-46BE-DFA657C482FAAF20 |  | 2011-04-04 | Use As Is
439 | FLG0037-060 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 51 | SMALL NICK ON KNIFE EDGE. | ogle | 2011-04-01 |  |  |  | HKPARK | 2011-04-04 | The vendor should replace this. tag it and store. | Scrap | [14 empty] | FLG0037-060 | ogle | 2011-04-01 | 13D71756-CE9D-4D05-0B9AB46EC1F0C06F |  | 2011-04-04 | Scrap
440 | FLG0037-061 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 52 | NICK ON KNIFE EDGE. | ogle | 2011-04-01 |  |  |  | HKPARK | 2011-04-04 | The vendor should replace this. tag it and store. | Scrap | [14 empty] | FLG0037-061 | ogle | 2011-04-01 | 13DABA01-A3E1-F40F-9F41CFB033877081 |  | 2011-04-04 | Scrap
441 | FLG0037-064 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 55 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | ogle | 2011-04-01 |  |  |  | HKPARK | 2011-04-04 | These thru holes are oversized. The deviation is low enough to make a connection | Use As Is | [14 empty] | FLG0037-064 | ogle | 2011-04-01 | 13E9BE9C-A8F5-BEDB-E03ADEF46A27D995 |  | 2011-04-04 | Use As Is
442 | FLG0037-067 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 58 | NICK ON KNIFE EDGE. | ogle | 2011-04-01 |  |  |  | HKPARK | 2011-04-04 | The vendor should replace this. tag it and store. | Scrap | [14 empty] | FLG0037-067 | ogle | 2011-04-01 | 13EB57E9-A829-35F9-EC9EA1BE539C427E |  | 2011-04-04 | Scrap
443 | FLG0037-068 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 59 | NICK ON KNIFE EDGE. SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | ogle | 2011-04-01 |  |  |  | HKPARK | 2011-04-04 | The vendor should replace this. tag it and store. | Scrap | [14 empty] | FLG0037-068 | ogle | 2011-04-01 | 13ECC630-AE9A-9443-DD27B717F30D56FC |  | 2011-04-04 | Scrap
444 | FLG0037-069 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 60 | NICK ON KNIFE EDGE. SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | ogle | 2011-04-01 |  |  |  | HKPARK | 2011-04-04 | The vendor should replace this. tag it and store. | Scrap | [14 empty] | FLG0037-069 | ogle | 2011-04-01 | 13EF2245-9396-153B-440FE4746B2CFB94 |  | 2011-04-04 | Scrap
445 | C100_RI_029 |  |  |  |  | C100-CAV-VTRF | R3 | 53 | C100_RI_029 quenching at 21.6 MV/m in the pi mode. Quench study suggests cell 3 |  | 2011-04-02 |  |  |  | MARHAUSE | 2011-04-04 | Follow-up discussion required. | Modify | [14 empty] | C100_RI_029 |  | 2011-04-02 |  |  | 2011-04-04 | Modify
446 | C100-RI-034 |  | 12070130001 | CAVITY |  | C100-CAV-INSP2 | R3 | 34 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | ogle | 2011-04-05 |  |  |  | MARHAUSE | 2011-04-06 |  | Use As Is | [14 empty] | C100-RI-034 | ogle | 2011-04-05 | 284E65C6-9CF1-374A-586B82D01F70E7D9 |  | 2011-04-06 | Use As Is
447 | [5 empty] | C100-CAV-VTRF | R3 | 54 | C100_RI_032 cavity lock frequency of 1496.8754 MHZ is higher than the high speci |  | 2011-04-06 |  |  |  | MARHAUSE | 2011-04-06 | frequency will drop in spec. after helium vessel welding | Use As Is | [16 empty] | 2011-04-06 |  |  | 2011-04-06 | Use As Is
448 | F067763-01 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R4 | 56 | One outer pattern tapped hole oversized. One inner pattern tapped hole doesn't m | georged | 2011-04-07 |  |  |  | KWILSON |  |  | Return To Vendor | [14 empty] | F067763-01 | georged | 2011-04-07 |  |  |  | Return To Vendor
449 | F067763-02 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R5 | 98 | Ding and scratches in seal surface on the small flange. | georged | 2011-04-07 |  |  |  | KWILSON |  | As dent in surface of small flange cannot be repaired,Return To Vendor. | Return To Vendor | [14 empty] | F067763-02 | georged | 2011-04-07 |  |  |  | Return To Vendor
450 | ISI-0001 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 98 | TOOL CHATTER IN SEAL GROOVE. CONTAMINATION INTERNAL NEAR SAPPHIRE INSERT CONNEC | ogle | 2011-04-07 |  |  |  | HKPARK |  | cleaned, polished and used | [15 empty] | ISI-0001 | ogle | 2011-04-07 |  |  |  | Modify
451 | ISI-0002 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 99 | TOOL CHATTER IN SEAL GROOVE. PROBE HAS BEEN POLISHED. CONTAMINATION INTERNAL C | ogle | 2011-04-07 |  |  |  | HKPARK |  | cleaned, polished then used but later taken off. | [15 empty] | ISI-0002 | ogle | 2011-04-07 |  |  |  | Scrap
452 | ISI-0003 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 100 | TOOL CHATTER IN SEAL GROOVE. PROBE HAS BEEN POLISHED. CONTAMINATION INTERNAL; | ogle | 2011-04-07 |  |  |  | HKPARK |  | cleaned, polished and used. | [15 empty] | ISI-0003 | ogle | 2011-04-07 |  |  |  | Modify
453 | RI-PRO-02 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 101 | LENGTH .152-.166 CHECKS: .170 | ogle | 2011-04-07 |  |  |  | HKPARK | [17 empty] | RI-PRO-02 | ogle | 2011-04-07 |  |  |  | Use As Is
454 | RI-PRO-05 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 102 | .156-.166 DEPTH CHECKS:.169 PROBE ANGLE | ogle | 2011-04-07 |  |  |  | HKPARK | 2011-08-18 | See other NCR associated to this part. | Scrap | [14 empty] | RI-PRO-05 | ogle | 2011-04-07 |  |  | 2011-08-18 | Scrap
455 | OML-0056 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 83 | SMALL AMOUNT OF BRAZE MATERIAL ON INSIDE EDGE OF COPPER EYELET. | ogle | 2011-04-07 |  |  |  | HKPARK | 2011-05-13 | RF test will determine the effect of the braze material. | Use As Is | [14 empty] | OML-0056 | ogle | 2011-04-07 | 322133BF-C6F5-A30D-6846AA0799F6F72C |  | 2011-05-13 | Use As Is
456 | OML-0057 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 84 | BLACK RESIDUE ON COPPER EYELET SURFACE. LIGHT SMUDGE SPOTS ON CERAMIC SURFACE. | ogle | 2011-04-07 |  |  |  | HKPARK | 2011-05-13 | Clean the residue. The smudge found harmless by earlier RF test. | Use As Is | [14 empty] | OML-0057 | ogle | 2011-04-07 | 32225AC6-B3D9-E4E4-307B8C3526F64B0C |  | 2011-05-13 | Use As Is
457 | OML-0058 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 85 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | ogle | 2011-04-07 |  |  |  | HKPARK | 2011-05-17 | The radii has changed to .688+/-.005. See the updated drawing crm-120-7075-0027. | Use As Is | [14 empty] | OML-0058 | ogle | 2011-04-07 | 3223EE54-DF74-137D-FF2ED62815994B96 |  | 2011-05-17 | Use As Is
458 | OML-0059 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 86 | BLACK RESIDUE ON COPPER EYELET.SPOT OF BRAZE MATERIAL ON CERAMIC SURFACE. | ogle | 2011-04-07 |  |  |  | HKPARK | 2011-05-13 | The braze spot should be removed. If it is not removed do not RF test yet. | Modify | [14 empty] | OML-0059 | ogle | 2011-04-07 | 32260559-DC4D-7A5D-D8B348EEF1465EB8 |  | 2011-05-13 | Modify | Black brazed spot has been removed by bead-blasting and has been cleaned. BC 06/
459 | OML-0060 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 87 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | ogle | 2011-04-07 |  |  |  | HKPARK | 2011-05-17 | The radii has changed to .688+/-.005. See the updated drawing crm-120-7075-0027. | Use As Is | [14 empty] | OML-0060 | ogle | 2011-04-07 | 3228780D-AEA4-B84C-8EA497E773515F26 |  | 2011-05-17 | Use As Is
460 | OML-0061 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 88 | SMALL BRAZE SPOT ON CERAMIC SURFACE. SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORM | ogle | 2011-04-07 |  |  |  | HKPARK | 2011-05-17 | Remove the braze spot by bead blast orScraping with glass edge. The radii has | Modify | [14 empty] | OML-0061 | ogle | 2011-04-07 | 322A51FA-0833-D803-41A1D8438C61614D |  | 2011-05-17 | Modify | Bead blasted and OK to weld-HyeKyoung 1/10/12
461 | OML-0062 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 89 | SMUDGE SPOTS ON CERAMIC SURFACE.LIGHT BLACK RESIDUE ON COPPER EYELET. | ogle | 2011-04-07 |  |  |  | HKPARK | 2011-05-13 | Clean the residue. The smudge found harmless by earlier RF test. The RF test wil | Use As Is | [14 empty] | OML-0062 | ogle | 2011-04-07 | 322C1B48-00F1-0FE0-BDE83DF78544B102 |  | 2011-05-13 | Use As Is | Part visually clean, shipping out for ion implant-HK 8/10/11
462 | OML-0063 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 90 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | ogle | 2011-04-07 |  |  |  | HKPARK | 2011-05-17 | The radii has changed to .688+/-.005. See the updated drawing crm-120-7075-0027. | Use As Is | [14 empty] | OML-0063 | ogle | 2011-04-07 | 322DE21D-93AC-38B8-EDB7046513908118 |  | 2011-05-17 | Use As Is
463 | OML-0064 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 91 | LIGHT SMUDGE SPOTS ON CERAMIC SURFACE. | ogle | 2011-04-07 |  |  |  | HKPARK | 2011-05-13 | The smudge found harmless by earlier RF test. The RF test will find out. | Use As Is | [14 empty] | OML-0064 | ogle | 2011-04-07 | 322F9B1A-0651-C195-FCD09F8A182ED65C |  | 2011-05-13 | Use As Is
464 | OML-0065 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 92 | BLACK RESIDUE ON COPPER EYELET. SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | ogle | 2011-04-07 |  |  |  | HKPARK | 2011-05-17 | Remove the residue. The radii has changed to .688+/-.005. See the updated drawi | Modify | [14 empty] | OML-0065 | ogle | 2011-04-07 | 32310C58-9818-0534-06C771B492192CA9 |  | 2011-05-17 | Modify | Part visually clean, shipping out for ion implant.-HP 8/10/11
465 | OML-0066 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 93 | BLACK RESIDUE ON COPPER EYELET. SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | ogle | 2011-04-07 |  |  |  | HKPARK | 2011-05-17 | Remove the residue. The ceramic thickness is OK. | Modify | [14 empty] | OML-0066 | ogle | 2011-04-07 | 326D50EA-F05D-A189-8E9E7E548FE95B1E |  | 2011-05-17 | Modify | Part visually clean, shipping out for ion implant-HP 8/10/11
466 | OML-0067 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 94 | LIGHT SMUDGE SPOTS ON CERAMIC SURFACE. | ogle | 2011-04-07 |  |  |  | HKPARK | 2011-05-17 | The spots are proved to be harmless by earlier RF test. | Use As Is | [14 empty] | OML-0067 | ogle | 2011-04-07 | 32748F48-03F0-935E-4F06ABC3D4C69723 |  | 2011-05-17 | Use As Is
467 | OML-0068 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 95 | BRAZE SPOT ON CERAMIC SURFACE. | ogle | 2011-04-07 |  |  |  | HKPARK | 2011-05-13 | The braze spot should be removed before the RF test. if it is not removed do not | Modify | [14 empty] | OML-0068 | ogle | 2011-04-07 | 32791389-A810-DD99-A0B595B9BCD479FD |  | 2011-05-13 | Modify | Black brazed spot has been removed by bead-blasting and has been cleaned. BC 06/
468 | OML-0069 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 96 | BRAZE SPOT ON CERAMIC SURFACE. | ogle | 2011-04-07 |  |  |  | HKPARK | 2011-08-10 | Passed RF test. Perform visual inspection again-HP 8/10/11 | Use As Is | [14 empty] | OML-0069 | ogle | 2011-04-07 | 327A4D08-C2FC-25DC-D37A1329C8A6C713 |  | 2011-08-10 | Use As Is
469 | OML-070 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 97 | SMUDGE SPOTS ON CERAMIC SURFACE. SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | OGLE | 2011-04-07 |  |  |  | HKPARK | 2011-08-10 | Radii meet the new spec. | Use As Is | [14 empty] | OML-070 | OGLE | 2011-04-07 | 327BB656-A05F-61B6-3E8FD1F033669280 |  | 2011-08-10 | Use As Is
470 | OML-0071 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 98 | BRAZE SPOT ON CERAMIC SURFACE | ogle | 2011-04-07 |  |  |  | HKPARK | 2011-05-13 | The braze spot should be removed before RF test. If it is not removed do not RF | Modify | [14 empty] | OML-0071 | ogle | 2011-04-07 | 32819BB6-CD3D-63B9-FB4C77F7115FEB48 |  | 2011-05-13 | Modify | Passed RF test. use as is, shipping out for ion implant-HP 8/10/11
471 | OML-0073 |  | 12070750024 |  |  | C100-CAV-INSP-ADPT | R3 | 100 | BRAZE MATERIAL ON BARE ADAPTER FACE. SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORM | OGLE | 2011-04-07 |  |  |  | HKPARK | 2011-09-19 | Radius has been increased and the braze material on SS is OK as long as it does | Use As Is | [14 empty] | OML-0073 | OGLE | 2011-04-07 | 3283E3AA-CC6F-5A79-BEC8E51174E0849B |  | 2011-09-19 | Use As Is
472 | ISI-0011 |  | CRM-120-7014-0020 | HOM Feedthru |  | C100-CAV-INSP-HMFT | R3 | 110 | Probe angle is 0.9. Out to tolerance 0.4 | carpente | 2011-04-08 |  |  |  | HKPARK | 2011-04-08 | Tag it and store for now. It's possible that this feedthru can be used. | Scrap | [14 empty] | ISI-0011 | carpente | 2011-04-08 |  |  | 2011-04-08 | Scrap
473 | MPF-0049 |  | CRM-120-7014-0020 | HOM Feedthru |  |  |  |  | Broken sapphire, antenna and center pin loose. | carpente | 2011-04-08 |  |  |  | HKPARK | 2011-04-08 |  | Scrap | [14 empty] | MPF-0049 | carpente | 2011-04-08 |  | hkpark | 2011-04-08 | Scrap
474 | MPF-0057 |  | CRM-120-7014-0020 | HOM Feedthru |  |  |  |  | Cracked sapphire, antenna and center pin tight. | carpente | 2011-04-08 |  |  |  | HKPARK | 2011-04-08 |  | Scrap | [14 empty] | MPF-0057 | carpente | 2011-04-08 |  | hkpark | 2011-04-08 | Scrap
475 | ISI-0005 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 104 | SMALL SCRATCH IN SEAL GROOVE CONNECTOR DEPTH .152-.166CKS. .175 SEE ATTACHME | ogle | 2011-04-08 |  |  |  | HKPARK | 2011-04-22 | The probe tip length is too short. Do not use unless matching HOM can is found. | Scrap | [14 empty] | ISI-0005 | ogle | 2011-04-08 | 37A7E241-DD31-5AC6-882BCE815C6D25AD |  | 2011-04-22 | Scrap
476 | ISI-0006 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 105 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | ogle | 2011-04-08 |  |  |  | HKPARK | 2011-08-18 | Leaker | Scrap | [14 empty] | ISI-0006 | ogle | 2011-04-08 | 37AB97B4-01A9-D306-F5622AF4ED38EB90 |  | 2011-08-18 | Scrap
477 | C100-RI-054 |  | 12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 64 | HOM A:LIGHT PITTING ON SURFACES INCLUDING SEAL GROOVE SEE ATTACHMENT FOR DIMEN | ogle | 2011-04-08 |  |  |  | MARHAUSE | 2011-04-09 | Polish HOM A sealing surface thoroughly. Out-of-tolerance dimensions are accpeta | Modify | [14 empty] | C100-RI-054 | ogle | 2011-04-08 | 38372DB2-CF3A-ADE3-C5549733BB96B77B |  | 2011-04-09 | Modify
478 | C100-RI-035 |  | 12070130001 | CAVITY |  | C100-CAV-INSP2 | R3 | 36 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | ogle | 2011-04-11 |  |  |  | MARHAUSE | 2011-04-11 |  | Use As Is | [14 empty] | C100-RI-035 | ogle | 2011-04-11 | 462438F3-AED5-1843-D64B77111EF16AAA |  | 2011-04-11 | Use As Is
479 | F067764-03 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R5 | 96 | O-ring groove has light scratches | georged | 2011-04-12 |  |  |  | KWILSON |  | Repair scratches. | Modify | [14 empty] | F067764-03 | georged | 2011-04-12 |  |  |  | Modify
480 | RI-PRO-006 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 111 | SCRATCHES IN SEAL GROOVE. SEAL SURFACE VISUALLY NOT FLAT. DEPTH .152-.166: CH | ogle | 2011-04-12 |  |  |  | HKPARK | 2011-05-17 | Polish the seal groove. Measure the step depth .187/.207. | Remeasure | [14 empty] | RI-PRO-006 | ogle | 2011-04-12 | 4B866232-F3F5-CF24-D9C00D2D5BB8E8BC |  | 2011-05-17 | Remeasure | Step depth .187/.207. measures .213 05/18/11 BC Polished and leak tested. 05201 | 0EC587DC-AC98-073E-ECB93DBEF08F724C
481 | RI-PRO-007 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 112 | LIGHT SCRATCHES IN SEAL GROOVE.SEAL GROOVE NOT VISUALLY FLAT. SEE ATTACHMENT | ogle | 2011-04-12 |  |  |  | HKPARK | 2011-04-22 | Returned to the vendor -6/27/11 HyeKyoung | Return To Vendor | [14 empty] | RI-PRO-007 | ogle | 2011-04-12 | 4B8BADB5-FBE1-94BD-670F259A98AB878B |  | 2011-04-22 | Return To Vendor
482 | RI-PRO-008 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 113 | .156-.166 DEPTH CHECKS:.145 | ogle | 2011-04-12 |  |  |  | CARPENTE | 2011-05-17 | Inspect .187/.207 dimension. | Use As Is | [14 empty] | RI-PRO-008 | ogle | 2011-04-12 |  |  | 2011-05-17 | Use As Is | .187/.207 dimension measures .211 BC 05/18/11
483 | RI-PRO-009 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 114 | SEAL GROOVE NOT VISUALLY FLAT. .152-.166 DEPTH CHECKS:.145 | ogle | 2011-04-12 |  |  |  | CARPENTE | 2011-05-17 | Inspect .187/.207 dimension. Leak test worse seal groove of batch with alum. se | Remeasure | [14 empty] | RI-PRO-009 | ogle | 2011-04-12 | 4B9420BF-97F2-29D4-AC706FEA0A957C8F |  | 2011-05-17 | Remeasure | Leak tested on 051911 BC | 0EC67EEB-EA81-9BE9-4B86351EB9F35634
484 | RI-PRO-011 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 115 | SCRATCHES IN SEAL GROOVE. .152-.166 DEPTH CHECKS:.150 | ogle | 2011-04-12 |  |  |  | HKPARK | 2011-05-06 | The scratch is not right under the seal. Ultimately the step depth should be mea | Use As Is | [14 empty] | RI-PRO-011 | ogle | 2011-04-12 | 4B96EF33-C1BE-58E0-B391033C62E988DA |  | 2011-05-06 | Use As Is
485 | RI-PRO-013 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 116 | SCRATCHES IN SEAL GROOVE .152-.166 DEPTH CHECKS:.147 | ogle | 2011-04-12 |  |  |  | HKPARK | 2011-05-17 | Polish seal surface. measure step depth .187/.207. | Remeasure | [14 empty] | RI-PRO-013 | ogle | 2011-04-12 | 4B992757-B22C-BF30-5D03B23F4C8778CC |  | 2011-05-17 | Remeasure | Seal groove polished, step depth measured at .212 BC 5-26-11
486 | RI-PRO-014 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 117 | SCRATCH IN SEAL GROOVE. .152-.166 DEPTH CHECKS:.147 | ogle | 2011-04-12 |  |  |  | HKPARK | 2011-05-17 | Polish seal surface. measure step depth .187/.207. | Remeasure | [14 empty] | RI-PRO-014 | ogle | 2011-04-12 | 4BA08BAD-A47A-A583-F7D3A99E1817B148 |  |  | Return To Vendor | Seal surface has been polished. Step depth measures .210" High RF loss-returned to the vendor-11/2/11 HKP Use as is.
487 | RI-PRO-015 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 118 | LIGHT SCRATCH IN SEAL GROOVE. | ogle | 2011-04-12 |  |  |  | HKPARK | 2011-05-17 | Polish seal surface. Returned to vendor-6/27/11 HyeKyoung | Return To Vendor | [14 empty] | RI-PRO-015 | ogle | 2011-04-12 | 4BA2D30A-E946-3A69-E396815F6454AE76 |  | 2011-05-17 | Return To Vendor
488 | RI-PRO-016 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 119 | SEAL GROOVE NOT VISUALLY FLAT. .152-.166 DEPTH CHECKS:.144 | ogle | 2011-04-12 |  |  |  | HKPARK | 2011-05-17 | The worst case of the seal surface will be tested with aluminum seal. Measure | Remeasure | [14 empty] | RI-PRO-016 | ogle | 2011-04-12 | 4BB84F13-C970-F575-B2E54035801B0E7B |  |  | Use As Is
489 | RI-PRO-017 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 120 | SEAL SURFACE NOT VISUALLY FLAT. | ogle | 2011-04-12 |  |  |  | CARPENTE | 2011-05-17 | Leak test worse seal groove of batch with alum. seal. | Remeasure | [14 empty] | RI-PRO-017 | ogle | 2011-04-12 | 4BBB5487-A042-7251-C4EA06AA88A82011 |  | 2011-05-17 | Remeasure | Leak tested on 051911 BC | 0EC711A3-CF4A-DF80-74ED9BEBBEE3B1CD
490 | RI-PRO-018 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 121 | SCRATCH IN SEAL GROOVE. SPOT OF UNKNOWN MATERIAL ADHERED TO SEAL GROOVE SURFACE. | ogle | 2011-04-12 |  |  |  | CARPENTE | 2011-05-17 | Inspect .187/.207 dimension. Polish seal groove 17-May-2011 15:45 | Use As Is | [14 empty] | RI-PRO-018 | ogle | 2011-04-12 | 4BBEE2A2-FB24-3057-974D4811CDCD650F |  | 2011-05-17 | Use As Is | Seal groove has been polished on 051811 BC. Step depth .187/.207 From the pi | 0EC85BF9-DA62-0C9C-8541C00CD5BC2147
491 | RI-PRO-019 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 122 | SCRATCH IN SEAL GROOVE. SEAL GROOVE NOT VISUALLY FLAT. .152-.166 DEPTH CHECKS:. | ogle | 2011-04-12 |  |  |  | CARPENTE | 2011-05-17 | Inspect .187/.207 dimension. Polish seal groove 17-May-2011 15:48 | Use As Is | [14 empty] | RI-PRO-019 | ogle | 2011-04-12 | 4BC39B8C-9694-63E4-0D22A5B5981B3FD0 |  | 2011-05-17 | Use As Is | Seal groove has been polished on 051811 BC. Step depth .187/.207 From the pin
492 | RI-PRO-020 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 123 | .152-.166 DEPTH CHECKS:.147 SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | ogle | 2011-04-12 |  |  |  | HKPARK |  | Inspect .187/.207 dimension. | Remeasure | [14 empty] | RI-PRO-020 | ogle | 2011-04-12 | 4BFCA8F1-FA7B-AEF2-81893AC5DB3CDCF6 |  |  | Use As Is | Step depth .187/.207 From the pin end to the connector step:.Checks .210 5/17/1
493 | RI-PRO-021 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 124 | SCRATCH IN SEAL GROOVE .152-.166 CHECKS:.147 SEE ATTACHMENT FOR ADDITIONAL D | ogle | 2011-04-12 |  |  |  | CARPENTE | 2011-05-17 | Inspect .187/.207 dimension. Polish seal groove | Use As Is | [14 empty] | RI-PRO-021 | ogle | 2011-04-12 | 4C029E58-0C0F-BA5B-51ABE54C1A6FF016 |  | 2011-05-17 | Use As Is | Seal groove polished. 05/18/11 BC .187/.207 From the pin end to the connector
494 | RI-PRO-022 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 125 | SCRATCHES IN SEAL GROOVE. BRAZE SPOT IN SEAL GROOVE. | ogle | 2011-04-12 |  |  |  | HKPARK | 2011-05-17 |  | Use As Is | [14 empty] | RI-PRO-022 | ogle | 2011-04-12 | 4C0B940B-B53B-8BFE-953D126A8C7DCFBA |  | 2011-05-17 | Use As Is | This feedthru has passed a cold leak test. BCC 05/04/11
495 | RI-PRO-023 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 126 | LIGHT SCRATCH IN SEAL GROOVE. SEAL SURFACE NOT VISUALLY FLAT. | ogle | 2011-04-12 |  |  |  | HKPARK | 2011-05-17 | Polish seal groove Leak test worst of batch with alum. seal | Modify | [14 empty] | RI-PRO-023 | ogle | 2011-04-12 | 4C0EE5B8-92A9-F17A-1F8812F739B88A7D |  | 2011-05-17 | Modify | Leak tested on 051911 BC | 0EC928D7-F203-7E84-46EB19E2B31CDD3A
496 | RI-PRO-024 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 127 | SPOT OF UNKNOWN MATERIAL ADHERED IN SEAL GROOVE. CONNECTOR PIN ATTACHED BUT LOO | ogle | 2011-04-12 |  |  |  | HKPARK | 2011-06-28 |  | Return To Vendor | [14 empty] | RI-PRO-024 | ogle | 2011-04-12 | 4C110374-EA23-879C-AAE40E0E310FBEB4 |  | 2011-06-28 | Return To Vendor
497 | C100-RI-036 |  | 12070130001 | CAVITY |  | C100-CAV-INSP2 | R3 | 37 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | ogle | 2011-04-12 |  |  |  | MARHAUSE | 2011-04-14 |  | Use As Is | [14 empty] | C100-RI-036 | ogle | 2011-04-12 | 4C84F6B0-9AF2-B1FC-87E7309C3135BB41 |  | 2011-04-14 | Use As Is
498 | C100-RI-055 |  | 12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 65 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | ogle | 2011-04-12 |  |  |  | MARHAUSE | 2011-04-14 |  | Use As Is | [14 empty] | C100-RI-055 | ogle | 2011-04-12 | 4CDDB042-E022-1306-CC2FD87E4EF1A44E |  | 2011-04-14 | Use As Is
499 | ISI-0006 |  | CRM1207014-0020 | HOM Feedthru |  | C100-CAV-INSP-HMFT | R3 | 105 | Leak found during receiving inspection. Sapphire appears to be intact, antenna i | carpente | 2011-04-13 |  |  |  | HKPARK | 2011-04-13 | The loose pin is not a leak pass. It's possible it has hair line crack at the sa | Scrap | [14 empty] | ISI-0006 | carpente | 2011-04-13 |  |  | 2011-04-13 | Scrap
500 | RI-PRO-05 |  | CRM12070140020 | HOM Feedthru |  | C100-CAV-INSP-HMFT | R3 | 102 | This feedthru has been recycled by polishing the antenna. It has failed a prelim | carpente | 2011-04-13 |  |  |  | HKPARK | 2011-04-13 |  | Scrap | [14 empty] | RI-PRO-05 | carpente | 2011-04-13 |  |  | 2011-04-13 | Scrap
501 | C100-RI-024 |  | 12070130001 | CAVITY |  | C100-CAV-INSP2 | R3 | 38 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | ogle | 2011-04-13 |  |  |  | ABURRILL | 2011-08-12 | dimensional deviation acceptable | Use As Is | [14 empty] | C100-RI-024 | ogle | 2011-04-13 | 51804D66-C800-53CD-0EE0236031C65520 |  | 2011-08-12 | Use As Is
502 | C100-RI-037 |  | 12070130001 | CAVITY |  | C100-CAV-INSP2 | R3 | 39 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | ogle | 2011-04-15 |  |  |  | ABURRILL | 2011-08-12 | dimensional deviation acceptable | Use As Is | [14 empty] | C100-RI-037 | ogle | 2011-04-15 | 5AF9AD91-DE14-1203-969FE19028EAC426 |  | 2011-08-12 | Use As Is
503 | F067764-17 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-RWRK-WGD | R3 | 25 | Flatness out of tolerance, measured 0.005" | georged | 2011-04-18 |  |  |  | KWILSON |  | Flatness ok. | Use As Is | [14 empty] | F067764-17 | georged | 2011-04-18 |  | carpente |  | Use As Is
504 | F063134-06 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R5 | 102 | Scratches on small flange. Damage inside o-ring groove. Nicks on knife. | georged | 2011-04-18 |  |  |  | KWILSON |  | Please repair scratches and nicks. | Modify | [14 empty] | F063134-06 | georged | 2011-04-18 |  |  |  | Modify
505 | C100-RI-056 |  | 12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 66 | FPC FLANGE HAS SCRATCHES WITHIN THE SEAL AREA. SEE ATTACHMENT FOR DIMENSIONAL N | OGLE | 2011-04-18 | 17124ACB-09F2-2F50-A3973E4FBA8DD376 |  |  | ABURRILL | 2011-10-04 | lap FPC flange, dimensional will be made up in welding process | Use As Is | [14 empty] | C100-RI-056 | ogle | 2011-04-18 | 6B7F21E1-E280-67E6-86039085E812B86C
506 | F067764-07 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R5 | 103 | Scratches on small flange. Surface finish out of tolerance (28 micro inches) | georged | 2011-04-19 |  |  |  | HOGAN |  | Repair the scratches using the C100-CAV-WRK-WGDR3 traveler. | Modify | [14 empty] | F067764-07 | georged | 2011-04-19 |  |  |  | Modify
507 | F356224-03 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R5 | 101 | Scratches on small flange. Surface finish out of tolerance, meas. 24 micro inche | georged | 2011-04-19 |  |  |  | HOGAN | 2011-04-21 | Repair the scratches and damage to knife edge using the C100-CAV-WRK-WGDR3 trave | Modify | [14 empty] | F356224-03 | georged | 2011-04-19 |  |  | 2011-04-21 | Modify
508 | C100-RI-057 |  | 12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 67 | SCRATCHES ON FPC FLANGE WITHIN SEAL AREA. SEE ATTACHMENT FOR DIMENSIONAL NON-CO | OGLE | 2011-04-19 | 171AB21C-FD1B-4852-56722BB54DA7F27F |  |  | ABURRILL | 2011-10-04 | dimensional deviation will be made up in assembly | Use As Is | [14 empty] | C100-RI-057 | ogle | 2011-04-19 | 704308DC-EA30-C7AA-10283228DD80FFD6
509 | C100-RI-038 |  | 12070130001 | CAVITY |  | C100-CAV-INSP2 | R3 | 40 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | ogle | 2011-04-20 |  |  |  | ABURRILL | 2011-09-19 | dimensional deviation will be made up in next assembly step | Use As Is | [14 empty] | C100-RI-038 | ogle | 2011-04-20 | 74A9DFB2-E97D-DD66-ED254D61D021F320 |  | 2011-09-19 | Use As Is
510 | OML-0077 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 104 | SMALL SPOT OF COPPER/BRAZE MATERIAL ON CERAMIC SURFACE. | ogle | 2011-04-20 |  |  |  | HKPARK | 2011-04-22 | The metal spot has to be removed. GentlyScrape with a glass or contact Scott Wi | Modify | [14 empty] | OML-0077 | ogle | 2011-04-20 | 7567013A-BF3A-F53B-1FD062731A4AC1E6 |  |  | Modify | Passed RF test. use as is.
511 | OML-0080 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 107 | SMALL SPOT OF COPPER/BRAZED MATERIAL ON CERAMIC SURFACE. SEE ATTACHMENT FOR DIM | ogle | 2011-04-20 |  |  |  | HKPARK | 2011-04-22 | The ceramic thickness is OK but the metal spot has to be removed. GentlyScrape | Modify | [14 empty] | OML-0080 | ogle | 2011-04-20 | 7568B0BF-C5D6-B63B-D8D1817DFAC3A470 |  |  | Modify | Passed RF test, use as is.
512 | F067764-19 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R6 | 117 | Nick on knife edge. Scratches on small flange. | georged | 2011-04-21 |  |  |  | KWILSON |  | Please repair nick and scratches. | Modify | [14 empty] | F067764-19 | georged | 2011-04-21 |  |  |  | Modify
513 | F356223_03 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R6 | 109 | Light scratches on small flange. Minor stains and mark inside of waveguide. FPC | carpente | 2011-04-22 |  |  |  | KWILSON |  | Please repair scratches and see if flatness can be improved at the same time. S | Modify | [14 empty] | F356223_03 | carpente | 2011-04-22 |  |  |  | Modify
514 | F067763_10 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R6 | 112 | Some dark oxidation spots inside wg. Burr on base of knife edge, no impact to se | carpente | 2011-04-22 |  |  |  | KWILSON |  | Please repair scratches and tap holes as needed. | Modify | [14 empty] | F067763_10 | carpente | 2011-04-22 |  |  |  | Modify
515 | F356224_01 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R6 | 110 | Small black oxidation spots in wg. Several nicks on knife edge. Light scratches | carpente | 2011-04-22 |  |  |  | KWILSON |  | Please tap holes as needed. Please try to repair scratches on flange and nicks | Modify | [14 empty] | F356224_01 | carpente | 2011-04-22 |  |  |  | Modify
516 | OML-0082 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 109 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | ogle | 2011-04-22 |  |  |  | HKPARK | 2011-05-13 | The level of ceramic thickness deviation OK. | Use As Is | [14 empty] | OML-0082 | ogle | 2011-04-22 | 8050E0D5-B97E-4E6B-3D5CE000911C2E4C |  | 2011-05-13 | Use As Is
517 | OML-0088 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 115 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | ogle | 2011-04-22 |  |  |  | HKPARK | 2011-05-13 | The level of ceramic thickness deviation OK. | Use As Is | [14 empty] | OML-0088 | ogle | 2011-04-22 | 80515A2A-A699-1598-D8A20903426A7FDB |  | 2011-05-13 | Use As Is
518 | OML-0090 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 117 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | ogle | 2011-04-22 |  |  |  | HKPARK | 2011-05-13 | The level of ceramic thickness deviation OK. | Use As Is | [14 empty] | OML-0090 | ogle | 2011-04-22 | 80551BCF-E277-1909-2B9967F51A65D9F5 |  | 2011-05-13 | Use As Is
519 | C100-RI-041 |  | 12070130001 | CAVITY |  | C100-CAV-INSP2 | R3 | 42 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | ogle | 2011-04-26 |  |  |  | ABURRILL | 2011-09-19 | dimensional deviation ok to use | Use As Is | [14 empty] | C100-RI-041 | ogle | 2011-04-26 | 93DEA0CA-A75B-C063-DA27CB376CF9E049 |  | 2011-09-19 | Use As Is
520 | F067764_14 |  | CRM1207070-0000 |  |  | C100-CAV-INSP-WGD | R6 | 115 | Scratches on the small flange. | carpente | 2011-04-28 |  |  |  | KWILSON |  | Please polish scratches. | Modify | [14 empty] | F067764_14 | carpente | 2011-04-28 |  |  |  | Modify
521 | FLG0037-074 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 65 | NICK ON KNIFE EDGE. SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | ogle | 2011-04-28 |  |  |  | HKPARK | 2011-04-29 | Tag and store them. It'll be sent back to the vendor for replacement. | Scrap | [14 empty] | FLG0037-074 | ogle | 2011-04-28 | 9E9F8558-E145-F1AF-B4AA21272F88309F |  | 2011-04-29 | Scrap
522 | FLG0037-077 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 68 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | ogle | 2011-04-28 |  |  |  | HKPARK | 2011-04-29 |  | Use As Is | [14 empty] | FLG0037-077 | ogle | 2011-04-28 | 9EA1C5AD-04CF-DC24-4FFB84F947700864 |  | 2011-04-29 | Use As Is
523 | FLG0037-080 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 71 | NICK ON KNIFE EDGE. | ogle | 2011-04-28 |  |  |  | HKPARK | 2011-04-29 | Tag and store. It'll be sent back to the vendor for replacement. | Scrap | [14 empty] | FLG0037-080 | ogle | 2011-04-28 | 9EA654B6-0ED9-8C5E-442128F4EB26E869 |  | 2011-04-29 | Scrap
524 | FLG0037-084 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 75 | NICK ON KNIFE EDGE | ogle | 2011-04-28 |  |  |  | HKPARK | 2011-04-29 | Tag and store. it'll be sent back to the vendor for replacement. | Scrap | [14 empty] | FLG0037-084 | ogle | 2011-04-28 | 9EAB0600-DAC5-F2EF-489D3CE37C1E1CB1 |  | 2011-04-29 | Scrap
525 | FLG0037-085 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 76 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | ogle | 2011-04-28 |  |  |  | HKPARK | 2011-04-29 | Before release check with the mating flange. | Use As Is | [14 empty] | FLG0037-085 | ogle | 2011-04-28 | 9EAD4437-C0AA-875E-F4FA7A43D3F3418C |  | 2011-04-29 | Use As Is
526 | FLG0037-087 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 78 | NICK ON KNIFE EDGE | ogle | 2011-04-28 |  |  |  | HKPARK | 2011-04-29 | Tag and store. It'll be sent back to the vendor for replacement. | Scrap | [14 empty] | FLG0037-087 | ogle | 2011-04-28 | 9EAEAB53-9548-4919-4190BC0D1CC9D689 |  | 2011-04-29 | Scrap
527 | C100-RI-032 |  | 12070130001 | CAVITY |  | C100-CAV-INSP2 | R3 | 43 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | ogle | 2011-04-28 |  |  |  | ABURRILL | 2011-09-19 | deviation will compensated for in next assembly step. | Use As Is | [14 empty] | C100-RI-032 | ogle | 2011-04-28 | 9EEE771A-E3CD-59E9-4A2D1D1FE94A0BA0 |  | 2011-09-19 | Use As Is
528 | RI-PRO-027 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 129 | LIGHT SCRATCH IN SEAL GROOVE. DIMENSIONAL: .152/.166 CKS..147 .187/.207 CKS . | ogle | 2011-05-03 |  |  |  | HKPARK | 2011-06-07 | Seal groove polished. The mentioned dimensions not affecting assembly. The inser | Use As Is | [14 empty] | RI-PRO-027 | ogle | 2011-05-03 | B755ADAE-BC12-6063-6512B5020E80A5C9 |  | 2011-06-07 | Use As Is
529 | RI-PRO-028 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 130 | CONTAMINATION INSIDE NEAR SAPPHIRE;CONNECTOR AREA LIGHT SCRATCH IN SEAL GROOVE. | ogle | 2011-05-03 |  |  |  | HKPARK | 2011-06-09 | Seal surface polished. Insertion loss below 8 db. | Use As Is | [14 empty] | RI-PRO-028 | ogle | 2011-05-03 | B75AA0C6-B996-D8AF-E3043D079905E2F8 |  | 2011-06-09 | Use As Is
530 | RI-PRO-029 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 131 | CONTAMINATION INSIDE NEAR SAPPHIRE;CONNECTOR AREA DIMENSIONAL: .152/.166 CKS.. | ogle | 2011-05-03 |  |  |  | HKPARK | 2011-06-28 |  | Return To Vendor | [14 empty] | RI-PRO-029 | ogle | 2011-05-03 |  |  | 2011-06-28 | Return To Vendor
531 | RI-PRO-030 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 132 | LIGHT SCRATCH IN SEAL GROOVE. SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | ogle | 2011-05-03 |  |  |  | HKPARK | 2011-06-07 | Seal groove polished. 1.464 +/-.010 OK. Angle deviation small enough. Insertion | Use As Is | [14 empty] | RI-PRO-030 | ogle | 2011-05-03 | B761D6CD-E44C-CAF2-88D15B92CC02CD7B |  | 2011-06-07 | Use As Is
532 | RI-PRO-031 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 133 | CONTAMINATION INSIDE NEAR SAPPHIRE;CONNECTOR AREA DIMENSIONAL: .152/.166 CKS.. | ogle | 2011-05-03 |  |  |  | HKPARK | 2011-06-07 | Clean the contamination. The mentioned dimension not affecting assembly. Inserti | Use As Is | [14 empty] | RI-PRO-031 | ogle | 2011-05-03 |  |  | 2011-06-07 | Use As Is
533 | RI-PRO-032 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 134 | CONTAMINATION INSIDE NEAR SAPPHIRE;CONNECTOR AREA DIMENSIONAL: .152/.166 CKS.. | ogle | 2011-05-03 |  |  |  | HKPARK | 2011-06-28 |  | Return To Vendor | [14 empty] | RI-PRO-032 | ogle | 2011-05-03 |  |  | 2011-06-28 | Return To Vendor
534 | RI-PRO-033 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 135 | LIGHT SCRATCH IN SEAL GROOVE. | ogle | 2011-05-03 |  |  |  | HKPARK | 2011-06-07 | Seal groove polished. Insertion loss below 5db. | Use As Is | [14 empty] | RI-PRO-033 | ogle | 2011-05-03 | B769554B-B906-35AE-42C270EF2434C0E4 |  | 2011-06-07 | Use As Is
535 | RI-PRO-034 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 136 | BRAZE ONLY HALF AROUND BASE OF PROBE CONTAMINATION INSIDE NEAR SAPPHIRE;CONNECT | ogle | 2011-05-03 |  |  |  | HKPARK | 2011-06-28 | Confirm the braze strength with torque test. Locate the seal line with the al. | Remeasure | [14 empty] | RI-PRO-034 | ogle | 2011-05-03 | B76D9A54-DB1E-918F-A38F832AA88CE393 |  | 2011-06-28 | Remeasure | TORQUE TEST ON CONNECTOR GOOD. BRAZE SPOT OUTSIDE SEAL AREA AS DETERMINED BY SE
536 | RI-PRO-036 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 137 | CONTAMINATION INSIDE NEAR SAPPHIRE;CONNECTOR AREA DAMAGE SPOT IN SEAL GROOVE. | ogle | 2011-05-03 |  |  |  | HKPARK | 2011-06-09 | Clean contamination, seal surface polished. RF loss below 8 dB. | Use As Is | [14 empty] | RI-PRO-036 | ogle | 2011-05-03 | B7705AE1-DD30-9615-927EA69EE7571C26 |  | 2011-06-09 | Use As Is
537 | RI-PRO-037 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 138 | CONTAMINATION INSIDE NEAR SAPPHIRE;CONNECTOR AREA DIMENSIONAL: .152/.166 CKS.. | ogle | 2011-05-03 |  |  |  | HKPARK | 2011-06-09 | clean contamination, seal surface polished. RF loss below 8 dB. | Use As Is | [14 empty] | RI-PRO-037 | ogle | 2011-05-03 |  |  | 2011-06-09 | Use As Is
538 | RI-PRO-038 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 139 | DIMENSIONAL: .152/.166 CKS..146 .187/.207 CKS.211 | OGLE | 2011-05-03 |  |  | need disposition to finish this one. BC | HKPARK |  | If the connector can be safely engaged, use as is. | [15 empty] | RI-PRO-038 | OGLE | 2011-05-03 |  |  |  | Use As Is
539 | RI-PRO-039 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 140 | CONTAMINATION INSIDE NEAR SAPPHIRE;CONNECTOR AREA LIGHT SCRATCH IN SEAL GROOVE. | ogle | 2011-05-03 |  |  |  | HKPARK | 2011-06-07 | Clean the contamination. Seal groove polished. The mentioned dimensions not affe | Use As Is | [14 empty] | RI-PRO-039 | ogle | 2011-05-03 | B7768928-0CD0-6FE4-9DAF8D7108499C6F |  | 2011-06-07 | Use As Is
540 | RI-PRO-040 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 141 | CONTAMINATION INSIDE NEAR SAPPHIRE;CONNECTOR AREA LIGHT SCRATCH IN SEAL GROOVE | ogle | 2011-05-03 |  |  |  | HKPARK | 2011-06-09 | clean contamination, seal surface polished. Angle deviation small. RF loss below | Use As Is | [14 empty] | RI-PRO-040 | ogle | 2011-05-03 | B790E289-FF26-053E-637CD91E1EA29B5A |  | 2011-06-09 | Use As Is
541 | RI-PRO-042 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 142 | CONTAMINATION INSIDE NEAR SAPPHIRE;CONNECTOR AREA SEE ATTACHMENT FOR DIMENSIONA | ogle | 2011-05-03 |  |  |  | HKPARK | 2011-06-28 |  | Return To Vendor | [14 empty] | RI-PRO-042 | ogle | 2011-05-03 | B7959CD5-C7FE-393A-1CD987C14A54564F |  | 2011-06-28 | Return To Vendor
542 | RI-PRO-043 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 143 | CONTAMINATION INSIDE NEAR SAPPHIRE;CONNECTOR AREA LIGHT SCRATCH IN SEAL GROOVE | ogle | 2011-05-03 |  |  |  | HKPARK | 2011-06-07 | Clean the contamination. Seal groove polished. The mentioned dimensions not affe | Use As Is | [14 empty] | RI-PRO-043 | ogle | 2011-05-03 | B798C166-0FB7-D2D1-5BDEC7D24A1BD07D |  | 2011-06-07 | Use As Is
543 | RI-PRO-047 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 144 | CONTAMINATION INSIDE NEAR SAPPHIRE;CONNECTOR AREA LIGHT SCRATCH IN SEAL GROOVE | ogle | 2011-05-03 |  |  |  | HKPARK | 2011-06-28 |  | Return To Vendor | [14 empty] | RI-PRO-047 | ogle | 2011-05-03 | B7A4A690-EE2E-82E4-304A7887B9E11B95 |  | 2011-06-28 | Return To Vendor
544 | OML-0092 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 119 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | ogle | 2011-05-04 |  |  |  | HKPARK | 2011-05-05 | Use as is if it fits in the flange pocket. | Use As Is | [14 empty] | OML-0092 | ogle | 2011-05-04 | BCA504AA-EA22-0E88-5CCA62796D50F08D |  | 2011-05-05 | Use As Is
545 | OML-0093 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 120 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | ogle | 2011-05-04 |  |  |  | HKPARK | 2011-05-05 | Use as is if it fits in the flange pocket. | Use As Is | [14 empty] | OML-0093 | ogle | 2011-05-04 | BCC72772-AFEB-8A41-21C11F7EF604A686 |  | 2011-05-05 | Use As Is
546 | OML-0094 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 121 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | ogle | 2011-05-04 |  |  |  | HKPARK | 2011-05-05 | The thickness level is OK. | Use As Is | [14 empty] | OML-0094 | ogle | 2011-05-04 | BCCDBF90-E127-F5C2-90A4F2D1FB0542B6 |  | 2011-05-05 | Use As Is
547 | OML-0098 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 125 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | ogle | 2011-05-04 |  |  |  | HKPARK | 2011-05-05 | Use as is if it fits in the flange pocket. | Use As Is | [14 empty] | OML-0098 | ogle | 2011-05-04 | BD017969-BDB7-01C1-A20EA41900777CD9 |  | 2011-05-05 | Use As Is
548 | OML-0099 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 126 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | ogle | 2011-05-04 |  |  |  | HKPARK | 2011-05-05 | Use as is if it fits in the flange pocket. | Use As Is | [14 empty] | OML-0099 | ogle | 2011-05-04 | BD1BE241-B453-1222-20602C79B2113FC7 |  | 2011-05-05 | Use As Is
549 | C100-RI-058 |  | 12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 68 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | ogle | 2011-05-04 | 17097FB7-EFA2-4D9D-A4FBB2A8E9320370 |  |  | ABURRILL | 2011-10-04 |  | Use As Is | [14 empty] | C100-RI-058 | ogle | 2011-05-04 | BD5DFFD2-B0EC-0043-7088934202BFF2C3
550 | C100-RI-059 |  | 12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 69 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-05-04 | 1724B583-B7F0-4D41-241E0A586462C5CD |  |  | ABURRILL | 2011-10-04 | deviation will be made up in cryomodule assembly | Use As Is | [14 empty] | C100-RI-059 | ogle | 2011-05-04 | BDBCFF86-0017-AAD1-A0E056BF08259EFC
551 | F892522-09 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R6 | 120 | Ding on Cu plating, located on inside corner of flg face. Nick on knife edge. Su | GEORGED | 2011-05-05 |  |  |  | KWILSON |  | Please repair knife edge. Ding ok. | Modify | [14 empty] | F892522-09 | GEORGED | 2011-05-05 |  |  |  | Modify
552 | F890195-05 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R6 | 121 | Scratches on small flange. Copper missing on large flange face. | georged | 2011-05-05 |  |  |  | KWILSON |  | Repair scratches on small flange. Copper ok. | Modify | [14 empty] | F890195-05 | georged | 2011-05-05 |  |  |  | Modify
553 | FLG0037-090 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 81 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | ogle | 2011-05-06 |  |  |  | HKPARK | 2011-05-09 | The adapter fits and the dimension doesn't affect welding. | Use As Is | [14 empty] | FLG0037-090 | ogle | 2011-05-06 | C7884449-D38F-796D-D2A8F8D914576349 |  | 2011-05-09 | Use As Is
554 | FLG0037-094 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 85 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | ogle | 2011-05-06 |  |  |  | HKPARK | 2011-05-09 | The adapter fits and the dimension doesn't affect welding. | Use As Is | [14 empty] | FLG0037-094 | ogle | 2011-05-06 | C7AAC141-0972-072B-7E74506CB5DBEC19 |  | 2011-05-09 | Use As Is
555 | FLG0037-095 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 86 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | ogle | 2011-05-06 |  |  |  | HKPARK | 2011-05-09 | The bolt hole location deviation value is OK but make sure the fitting with the | Use As Is | [14 empty] | FLG0037-095 | ogle | 2011-05-06 | C7B340D1-F786-F744-CC25A86C773DBDA8 |  | 2011-05-09 | Use As Is
556 | FLG0037-096 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 87 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | ogle | 2011-05-06 |  |  |  | HKPARK | 2011-05-09 | The adapter fits and the dimension doesn't affect welding. The bold hole deviati | Use As Is | [14 empty] | FLG0037-096 | ogle | 2011-05-06 | C7C84AE2-A162-FE71-0D85C7A75F9EB21E |  | 2011-05-09 | Use As Is
557 | FLG0037-097 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 88 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | ogle | 2011-05-06 |  |  |  | HKPARK | 2011-05-09 | The adapter fits and the dimension doesn't affect welding. The bold hole deviati | Use As Is | [14 empty] | FLG0037-097 | ogle | 2011-05-06 | C7D16966-EB16-5912-8BB5664992004758 |  | 2011-05-09 | Use As Is
558 | FLG0037-099 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 90 | DAMAGE TO KNIFE EDGE. SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | ogle | 2011-05-06 |  |  |  | HKPARK | 2011-05-09 |  | Return To Vendor | [14 empty] | FLG0037-099 | ogle | 2011-05-06 | C818D52C-AF3B-1410-F3E6DA0F341801D0 |  | 2011-05-09 | Return To Vendor
559 | FLG0037-100 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 91 | DAMAGED KNIFE EDGE | ogle | 2011-05-06 |  |  |  | HKPARK | 2011-05-09 |  | Return To Vendor | [14 empty] | FLG0037-100 | ogle | 2011-05-06 | C81F4CA8-BC2B-D108-FF74CD1C9CBF537E |  | 2011-05-09 | Return To Vendor
560 | FLG0037-101 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 92 | DAMAGED KNIFE EDGE. SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | ogle | 2011-05-06 |  |  |  | HKPARK | 2011-05-09 |  | Return To Vendor | [14 empty] | FLG0037-101 | ogle | 2011-05-06 | C828FA6B-E6AB-72C4-402C0146DE32B736 |  | 2011-05-09 | Return To Vendor
561 | F356224-02 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R6 | 118 | Threads bad on one outer pattern hole. Scratches on small flange. | georged | 2011-05-09 |  |  |  | KWILSON |  | Please repair scratches and threads. Thanks. | Modify | [14 empty] | F356224-02 | georged | 2011-05-09 |  |  |  | Modify
562 | F890196-01 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R6 | 125 | Nick on knife edge. Scratches on small flange. Surface finish out of tolerance, | georged | 2011-05-10 |  |  |  | KWILSON |  | Please repair scratches on small flange and nicks on knife edge. | Modify | [14 empty] | F890196-01 | georged | 2011-05-10 |  |  |  | Modify
563 | ISI-0012 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 145 | Sapphire broken, probe tip very loose. HEAVY NICKS ON PROBE (SEE TRAVELER PHOTO | OGLE | 2011-05-10 |  |  |  | HKPARK | [17 empty] | ISI-0012 | OGLE | 2011-05-10 | DC150A67-C01D-0D74-9E2E98E9595B71FA |  |  | Return To Vendor
564 | ISI-0013 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 146 | HEAVY NICKS ON PROBE (SEE TRAVELER PHOTO) LIGHT SCRATCH IN SEAL GROOVE (SEE T | ogle | 2011-05-10 |  |  |  | HKPARK | 2011-06-10 | SS connector found loose at torque test. Returned to the vendor. | Return To Vendor | [14 empty] | ISI-0013 | ogle | 2011-05-10 | DC17789B-91E5-4B83-8E8C9B7FF5B93509 |  | 2011-06-10 | Return To Vendor
565 | ISI-0014 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 147 | HEAVY NICKS ON PROBE (SEE TRAVELER PHOTO) DEPTH:.152/.166 CKS..168 | ogle | 2011-05-10 |  |  |  | HKPARK | 2011-06-28 | The heavy nicks are in low field area per Reece. Vendor agreed to pay for the re | Modify | [14 empty] | ISI-0014 | ogle | 2011-05-10 | DC1C1C15-F468-9984-5D2829A09FFB3AD2 |  | 2011-06-28 | Modify | BCP COMPLETE RELEASE FOR PRODUCTION.
566 | ISI-0015 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 148 | HEAVY NICKS ON PROBE (SEE TRAVELER PHOTO) | ogle | 2011-05-10 |  |  |  | HKPARK | 2011-06-28 | The heavy nicks are in low field area per Reece. Vendor agreed to pay for the re | Modify | [14 empty] | ISI-0015 | ogle | 2011-05-10 | DC1E1F92-A659-A3A3-F7BA1AE57B9A00C6 |  | 2011-06-28 | Modify | BCP COMPLETE RELEASE FOR PRODUCTION.
567 | ISI-0016 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 149 | HEAVY NICKS ON PROBE (SEE TRAVELER PHOTO) DEPTH:.152/.166 CKS..168 | ogle | 2011-05-10 |  |  |  | HKPARK | 2011-06-28 | The heavy nicks are in low field area per Reece. Vendor agreed to pay for the re | Modify | [14 empty] | ISI-0016 | ogle | 2011-05-10 | DC20B2D6-9A5B-B064-997D8F37CE5E49B5 |  | 2011-06-28 | Modify | BCP COMPLETE RELEASE FOR PRODUCTION.
568 | ISI-0017 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 150 | HEAVY NICKS ON PROBE (SEE TRAVELER PHOTO) SEE ATTACHMENT FOR DIMENSIONAL-NON C | ogle | 2011-05-10 |  |  |  | HKPARK | 2011-06-28 | The heavy nicks are in low field area per Reece. Vendor agreed to pay for the re | Modify | [14 empty] | ISI-0017 | ogle | 2011-05-10 | DC23F797-D45F-0C2B-F292113F5B4B52B6 |  | 2011-06-28 | Modify | BCP COMPLETE RELEASE FOR PRODUCTION.
569 | ISI-0018 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 151 | HEAVY NICKS ON PROBE (SEE TRAVELER PHOTO) | ogle | 2011-05-10 |  |  |  | HKPARK | 2011-06-28 | The heavy nicks are in low field area per Reece. Vendor agreed to pay for the re | Modify | [14 empty] | ISI-0018 | ogle | 2011-05-10 |  |  | 2011-06-28 | Modify | BCP COMPLETE RELEASE FOR PRODUCTION.
570 | ISI-0019 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 152 | SCRATCH IN SEAL GROOVE HEAVY NICKS ON PROBE (SEE TRAVELER FOR PHOTOS) | ogle | 2011-05-10 |  |  |  | HKPARK | 2011-06-28 | The heavy nicks are in low field area per Reece. Vendor agreed to pay for the re | Modify | [14 empty] | ISI-0019 | ogle | 2011-05-10 |  |  | 2011-06-28 | Modify | BCP COMPLETE RELEASE FOR PRODUCTION.
571 | ISI-0020 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 153 | HEAVY NICKS ON PROBE (SEE TRAVELER FOR PHOTOS) SEE ATTACHMENT FOR DIMENSIONAL | ogle | 2011-05-10 |  |  |  | HKPARK | 2011-06-28 |  | Return To Vendor | [14 empty] | ISI-0020 | ogle | 2011-05-10 | DCADB0FC-F08E-2552-7415F0DA64573F97 |  | 2011-06-28 | Return To Vendor
572 | ISI-0021 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 154 | HEAVY NICKS ON PROBE (SEE TRAVELER FOR PHOTO) | OGLE | 2011-05-10 |  |  |  | HKPARK | 2011-06-28 | The heavy nicks are in low field area per Reece. Vendor agreed to pay for the re | Modify | [14 empty] | ISI-0021 | OGLE | 2011-05-10 |  |  | 2011-06-28 | Modify | BCP COMPLETE RELEASE FOR PRODUCTION.
573 | ISI-0022 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 155 | HEAVY NICKS ON PROBE (SEE TRAVELER FOR PHOTO) DEPTH:.152/.166 CKS..170 SEE AT | ogle | 2011-05-10 |  |  |  | HKPARK | 2011-06-28 | The heavy nicks are in low field area per Reece. Vendor agreed to pay for the re | Modify | [14 empty] | ISI-0022 | ogle | 2011-05-10 | DCB6FF68-F2FC-1880-BFDD201296AB2AD7 |  | 2011-06-28 | Modify | BCP COMPLETE RELEASE FOR PRODUCTION.
574 | ISI-0023 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 156 | HEAVY NICKS ON PROBE (SEE TRAVELER FOR PHOTO) | ogle | 2011-05-10 |  |  |  | HKPARK | 2011-06-28 | The heavy nicks are in low field area per Reece. Vendor agreed to pay for the re | Modify | [14 empty] | ISI-0023 | ogle | 2011-05-10 |  |  | 2011-06-28 | Modify | BCP COMPLETE RELEASE FOR PRODUCTION.
575 | ISI-0024 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 157 | HEAVY NICKS ON PROBE (SEE TRAVELER FOR PHOTO) | OGLE | 2011-05-10 |  |  | BCP COMPLETE RELEASE FOR PRODUCTION. | HKPARK | 2011-06-28 | The heavy nicks are in low field area per Reece. Vendor agreed to pay for the re | Modify | [14 empty] | ISI-0024 | OGLE | 2011-05-10 |  |  | 2011-06-28 | Modify | BCP COMPLETE RELEASE FOR PRODUCTION.
576 | ISI-0025 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 158 | HEAVY NICKS ON PROBE (SEE TRAVELER FOR PHOTO) | ogle | 2011-05-10 |  |  |  | HKPARK | 2011-06-28 | The heavy nicks are in low field area per Reece. Vendor agreed to pay for the re | Modify | [14 empty] | ISI-0025 | ogle | 2011-05-10 |  |  | 2011-06-28 | Modify | BCP COMPLETE RELEASE FOR PRODUCTION.
577 | F067764-18 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R6 | 116 | Nick on knife edge. Scratches on small flange. Stains inside waveguide. Machin | GEORGED | 2011-05-12 |  |  |  | KWILSON |  |  | Return To Vendor | [14 empty] | F067764-18 | GEORGED | 2011-05-12 |  |  |  | Return To Vendor
578 | C100-RI-039 |  | 12070130001 | CAVITY |  | C100-CAV-INSP2 | R3 | 45 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-05-12 | 172F1BCC-AE08-DC4E-6DD12301AEFE75A6 |  | 2mils out of spec should be OK. Thermal contraction and the flexible waveguide t | HAIPENG | 2011-10-04 | LOC1 dimension needs to be QC by RF bench measurement. | Use As Is | [14 empty] | C100-RI-039 | ogle | 2011-05-12 | E6455CCC-CA6B-F767-A5C80DDC17B86004 |  |  | Remeasure
579 | OML-0100 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 127 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-05-16 |  |  |  | HKPARK | 2011-05-17 | The drawing has been changed to have radii .688+/-.005 to avoid interference. | Use As Is | [14 empty] | OML-0100 | OGLE | 2011-05-16 | FBCDB3C7-C348-EEFB-FEFB408981AB2BC2 |  | 2011-05-17 | Use As Is
580 | OML-0101 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 128 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-05-16 |  |  |  | HKPARK | 2011-05-17 | The drawing has been changed to have radii .688+/-.005 to avoid interference. | Use As Is | [14 empty] | OML-0101 | OGLE | 2011-05-16 | FBCEE6B6-D5CF-A4BD-1DBD3F287EC8212E |  | 2011-05-17 | Use As Is
581 | OML-0102 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 129 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-05-16 |  |  |  | HKPARK | 2011-05-17 | The drawing has been changed to have radii .688+/-.005 to avoid interference. | Use As Is | [14 empty] | OML-0102 | OGLE | 2011-05-16 | FBD20E79-C800-6A68-71F86AFFD4D8F5F3 |  | 2011-05-17 | Use As Is
582 | OML-0103 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 130 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-05-16 |  |  |  | HKPARK | 2011-05-17 | The drawing has been changed to have radii .688+/-.005 to avoid interference. | Use As Is | [14 empty] | OML-0103 | OGLE | 2011-05-16 | FBD2FE31-A849-E0BE-D82038047C3D99A0 |  | 2011-05-17 | Use As Is
583 | OML-0104 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 131 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-05-16 |  |  |  | HKPARK | 2011-05-17 | The drawing has been changed to have radii .688+/-.005 to avoid interference. | Use As Is | [14 empty] | OML-0104 | OGLE | 2011-05-16 | FBD3EDA6-A45B-12D9-28E65CCDC2B3B10C |  | 2011-05-17 | Use As Is
584 | OML-0105 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 132 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-05-16 |  |  |  | HKPARK | 2011-05-17 | The drawing has been changed to have radii .688+/-.005 to avoid interference. | Use As Is | [14 empty] | OML-0105 | OGLE | 2011-05-16 | FBD4E235-DEBD-BDD2-DC0A99B45641CF13 |  | 2011-05-17 | Use As Is
585 | OML-0107 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 134 | BRAZE SPOT ON CERAMIC SURFACE | OGLE | 2011-05-16 |  |  |  | HKPARK | 2011-05-17 | Remove braze spot by bead blast orScraping with glass edge. | Modify | [14 empty] | OML-0107 | OGLE | 2011-05-16 | FBD91358-D712-AC16-6103AC85BD985ECF |  | 2011-05-17 | Modify | RF tested and sent to the vendor for re-cleaning-HP 8/17/11
586 | OML-0111 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 138 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-05-16 |  |  |  | HKPARK | 2011-05-17 | The drawing has been changed to have radii .688+/-.005 to avoid interference. | Use As Is | [14 empty] | OML-0111 | OGLE | 2011-05-16 | FBDB09AD-EB04-CD72-AE1FB5E64BFBA4E3 |  | 2011-05-17 | Use As Is
587 | OML-0113 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 140 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-05-16 |  |  |  | HKPARK | 2011-05-17 | The drawing has been changed to have radii .688+/-.005 to avoid interference. | Use As Is | [14 empty] | OML-0113 | OGLE | 2011-05-16 | FBDCE89F-E4F8-C6B5-5C761D8DE55E6DBE |  | 2011-05-17 | Use As Is
588 | OML-0114 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 141 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-05-16 |  |  |  | HKPARK | 2011-05-17 | Ceramic thickness OK. | Use As Is | [14 empty] | OML-0114 | OGLE | 2011-05-16 | FBDE30E8-CD16-171C-18AAE338A3354AEE |  | 2011-05-17 | Use As Is
589 | OML-0115 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 142 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-05-16 |  |  |  | HKPARK | 2011-05-17 | The drawing has been changed to have radii .688+/-.005 to avoid interference. | Use As Is | [14 empty] | OML-0115 | OGLE | 2011-05-16 | FBDEFB6F-C911-4BA7-8564B2A0FB71A8A4 |  | 2011-05-17 | Use As Is
590 | ISI-0005 |  | CRM-120-7014-0020 | HOM Feedthru |  | C100-CAV-INSP-HMFT | R3 | 104 | Threaded connector end has broken-off during assembly in CR on C100-36 | CARPENTE | 2011-05-17 |  |  |  | HKPARK | 2011-05-17 | This feedthru was used for the load test. | Scrap | [14 empty] | ISI-0005 | CARPENTE | 2011-05-17 |  |  | 2011-05-17 | Scrap
591 | C100-RI-060 |  | 12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 70 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-05-17 | 173979E0-0CEA-876B-F6889621124F5134 |  |  | HAIPENG | 2011-06-17 | LOC1 dimension error needs to be QC by RF bench measurement. LOC17 dimension can | Modify | [9 empty] | ABURRILL | 2011-10-07 | Dimensional issues will be resolved in HV welding step. | Use As Is |  | C100-RI-060 | OGLE | 2011-05-17 | 0070532B-E05C-CD32-337804FDCE4FE90B |  | 2011-06-17 | Modify
592 | C100-RI-061 |  | 12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 71 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-05-17 | B27B29E5-F1F1-0A88-E9A7AC8F1316E68D |  |  | ABURRILL | 2011-10-07 | Dimensional issues will be resolved in HV welding | Use As Is | [14 empty] | C100-RI-061 | OGLE | 2011-05-17 | 00F85CAA-0A20-C7F5-DC82CFE1ED93C2F1 |  |  | Use As Is
593 | F067764-08 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R6 | 111 | Flatness out of tolerance, meas. .005" | GEORGED | 2011-05-18 |  |  |  | KWILSON |  |  | Use As Is | [14 empty] | F067764-08 | GEORGED | 2011-05-18 |  |  |  | Use As Is
594 | FLG0037-103 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 94 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-05-18 |  |  |  | HKPARK | 2011-06-08 | Deviation small enough for assembly | Use As Is | [14 empty] | FLG0037-103 | OGLE | 2011-05-18 | 063BFA4E-F6A7-0DDC-68B85D37F5D8B48A |  | 2011-06-08 | Use As Is
595 | FLG0037-106 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 97 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-05-18 |  |  |  | HKPARK | 2011-06-08 |  | Use As Is | [14 empty] | FLG0037-106 | OGLE | 2011-05-18 | 063D891B-B87D-9F70-2C68DF783341EACB |  | 2011-06-08 | Use As Is
596 | FLG0037-108 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 99 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-05-18 |  |  |  | HKPARK | 2011-06-08 | Make sure it fits with the mating flange. | Use As Is | [14 empty] | FLG0037-108 | OGLE | 2011-05-18 | 063EED57-B4F5-8157-A86175E1BD3493E1 |  | 2011-06-08 | Use As Is
597 | F067764-13 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R6 | 114 | Scratches on small flange. Flatness and surface finish out of tolerance. Flatnes | GEORGED | 2011-05-19 |  |  |  | KWILSON |  | Flatness and surface finish. Please polish scrathes. | Modify | [14 empty] | F067764-13 | GEORGED | 2011-05-19 |  |  |  | Modify
598 | F356223-05 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R6 | 108 | Scratches on small flange. On last leak check wg will only seal if torqued to at | GEORGED | 2011-05-20 |  |  |  | KWILSON |  | RTV for difficulty sealing during leak check | Return To Vendor | [14 empty] | F356223-05 | GEORGED | 2011-05-20 |  |  |  | Return To Vendor
599 | FLG0037-115 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 106 | DAMAGED KNIFE EDGE | OGLE | 2011-05-20 |  |  |  | HKPARK | 2011-06-03 |  | Return To Vendor | [14 empty] | FLG0037-115 | OGLE | 2011-05-20 | 1002CA86-A33A-567C-51A3A9947172460A |  | 2011-06-03 | Return To Vendor
600 | FLG0037-116 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 107 | DAMAGED KNIFE EDGE SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-05-20 |  |  |  | HKPARK | 2011-06-03 |  | Return To Vendor | [14 empty] | FLG0037-116 | OGLE | 2011-05-20 | 100C9E60-D767-C37B-7C32FA10E50693FB |  | 2011-06-03 | Return To Vendor
601 | FLG0037-117 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 108 | KNIFE EDGE DAMAGE TWO LOCATIONS | OGLE | 2011-05-20 |  |  |  | HKPARK | 2011-06-03 |  | Return To Vendor | [14 empty] | FLG0037-117 | OGLE | 2011-05-20 | 101127FA-9C06-0736-AE0402100D524F71 |  | 2011-06-03 | Return To Vendor
602 | F356223-01 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R6 | 106 | Scratches on small flange | GEORGED | 2011-05-23 |  |  |  | KWILSON |  | Please repair scratches | Modify | [14 empty] | F356223-01 | GEORGED | 2011-05-23 |  |  |  | Modify
603 | F067764-19 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R6 | 117 | Flatness out of tolerance, meas. .004" | GEORGED | 2011-05-23 |  |  |  | KWILSON |  |  | Use As Is | [14 empty] | F067764-19 | GEORGED | 2011-05-23 |  |  |  | Use As Is
604 | F067763-12 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R6 | 113 | Scratches on small flange | GEORGED | 2011-05-23 |  |  |  | KWILSON |  | Please repair scratches. | Modify | [14 empty] | F067763-12 | GEORGED | 2011-05-23 |  |  |  | Modify
605 | F067764-14 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R6 | 115 | Surface finish out of tolerance, meas. 19 micro inches. Flatness out of toleranc | GEORGED | 2011-05-24 |  |  |  | KWILSON |  |  | Use As Is | [14 empty] | F067764-14 | GEORGED | 2011-05-24 |  |  |  | Use As Is
606 | ISI-0040 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 159 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-05-24 |  |  |  | HKPARK | 2011-06-03 | The dimension is within 1.464 +/- .010 | Use As Is | [14 empty] | ISI-0040 | OGLE | 2011-05-24 | 2381EE74-CC87-56C7-4451F1A07962ED8C |  | 2011-06-03 | Use As Is
607 | ISI-0041 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 160 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-05-24 |  |  |  | HKPARK | 2011-06-03 | The dimension is within 1.464 +/- .010 | Use As Is | [14 empty] | ISI-0041 | OGLE | 2011-05-24 | 23A55888-A451-332D-5DB982B11AEED214 |  | 2011-06-03 | Use As Is
608 | ISI-0042 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 161 | SCRATCH IN SEAL GROOVE | OGLE | 2011-05-24 |  |  |  | HKPARK | 2011-06-10 | Polish seal surface | Modify | [14 empty] | ISI-0042 | OGLE | 2011-05-24 | 23AB874F-C07B-847D-6FA40F94DA7A81A6 |  | 2011-06-10 | Modify | Polished on 06/13/11 BC
609 | ISI-0044 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 163 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-05-24 |  |  |  | HKPARK | 2011-06-10 | Found ss connector loose at torque test. Returned to the vendor. | Return To Vendor | [14 empty] | ISI-0044 | OGLE | 2011-05-24 | 23BCA717-AA3A-D0C9-50D3E2D34F45E149 |  | 2011-06-10 | Return To Vendor
610 | ISI-0045 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 164 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-05-24 |  |  |  | HKPARK | 2011-06-03 | The dimension is within 1.464 +/- .010 | Use As Is | [14 empty] | ISI-0045 | OGLE | 2011-05-24 | 23BF54DA-CE75-7D9E-BD6E49901B8198D0 |  | 2011-06-03 | Use As Is
611 | ISI-0046 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 165 | SCRATCH IN SEAL GROOVE | OGLE | 2011-05-24 |  |  |  | HKPARK | 2011-08-19 | Later found the conntor loose, returned to the vendor. | Return To Vendor | [14 empty] | ISI-0046 | OGLE | 2011-05-24 | 23C2C016-9739-792B-4988FBE8225CF503 |  | 2011-08-19 | Return To Vendor
612 | ISI-0047 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 166 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-05-24 |  |  |  | HKPARK | 2011-06-10 | Found ss connector loose at torque test. Returned to the vendor. | Return To Vendor | [14 empty] | ISI-0047 | OGLE | 2011-05-24 | 23C75695-A452-E41C-0522649D32D1DB73 |  | 2011-06-10 | Return To Vendor
613 | ISI-0048 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 167 | CONNECTOR PIN DEPTH: .152/.166 CHECKS .134 DEPTH .187/.207 CHECKS .223 | OGLE | 2011-05-25 |  |  |  | HKPARK | 2011-07-05 | If it makes a connection with a cable without a problem, use as is. | Use As Is | [14 empty] | ISI-0048 | OGLE | 2011-05-25 |  |  | 2011-07-05 | Use As Is
614 | ISI-0049 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 168 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-05-25 |  |  |  | HKPARK | 2011-06-29 | Angle slightly out tol but the probe tip length still meets the spec. | Use As Is | [14 empty] | ISI-0049 | OGLE | 2011-05-25 | 2A2EC255-E360-8366-4DE8631E48CB0B9D |  | 2011-06-29 | Use As Is
615 | ISI-0051 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 170 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-05-25 |  |  |  | HKPARK | 2011-06-03 | The dimension is within 1.464 +/- .010 | Use As Is | [14 empty] | ISI-0051 | OGLE | 2011-05-25 | 2A30B4E0-D6F4-1805-2D789D72AF3A17DE |  | 2011-06-03 | Use As Is
616 | ISI-0012 |  | CRM-120-7014-0020 | HOM Feedthru |  | C100-CAV-INSP-HMFT | R4 | 145 | Dropped during leak testing, which broke sapphire. | CARPENTE | 2011-05-26 |  |  |  | CARPENTE |  |  | Scrap | [14 empty] | ISI-0012 | CARPENTE | 2011-05-26 |  |  |  | Scrap
617 | OML-0117 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 143 | STAINS ON ADAPTER SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-06-02 |  |  |  | HKPARK | 2011-06-03 | The stain on ss adapter is OK. The thickness OK. | Use As Is | [14 empty] | OML-0117 | OGLE | 2011-06-02 | 534DED7F-F984-1EB1-0DBA82541BFF3892 |  | 2011-06-03 | Use As Is
618 | OML-0119 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 145 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-06-02 |  |  |  | HKPARK | 2011-06-03 | Ceramic thickness OK. The radii updated to nominal .688 | Use As Is | [14 empty] | OML-0119 | OGLE | 2011-06-02 | 535046C5-AD0D-1C03-9BB96E981452F3F2 |  | 2011-06-03 | Use As Is
619 | OML-0120 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 146 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-06-02 |  |  |  | HKPARK | 2011-08-18 | Radius has been changed. See the latest drawing. | Use As Is | [14 empty] | OML-0120 | OGLE | 2011-06-02 | 535224B2-F7F7-9ECB-09EBEAA5777B5FFD |  | 2011-08-18 | Use As Is
620 | OML-0121 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 147 | BRAZE SPOT ON CERAMIC SURFACE SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-06-02 |  |  |  | HKPARK |  | Radii are OK. braze spot should be cleaned by the vendor. | Return To Vendor | [14 empty] | OML-0121 | OGLE | 2011-06-02 | 5353E5F7-97E8-BA07-D30DA7294C3FAE6C |  |  | Return To Vendor
621 | OML-0122 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 148 | BLACK SPOTS ON CERAMIC SURFACE SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-06-02 |  |  |  | HKPARK | 2011-08-18 | Radii are OK. Remove the black spot. | Modify | [14 empty] | OML-0122 | OGLE | 2011-06-02 | 53560B5E-D879-A4EA-A73E32C653CA707C |  |  | Modify | cleaned and coated-11/2/11 HKP
622 | C100-RI-062 |  | 12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 72 | FPC FLANGE: SCRATCHES AND NICKS IN SEAL AREA HOM FLANGE "B":FLANGE FACE HAS RO | OGLE | 2011-06-02 |  |  |  | aburrill | 2011-06-17 | The HOM B flange needs to be replaced for a proper sealing seat. See Bill Clemen Cavity to be taken to machine shop for rework | Modify | 30827F05-AF50-3214-E06699EE0E53E5B2 | [13 empty] | C100-RI-062 | OGLE | 2011-06-02 | 53817974-CF4A-64E0-50BBF56EFB62AE71 | aburrill |  | Modify | HOM B flange seal surface re-faced at Jlab machine shop by using 30 um diamond lapping paste and a brass lapping tool. HOM A flange also touched up. Visual & dimensional inspection good. BC 11/18/11
623 | ISI-0053 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 172 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-06-03 |  |  |  | HKPARK | 2011-06-06 | 1.464+/-.010 is acceptable. | Use As Is | [14 empty] | ISI-0053 | OGLE | 2011-06-03 | 58786D15-913B-E3BF-A1C8F6388373774E |  | 2011-06-06 | Use As Is
624 | ISI-0054 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 173 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-06-03 |  |  |  | HKPARK | 2011-06-06 | 1.464+/-.010 is acceptable. | Use As Is | [14 empty] | ISI-0054 | OGLE | 2011-06-03 | 5879C8EF-AB32-2BB1-380D320CC3F802E5 |  | 2011-06-06 | Use As Is
625 | ISI-0055 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 174 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-06-03 |  |  |  | HKPARK | 2011-06-10 | Deviation on angle small. | Use As Is | [14 empty] | ISI-0055 | OGLE | 2011-06-03 | 587ADB54-92F5-E735-6CEDB7FE2A57C211 |  | 2011-06-10 | Use As Is
626 | ISI-0056 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 175 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-06-03 |  |  |  | HKPARK | 2011-06-06 | 1.464+/-.010 is acceptable. | Use As Is | [14 empty] | ISI-0056 | OGLE | 2011-06-03 | 587C00E4-9C48-9392-1B0176890CA0704C |  | 2011-06-06 | Use As Is
627 | ISI-0057 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 176 | SCRATCH IN SEAL GROOVE | OGLE | 2011-06-03 |  |  |  | HKPARK | 2011-06-10 | Polish seal groove. | Modify | [14 empty] | ISI-0057 | OGLE | 2011-06-03 | 587D59A0-AE22-B3DD-F4E2D27A1F722C47 |  | 2011-06-10 | Modify | Seal groove has been successfully polished. BC 08/15/11
628 | C100-RI-063 |  | 12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 73 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-06-03 | B292BE9D-E971-C50D-AD647CC9CCAB9F18 |  |  | ABURRILL | 2011-10-07 | Dimensional issues will be resolved in HV welding | Use As Is | [14 empty] | C100-RI-063 | OGLE | 2011-06-03 | 58AD055B-AEBE-2FF6-BE952C7B489D5C87 |  |  | Remeasure
629 | C100-RI-024 |  |  | C100 cavuty |  | C100-CAV-VTRF | R3 | 67 | cavity did not meet lock frequency specification. Lock frequency was measured a | FOREHAND | 2011-06-06 |  |  | C100-RI-024 did not meet lock frequency specification. Frequency was measured a | HOGAN | 2011-06-10 | C100-RI-024 is not installed in HV yet. The cavity will still be tuned before HV | Use As Is | [14 empty] | C100-RI-024 | FOREHAND | 2011-06-06 |  |  | 2011-06-10 | Use As Is
630 | F356223-03 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-RWRK-WGD | R4 | 60 | Flatness out of tolerance, measured .006" after lapping | GEORGED | 2011-06-07 |  |  |  | KWILSON |  |  | Use As Is | [14 empty] | F356223-03 | GEORGED | 2011-06-07 |  |  |  | Use As Is
631 | F067764-13 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-RWRK-WGD | R4 | 54 | Flatness out of tolerance, meas. .005" | GEORGED | 2011-06-07 |  |  |  | KWILSON |  |  | Use As Is | [14 empty] | F067764-13 | GEORGED | 2011-06-07 |  |  |  | Use As Is
632 | F067764-11 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R6 | 124 | Scratches on small flange. Surface finish out of tolerance, meas. 17 micro inche | GEORGED | 2011-06-08 |  |  |  | KWILSON |  | Please repair scratches. Surface finish ok. Might get improved with scratch po | Modify | [14 empty] | F067764-11 | GEORGED | 2011-06-08 |  |  |  | Modify
633 | ISI-0044 |  | CRM-120-7014-0020 | HOM Feedthru |  | C100-CAV-INSP-HMFT | R4 | 163 | Stainless Steel connector broke during final RF test. | CARPENTE | 2011-06-08 |  |  |  | HKPARK | 2011-06-08 | This can be salvaged so return to the vendor. | Return To Vendor | [14 empty] | ISI-0044 | CARPENTE | 2011-06-08 | 6FA2B70A-F1E0-05F5-72E2B0A3D28C257E |  | 2011-06-08 | Return To Vendor
634 | ISI-0047 |  | CRM-120-7014-0020 | HOM Feedthru |  | C100-CAV-INSP-HMFT | R4 | 166 | Stainless Steel connector broke during final RF test. | CARPENTE | 2011-06-08 |  |  |  | HKPARK | 2011-06-08 |  | Return To Vendor | [14 empty] | ISI-0047 | CARPENTE | 2011-06-08 | 6FA3A773-0D0C-E64F-D9BAA5C605CEAC79 |  | 2011-06-08 | Return To Vendor
635 | F892523-02 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R6 | 119 | Scratches on small flange. Flatness out of tolerance (meas. .004) Fingerprints o | GEORGED | 2011-06-08 |  |  |  | KWILSON |  | Repair scratches. Clean. Flatness ok. | Modify | [14 empty] | F892523-02 | GEORGED | 2011-06-08 |  |  |  | Modify
636 | F356224-07 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R6 | 123 | Scratches on small flange. Surface finish meas. 19 micro inches. Nicks on knife | GEORGED | 2011-06-08 |  |  |  | KWILSON |  | Please repair scratches and nicks. | Modify | [14 empty] | F356224-07 | GEORGED | 2011-06-08 |  |  |  | Modify
637 | ISI-0040 |  | CRM120-7014-0020 | HOM Feedthrough |  | C100-CAV-INSP-HMFT | R4 | 159 | Connector end of feedthru broke during torque test. Did not do pull test. | CARPENTE | 2011-06-09 |  |  |  | HKPARK | 2011-06-09 |  | Return To Vendor | [14 empty] | ISI-0040 | CARPENTE | 2011-06-09 |  |  | 2011-06-09 | Return To Vendor
638 | ISI-0046 |  | CRM120-7014-0020 | HOM Feedthrough |  | C100-CAV-INSP-HMFT | R4 | 165 | Connector end of feedthru broke during torque test. Did not do pull test. | CARPENTE | 2011-06-09 |  |  |  | HKPARK | 2011-06-09 |  | Return To Vendor | [14 empty] | ISI-0046 | CARPENTE | 2011-06-09 |  |  | 2011-06-09 | Return To Vendor
639 | ISI-0050 |  | CRM120-7014-0020 | HOM Feedthrough |  | C100-CAV-INSP-HMFT | R4 | 169 | Connector end of feedthru broke during torque test. Did not do pull test. | CARPENTE | 2011-06-09 |  |  |  | HKPARK | 2011-06-09 |  | Return To Vendor | [14 empty] | ISI-0050 | CARPENTE | 2011-06-09 |  |  | 2011-06-09 | Return To Vendor
640 | ISI-0054 |  | CRM120-7014-0020 | HOM Feedthrough |  | C100-CAV-INSP-HMFT | R4 | 173 | Connector end of feedthru broke during torque test. | CARPENTE | 2011-06-09 |  |  |  | HKPARK | 2011-06-09 |  | Return To Vendor | [14 empty] | ISI-0054 | CARPENTE | 2011-06-09 |  |  | 2011-06-09 | Return To Vendor
641 | ISI-0013 |  | CRM120-7014-0020 | HOM Feedthrough |  | C100-CAV-INSP-HMFT | R4 | 146 | Failed torque test. 06/09/11 BC | CARPENTE | 2011-06-09 |  |  |  | HKPARK | 2011-06-09 |  | Return To Vendor | [14 empty] | ISI-0013 | CARPENTE | 2011-06-09 |  |  | 2011-06-09 | Return To Vendor
642 | ISI-JL-48 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 150 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | OGLE | 2011-06-09 |  |  |  | HKPARK | 2011-06-10 | Ceramic thickness OK. | Use As Is | [14 empty] | ISI-JL-48 | OGLE | 2011-06-09 | 769AD294-0BDB-34D4-E670220C45BBD24D |  | 2011-06-10 | Use As Is
643 | ISI-JL-49 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 151 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-06-09 |  |  |  | HKPARK | 2011-06-10 | Ceramic thickness OK. | Use As Is | [14 empty] | ISI-JL-49 | OGLE | 2011-06-09 | 76C8124D-C219-0DFC-779F9FFEB67B5715 |  | 2011-06-10 | Use As Is
644 | ISI-JL-51 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 153 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-06-09 |  |  |  | HKPARK | 2011-06-10 | Ceramic thickness OK. | Use As Is | [14 empty] | ISI-JL-51 | OGLE | 2011-06-09 | 76CA10B2-EF1B-CBB5-4228C633570A5190 |  | 2011-06-10 | Use As Is
645 | ISI-JL-52 |  | 12070750024 | BRAZED WINDOW ADAPTER |  | C100-CAV-INSP-ADPT | R3 | 154 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-06-09 |  |  |  | HKPARK | 2011-06-10 | Ceramic thickness OK. | Use As Is | [14 empty] | ISI-JL-52 | OGLE | 2011-06-09 | 76CC1468-0F49-B133-74DDD2726B8BB7C3 |  | 2011-06-10 | Use As Is
646 | C100_RI_037 |  |  |  | Travel C100-CAV-VTRF seq #68 | C100-CAV-VTRF | R3 | 68 | Cavity C100_RI_037 didn't pass VTA "canary" test: Radiation onset at 7 MV/m, Qo | HOGAN | 2011-06-10 |  |  |  | HAIPENG | 2011-06-17 | Reprocess it. | Modify | [9 empty] | ABURRILL | 2011-10-07 | Cavity meets spec | Use As Is |  | C100_RI_037 | HOGAN | 2011-06-10 |  |  | 2011-06-17 | Modify
647 | C100-RI-026 |  | 12070130001 | CAVITY |  | C100-CAV-INSP2 | R3 | 44 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-06-10 |  |  |  | HAIPENG |  | We need a RF bench measurement to determine the outcome of this cavity by both f | Remeasure | [14 empty] | C100-RI-026 | OGLE | 2011-06-10 | 7BB86EC9-BC55-7700-B8F9C13F9E99680A |  |  | Use As Is
648 | C100-RI-032 |  | 12070130001 | CAVITY |  | C100-CAV-INSP2 | R3 | 46 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-06-10 |  |  |  | HAIPENG | 2011-06-17 | This dimension could affect Qext of the cavity, we need a RF bench measurement t | Use As Is | [14 empty] | C100-RI-032 | OGLE | 2011-06-10 | 7C10E272-965E-E66A-7A334CE9E78AAB4E |  | 2011-06-17 | Use As Is | This cavity has been successfully tested in the VTA on 20-April-2011 with no iss
649 | ISI-0026 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 177 | Part failed final leak test. Step depth From the pin end to the connector step | OGLE | 2011-06-10 |  |  |  | HKPARK | 2011-08-01 | HyeKyoung will contact the vendor and get it ready to be returned. | Return To Vendor | [14 empty] | ISI-0026 | OGLE | 2011-06-10 |  |  | 2011-08-01 | Return To Vendor
650 |  |  |  | C100_RI_036 | Cavity lock frequency is above specification. Measured 2K lock frequency of cav | C100-CAV-VTRF | R3 | 71 |  | FOREHAND | 2011-06-13 |  |  |  | HAIPENG | 2011-06-14 | I saw the lock frequency at 2K was 1496.670MHz less than the spec of upper limit | Remeasure | [15 empty] | FOREHAND | 2011-06-13 |  |  |  | Use As Is
651 | C100-RI-064 |  | 12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 74 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-06-15 | B2FC24D5-D780-8C76-5F19B1FF07D0963C |  |  | ABURRILL | 2011-06-17 | Dimensional issues will be resolved in CM assembly | Use As Is | [14 empty] | C100-RI-064 | OGLE | 2011-06-15 | 95F01D4A-A6CF-87E6-C6F49B1A30FFD1D3 |  | 2011-06-17 | Modify
652 | MPF-0044 |  | CRM1207014-0020 | HOM Feedthru |  | C100-CAV-INSP-HMFT | R3 | 84 | This feedthru was found to be broken as it was being assembled onto cavity C100- | CARPENTE | 2011-06-16 |  |  |  | HKPARK | [17 empty] | MPF-0044 | CARPENTE | 2011-06-16 |  |  |  | Scrap
653 | C100-RI-065 |  | 12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 75 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING FPC flange surface finish = 24 ui | OGLE | 2011-06-16 | B30E0275-B14F-E6CB-A637B5BDB5FE9657 |  |  | HAIPENG | 2011-06-17 | If it is the FPC surface finish problem, we can try to re-lap the surface. If it | Modify | [9 empty] | ABURRILL | 2011-10-07 | Dimensional issues will be resolved in HV welding and CM assembly | Use As Is |  | C100-RI-065 | OGLE | 2011-06-16 | 9B77599A-EFCF-62FA-716D4310EB8A23D4 |  | 2011-06-17 | Modify
654 | ISI-0058 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R5 | 179 | CENTER PIN NOT FLEXIBLE AT ALL. NO BRAZE MATERIAL AROUND CONNECTOR. FAILED TOR | OGLE | 2011-06-20 |  |  |  | HKPARK | 2011-06-27 |  | Return To Vendor | [14 empty] | ISI-0058 | OGLE | 2011-06-20 | AE089000-A2AB-979D-85C2084E716EA71F |  | 2011-06-27 | Return To Vendor
655 | ISI-0059 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R5 | 180 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-06-20 |  |  |  | HKPARK | 2011-06-27 |  | Use As Is | [14 empty] | ISI-0059 | OGLE | 2011-06-20 | AE29001E-DE92-C1A0-0FB09CFAE9738C8E |  | 2011-06-27 | Use As Is
656 | ISI-0060 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R5 | 181 | SCRATCH IN SEAL GROOVE. MATERIAL COVERING SAPPHIRE SURFACE INSIDE CONNECTOR. SE | OGLE | 2011-06-20 |  |  |  | HKPARK | 2011-06-27 | Before returning it to the vendor RF test to see how it affects RF. | Return To Vendor | [14 empty] | ISI-0060 | OGLE | 2011-06-20 | AE6A912F-A338-D88C-B2BFE7418ECD639A |  | 2011-06-27 | Return To Vendor
657 | C100-RI-066 |  | 12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 76 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | OGLE | 2011-06-24 |  |  |  | ABURRILL | 2011-08-03 | Deviations from drawing dimensions are acceptable. | Use As Is | [14 empty] | C100-RI-066 | OGLE | 2011-06-24 | C2FDFF99-9899-37B7-C9CCD0D87AABE2C2 |  | 2011-08-03 | Use As Is
658 | C100-RI-067 |  | 12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 77 | FPC FLANGE RA CHECKS 19 MICRO INCH. SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMI | OGLE | 2011-06-24 | B319F7E2-9A78-81C5-877403A933383D03 |  |  | ABURRILL | 2011-10-07 | Dimensional issues will be resolved in HV welding and CM assembly | Use As Is | [14 empty] | C100-RI-067 | OGLE | 2011-06-24 | C3722E3B-FBB8-3BF2-FE2E65A30334720D
659 | F890195-04 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R6 | 122 | Scratches on the small flange. Nick on knife edge. | GEORGED | 2011-06-28 |  |  |  | KWILSON |  | Please repair scratches and nick. | Modify | [14 empty] | F890195-04 | GEORGED | 2011-06-28 |  |  |  | Modify
660 | F356223-04 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R6 | 129 | Scratches on the small flange. Length and flatness out of tolerance. Flatness me | GEORGED | 2011-06-29 |  |  |  | KWILSON |  | Repair scratches. Length and flatness ok. | Modify | [14 empty] | F356223-04 | GEORGED | 2011-06-29 |  |  |  | Modify
661 | F892522-09 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-RWRK-WGD | R4 | 40 | Scratches on the small flange | GEORGED | 2011-06-29 |  |  |  | KWILSON |  | Please repair scratches. | Modify | [14 empty] | F892522-09 | GEORGED | 2011-06-29 |  |  |  | Modify
662 | C100-RI-068 |  | 12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 78 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-06-29 |  |  |  | ABURRILL | 2011-08-03 | Deviations from drawing dimensions are acceptable. | Use As Is | [14 empty] | C100-RI-068 | OGLE | 2011-06-29 | DE4AD87C-ED67-35C8-03A2949FE025B5EC |  | 2011-08-03 | Use As Is
663 | C100-RI-069 |  | 12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 79 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-06-30 |  |  |  | ABURRILL | 2011-08-03 | Deviations from drawing dimensions are acceptable. | Use As Is | [14 empty] | C100-RI-069 | OGLE | 2011-06-30 | E3938F35-A4B4-0E8F-9A63CCB093C6DF99 |  | 2011-08-03 | Use As Is
664 | F893562-01 |  | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R6 | 130 | Scratches on the small flange. Nicks on knife edge. | GEORGED | 2011-07-05 |  |  |  | KWILSON |  | Please repair scratches and nicks. Thanks. | Modify | [14 empty] | F893562-01 | GEORGED | 2011-07-05 |  |  |  | Modify
665 | C100-RI-043 |  | 12070130001 | CAVITY |  | C100-CAV-INSP2 | R3 | 48 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-07-05 |  |  |  | ABURRILL | 2011-08-03 | Deviations from drawing dimensions are acceptable. | Use As Is | [14 empty] | C100-RI-043 | OGLE | 2011-07-05 | FC9B3A54-EBF1-43FD-F60C84B1CF7B146A |  | 2011-08-03 | Use As Is
666 | C100-RI-070 |  | 12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 80 | SCRATCHES NICKS AT APPROXIMATE DEPTH OF .0001"; AND STAINS ON FLANGE FACE. FPC | OGLE | 2011-07-05 |  |  |  | HAIPENG |  | See this FPC surface can be re-lapped to meet the spec. | Use As Is | [14 empty] | C100-RI-070 | OGLE | 2011-07-05 | FD506789-D963-605F-BAFD9B5B84E140F5 |  |  | Use As Is
667 | C100-RI-042 |  | 12070130001 | C100 Cavity |  | C100-CAV-INSP2 | R3 | 47 | Please see attached PDF | GEORGED | 2011-07-06 |  |  |  | ABURRILL | 2011-08-03 | Deviations from drawing dimensions are acceptable. | Use As Is | [14 empty] | C100-RI-042 | GEORGED | 2011-07-06 | 05BD806E-CFEF-5388-5F4A94A7C9E7B651 |  | 2011-08-03 | Use As Is
668 | C100-RI-071 |  | 12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 81 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-07-08 | B338BC19-F974-6FD0-E0914EB0A5EF01C9 |  |  | ABURRILL | 2011-10-07 | Dimensional issues will be resolved in HV welding and CM assembly | Use As Is | [14 empty] | C100-RI-071 | OGLE | 2011-07-08 | 0C204C02-E394-BE9C-75854E1E379FF9BF
669 | VTA-TH-07 |  | CRM1207015-0104 | Tophat |  | C100-CAV-LAP-THAT | R1 | 31 | Surface finish out of tolerance, measured 23 micro inches. | GEORGED | 2011-07-12 |  |  |  | ABURRILL | 2011-08-03 | Deviations acceptable. | Use As Is | [14 empty] | VTA-TH-07 | GEORGED | 2011-07-12 |  |  | 2011-08-03 | Use As Is | Awaiting return of other profilometer for cross reference and calibration
670 | WG-1 |  | CRM1207015-0104 | Tophat |  | C100-CAV-LAP-THAT | R1 | 32 | Surface finish out of tolerance, meas. 27 micro inches | GEORGED | 2011-07-12 |  |  |  | ABURRILL | 2011-08-03 | Deviation from required acceptable. | Use As Is | [14 empty] | WG-1 | GEORGED | 2011-07-12 |  |  | 2011-08-03 | Use As Is
671 | C100-RI-044 |  | 12070130001 | CAVITY |  | C100-CAV-INSP2 | R3 | 49 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | OGLE | 2011-07-14 |  |  |  | ABURRILL | 2011-08-03 | Deviations from drawing dimensions are acceptable. | Use As Is | [14 empty] | C100-RI-044 | OGLE | 2011-07-14 | 2B11332F-C48D-D8BB-294496ADC1313AE3 |  | 2011-08-03 | Use As Is
672 | F892522-08 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R6 | 135 | Surface finish out of tolerance (58 micro inches). Scratches on the small flange | GEORGED | 2011-07-15 |  |  |  | KWILSON |  | Please repair scratche and nick on knife edge | Modify | [14 empty] | F892522-08 | GEORGED | 2011-07-15 |  |  |  | Modify
673 | C100-RI-072 |  | 12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 82 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | OGLE | 2011-07-15 |  |  |  | ABURRILL | 2011-08-03 | Deviations from drawing dimensions are acceptable. | Use As Is | [14 empty] | C100-RI-072 | OGLE | 2011-07-15 | 302D7CA9-91C5-BF8F-4DC3EB5CB638FC69 |  | 2011-08-03 | Use As Is
674 | C100-RI-073 |  | 12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 83 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | OGLE | 2011-07-15 |  |  |  | ABURRILL | 2011-08-03 | Deviations from drawing dimensions are acceptable. | Use As Is | [14 empty] | C100-RI-073 | OGLE | 2011-07-15 | 30AD2E5A-D7C2-F1BD-6FB0311877F261DB |  | 2011-08-03 | Use As Is
675 | C100-RI-074 |  | 12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 84 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | OGLE | 2011-07-18 | B341DC08-9269-A581-6BD2487C1FAE5C2F |  |  | ABURRILL | 2011-10-07 | Dimensional issues will be resolved in HV welding and CM assembly | Use As Is | [14 empty] | C100-RI-074 | OGLE | 2011-07-18 | 4033D38A-9E59-737F-61E65CB9FC97AA79
676 | C100-RI-045 |  | 12070130001 | CAVITY |  | C100-CAV-INSP2 | R3 | 50 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-07-19 |  |  |  | ABURRILL | 2011-08-03 | Deviations from drawing dimensions are acceptable. | Use As Is | [14 empty] | C100-RI-045 | OGLE | 2011-07-19 | 44B15214-F3C8-8CE9-2C86642BCDAE4F8D |  | 2011-08-03 | Use As Is
677 | F063134-07 |  | CRM-12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R6 | 136 | Nicks on the knife edge. | GEORGED | 2011-07-20 |  |  |  | KWILSON |  | Please repair knife edge. Thanks. | Modify | [14 empty] | F063134-07 | GEORGED | 2011-07-20 |  |  |  | Modify
678 | C100-RI-075 |  | 12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 85 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. | OGLE | 2011-07-22 |  |  |  | ABURRILL | 2011-08-03 | Deviations from drawing dimensions are acceptable. | Use As Is | [14 empty] | C100-RI-075 | OGLE | 2011-07-22 | 483D008B-B114-5B1C-1E6953F45EFCA530 |  | 2011-08-03 | Use As Is
679 | C100-RI-076 |  | 12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 86 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-07-25 |  |  |  | ABURRILL | 2011-08-03 | Deviations from drawing dimensions are acceptable. | Use As Is | [14 empty] | C100-RI-076 | OGLE | 2011-07-25 | EA2C5108-E694-EE31-03BAC9538F0A9519 |  | 2011-08-03 | Use As Is
680 | ISI-0002 |  | CRM120-7014-0020 | HOM Feedthru |  | C100-CAV-D3 | R1 | 198 | Feedthru was broken after being dropped during dis-assembly of C100-035. | CARPENTE | 2011-07-27 |  |  |  | HKPARK | 2011-07-29 |  | Scrap | [14 empty] | ISI-0002 | CARPENTE | 2011-07-27 |  | hkpark | 2011-07-29 | Scrap
681 | C100-RI-077 |  | 12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 87 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-07-27 |  |  |  | ABURRILL |  | Deviations from drawing dimensions are acceptable. | Use As Is | [14 empty] | C100-RI-077 | OGLE | 2011-07-27 | 535DC5EA-F3CA-89D9-9F4D05146CF6D3FA |  |  | Use As Is
682 | F061359-11 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R6 | 140 | Extent of Cu plating on large flange face bad (0.082"). Raised edge on flange in | GEORGED | 2011-07-28 |  |  |  | KWILSON |  | Please continue with leak check. Raised area near knife edge ok if passes leak | Remeasure | [14 empty] | F061359-11 | GEORGED | 2011-07-28 |  |  |  | Remeasure
683 | F063952-04 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R3 | 23 | Scratches on the small flange. Surface finish out of tolerance (meas. 21 micro i | GEORGED | 2011-07-29 |  |  |  | KWILSON |  | Please repair scratches | Modify | [14 empty] | F063952-04 | GEORGED | 2011-07-29 |  |  |  | Modify
684 | F067764-16 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R5 | 100 | Scratches on the small flange. Damage inside o-ring groove. Flatness measured .0 | GEORGED | 2011-07-29 |  |  |  | KWILSON |  | Please repair scratches. | Modify | [14 empty] | F067764-16 | GEORGED | 2011-07-29 |  |  |  | Modify
685 | F892522-03 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R6 | 131 | Scratches on the small flange. Copper flaking on large flange face. Surface fini | GEORGED | 2011-07-29 |  |  |  | KWILSON |  | Please lightly scotchbrite flange with flaking copper to ensure there are no loo | Modify | [14 empty] | F892522-03 | GEORGED | 2011-07-29 |  |  |  | Modify
686 | F356225-01 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R6 | 133 | Part failed tape test,lots of copper comes off the large flange. Surface finish | GEORGED | 2011-08-01 |  |  |  | KWILSON |  |  | Return To Vendor | [14 empty] | F356225-01 | GEORGED | 2011-08-01 |  |  |  | Return To Vendor
687 | F356224-06 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R6 | 126 | Scratches on the small flange, finish meas. 22 micro inches. Nick on the knife e | GEORGED | 2011-08-02 |  |  |  | KWILSON |  | Please repair scratches and knife edge. | Modify | [14 empty] | F356224-06 | GEORGED | 2011-08-02 |  |  |  | Modify
688 | C100-RI-046 |  | CRM1207013-0001 | C100 Cavity |  | C100-CAV-INSP2 | R3 | 51 | FPC flange location and Beamline flange(FP) True position out of tolerance. See | CARPENTE | 2011-08-03 |  |  |  | ABURRILL | 2011-08-03 | Deviations from drawing dimensions are acceptable. | Use As Is | [14 empty] | C100-RI-046 | CARPENTE | 2011-08-03 | 9FF45C80-9685-9FAA-D4B7A190091AEEE2 |  | 2011-08-03 | Use As Is
689 | C100-RI-043 |  | 12070200000-INSP | HELIUM VESSEL FINAL WELD |  | C100-CAV-INST-HELV | R1 | 36 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-08-03 |  |  |  | HOGAN |  | The perpendicular dimension of the end flange is out of tolerance by 0.006". Th | Use As Is | [14 empty] | C100-RI-043 | OGLE | 2011-08-03 | BA096430-BB9F-7283-D579BBFCD3C71A78 |  |  | Use As Is
690 | C100-RI-078 |  | 12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 88 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-08-05 |  |  |  | HOGAN |  | The worse out of tolerance is .017" in perpendicularity. This can be corrected | Use As Is | [14 empty] | C100-RI-078 | OGLE | 2011-08-05 | 1186676C-BC08-152C-3813EC3B6C60CEE7 |  |  | Use As Is
691 | C100_RI_041 |  |  |  |  | C100-CAV-VTRF | R3 | 78 | Cavity lock frequency was not with in specification. Lock frequency measured 14 | FOREHAND | 2011-08-08 |  |  |  | ABURRILL | 2011-08-08 | Cavity does not have a helium vessel attached, so the lock frequency being out o | Use As Is | [14 empty] | C100_RI_041 | FOREHAND | 2011-08-08 |  |  | 2011-08-08 | Use As Is
692 | C100-RI-008 |  | 12070130001 | CAVITY |  | C100-CAV-INSP2 | R3 | 52 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-08-08 |  |  |  | ABURRILL | 2011-08-09 | Cavity dimensional discrepancy is acceptable for use. | Use As Is | [14 empty] | C100-RI-008 | OGLE | 2011-08-08 | AEAE9245-FBBA-6D4E-D5E45A2C464511D2 |  | 2011-08-09 | Use As Is
693 | C100-RI-079 |  | 12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 89 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-08-08 |  |  |  | ABURRILL | 2011-08-09 | Cavity dimensional discrepancy is acceptable for use. | Use As Is | [14 empty] | C100-RI-079 | OGLE | 2011-08-08 | B4DFF59C-038D-025E-86C8C6CF61A8C20D |  | 2011-08-09 | Use As Is
694 | FLG0037-118 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 109 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-08-08 |  |  |  | HKPARK | 2011-08-15 | Out tol. value acceptable for assembly. | Use As Is | [14 empty] | FLG0037-118 | OGLE | 2011-08-08 | BCA41EA0-BA3C-7390-2836EEC47E9B4ABB |  | 2011-08-15 | Use As Is
695 | FLG0037-124 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 115 | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING | OGLE | 2011-08-08 |  |  |  | HKPARK | 2011-08-15 | Out tol. value OK for assembly. | Use As Is | [14 empty] | FLG0037-124 | OGLE | 2011-08-08 | BCB5C1AD-9A98-A297-3B3491E96C9DF105 |  | 2011-08-15 | Use As Is
696 | FLG0037-125 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 116 | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING | OGLE | 2011-08-08 |  |  |  | HKPARK | 2011-08-15 | Out tol. value OK for assembly. | Use As Is | [14 empty] | FLG0037-125 | OGLE | 2011-08-08 | BCC79F46-9821-C98B-FC0F095D748652F6 |  | 2011-08-15 | Use As Is
697 | F892522-04 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R6 | 128 | O-ring groove has damage. Scratches on the small flange. Finish out of tolerance | GEORGED | 2011-08-09 |  |  |  | KWILSON |  | Please repair scratches. | Modify | [14 empty] | F892522-04 | GEORGED | 2011-08-09 |  |  |  | Modify
698 | C100-044 |  | CRM120-7020-1000 | Helium Vessel Assembly |  | C100-CAV-INST-HELV | R1 | 38 | Helium Port locations out of tolerance. See attachment. | CARPENTE | 2011-08-09 |  |  |  | FISCHER | 2011-09-13 |  | Use As Is | [14 empty] | C100-044 | CARPENTE | 2011-08-09 | DB78BF74-A6CE-178B-9B5ACFFE45FCDEBF |  | 2011-09-13 | Use As Is
699 | C100-RI-040 |  | CRM-120-7020-0000 | C100 Helium Vessel Assembly |  | C100-CAV-INST-HELV | R1 | 39 | Helium inlet location out of tolerance. See attachment. | CARPENTE | 2011-08-10 |  |  |  | HOGAN |  | The out of tolerance can be accommodated by the header assembly. | Use As Is | [14 empty] | C100-RI-040 | CARPENTE | 2011-08-10 | 0A27DD4F-BDFA-9E5D-127422BF50B70628 |  |  | Use As Is
700 | C100_RI_039 |  |  | Cavity | C100_RI_039A | C100-CAV-VTRF | R3 | 79 | Cavity lock frequency high out of specification. Locked at 1496.8668MHz. High Sp |  | 2011-08-10 |  |  |  | ABURRILL | 2011-08-12 | Cavity is not in a helium vessel where the target frequency will shift once weld | Use As Is | [14 empty] | C100_RI_039 |  | 2011-08-10 |  |  | 2011-08-12 | Use As Is
701 | C100-RI-085 |  | 12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 90 | See attachment for dimensional non-conforming. FPC Flange finish checks 19 micr | OGLE | 2011-08-10 | B25F4C15-B715-8AAE-5CA3D48CC29762C2 |  |  | ABURRILL | 2011-10-07 | Dimensional issues will be resolved in CM assembly | Use As Is | [14 empty] | C100-RI-085 | OGLE | 2011-08-10 | 1DC2BF88-99C3-5C3A-2D12F517157B117E
702 | C100-RI-086 |  | 12070130001 | CAVITY |  | C100-CAV-INSP | R3 | 91 | See attachment for dimensional non-conforming. FPC Finish checks 22 micro inch. | OGLE | 2011-08-10 | B2455EC1-AF96-9188-5B583642F6123353 |  |  | HAIPENG | 2011-08-25 | Please see Bill Clemens and John Fisher to see if this weld defect can be repair | Modify | [9 empty] | ABURRILL | 2011-10-07 | Dimensional issues will be resolved in CM assembly | Use As Is |  | C100-RI-086 | OGLE | 2011-08-10 | 233DBD90-B393-0713-1745475D88621E70 |  | 2011-08-25 | Modify
703 | C100-RI-047 |  | 12070130001 | CAVITY |  | C100-CAV-INSP2 | R3 | 53 | See attachment for dimensional non-conforming. | OGLE | 2011-08-11 |  |  |  | ABURRILL | 2011-08-12 | Dimensional offset will be taken up during assembly | Use As Is | [14 empty] | C100-RI-047 | OGLE | 2011-08-11 | 4FA58590-B8C4-C512-F0CCCC2ECA1804F3 |  | 2011-08-12 | Use As Is
704 | C100-RI-041 |  | 12070130001 | CAVITY |  | C100-CAV-INSP2 | R3 | 54 | Sorry correction: Part within tolerance. Disregard NCR. JTO | OGLE | 2011-08-11 |  |  |  | ABURRILL | 2011-08-12 | Part within tolerance | Use As Is | [14 empty] | C100-RI-041 | OGLE | 2011-08-11 | 5401B27D-BD1B-D6C5-25A6C7EA97AB30BB |  | 2011-08-12 | Use As Is
705 | C100-026 |  | CRM1207020-0000 | Helium Vessel Assembly |  | C100-CAV-INST-HELV | R1 | 40 | Helium outlet location is out of tolerance | CARPENTE | 2011-08-12 |  |  |  | HOGAN |  | The z-dimension is out of tolerance and has the most ability to be compliant at | Use As Is | [14 empty] | C100-026 | CARPENTE | 2011-08-12 | 76F67FCD-0939-57E7-6A6E130595D37CFD |  |  | Use As Is
706 | VTA-TH-08 |  | CRM-120-7015-0104 | C100 Cavity Tophat |  | C100-CAV-LAP-THAT | R1 | 34 | Surface finish out of tolerance, meas. 26 micro inches | GEORGED | 2011-08-16 |  |  |  | ABURRILL | [17 empty] | VTA-TH-08 | GEORGED | 2011-08-16 |  |  |  | Modify
707 | C100-RI-032 |  |  | Cavity |  | C100-CAV-TUNE | R3 | 44 | The cavity does not meet field flatness specification after the helium vessel wa | FOREHAND | 2011-08-17 |  |  |  | ABURRILL |  | Cavity suitable for use in string assembly | [15 empty] | C100-RI-032 | FOREHAND | 2011-08-17 | 783783E5-D1ED-BAF8-37FBE3A0DCB2050D |  |  | Use As Is
708 | C100-RI-008 |  | CRM-120-7070-0000 | C100 Helium Vessel Assembly |  | C100-CAV-INST-HELV | R1 | 42 | Cavity length and helium inlet port location out of tolerance. | CARPENTE | 2011-08-17 |  |  |  | FISCHER | 2011-09-13 |  | Use As Is | [14 empty] | C100-RI-008 | CARPENTE | 2011-08-17 | 7CC00BB7-9ABF-EFF6-E80F1DB86183109D |  | 2011-09-13 | Use As Is
709 | C100-RI-044 |  |  |  |  | C100-CAV-VTRF | R3 | 81 | The resonance frequency is out of range by 20 kHz | ABURRILL | 2011-08-18 |  |  |  | HAIPENG | 2011-09-01 | Tuner manual pre-load in the cryomodule assembly will take care of this small di | Use As Is | [14 empty] | C100-RI-044 | ABURRILL | 2011-08-18 |  |  | 2011-09-01 | Use As Is
710 | C100-RI-042 |  |  |  |  | C100-CAV-VTRF | R3 | 82 | Emax is 17 MV/m below specification of 19.2 MV/m. | ABURRILL | 2011-08-19 |  |  |  | ABURRILL |  | re-rinse | Use As Is | [14 empty] | C100-RI-042 | ABURRILL | 2011-08-19 |  |  |  | Use As Is
711 | [5 empty] | C100-CAV-VTRF | R3 | 82 | Passband processing was not successful. | ABURRILL | 2011-08-19 |  |  |  | ABURRILL |  | cavity requires defect repair or use below spec | [16 empty] | ABURRILL | 2011-08-19 |  |  |  | Modify
712 | C100-RI-040 |  |  |  |  | C100-CAV-VTRF | R3 | 83 | Cavity was field emission loaded below 16 MV/m. Cavity was in cleanroom during | ABURRILL | 2011-08-22 |  |  |  | ABURRILL | 2011-09-01 | re-rinse | Use As Is | [14 empty] | C100-RI-040 | ABURRILL | 2011-08-22 |  |  | 2011-09-01 | Use As Is
713 | C100-RI-026 |  |  |  |  | C100-CAV-VTRF | R3 | 84 | Cavity was quench limited below design specification. Cavity requires further E | ABURRILL | 2011-08-23 |  |  |  | haipeng |  | repair or accept below spec. | Modify | [14 empty] | C100-RI-026 | ABURRILL | 2011-08-23 |  |  |  | Modify | This cavity has been optically inspected. A cat eye defect had been found near equator. Based on ILC experience, this kind of defect normally caused a hard quench and a repair by remelting of the defect with a aimed EBW is possible. This cavity had been taken out of service not was installed into C100 strings.
714 | RI-PRO-034 |  | CRM120-7014-0020 | HOM Feedthru |  | C100-CAV-D3 | R1 | 222 | Center pin has de-mounted itself from HOM assembly. See D3 for more info. 1st p | CARPENTE | 2011-08-24 |  |  |  | HKPARK | 2011-08-24 | Keep the part until further notice. | Scrap | [14 empty] | RI-PRO-034 | CARPENTE | 2011-08-24 | E5A62A7B-A768-F17C-32F482FE3690D2EE | hkpark | 2011-08-24 | Scrap
715 | C100-RI-048 |  | 12070130001 | CAVITY |  | C100-CAV-INSP2 | R3 | 56 | See attachment for dimensional non conforming. | OGLE | 2011-08-24 |  |  |  | HAIPENG | 2011-08-25 | The waveguide bellows will take care this small deviation in the cryomodule asse | Use As Is | [14 empty] | C100-RI-048 | OGLE | 2011-08-24 | F01A29D7-BD21-BA9B-70693B674E6A8FB2 |  | 2011-08-25 | Use As Is
716 | He VESSEL 046 |  |  | FPC END HEAD 4810- |  | C100-CAV-INST-HELV | R2 | 45 | HEAD MEASURES 4.0 Mu ALL OVER EXCEPT TABS THEY MEASURE 1.2 Mu | MCCREA | 2011-08-25 |  |  | helium vessel was assembled into string and now in operation. Was deemed useable by Hogan. | FISCHER | 2012-05-31 | Cavity was tested in VTA ,no apparent effects see due to high Mu field. Unit to be used...Hogan. | Use As Is | [14 empty] | He VESSEL 046 | MCCREA | 2011-08-25 |  |  |  | Use As Is
717 | F892523-05 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R6 | 127 | Scratches on the small flange | GEORGED | 2011-08-25 |  |  |  | KWILSON |  | Please polish. | Modify | [14 empty] | F892523-05 | GEORGED | 2011-08-25 |  |  |  | Modify
718 | FLG0037-132 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 123 | Damaged knife edge. | OGLE | 2011-08-29 |  |  |  | HKPARK | 2011-08-31 |  | Return To Vendor | [14 empty] | FLG0037-132 | OGLE | 2011-08-29 | E64095F7-EF2A-1114-75C837DFC3EE2486 |  | 2011-08-31 | Return To Vendor
719 | FLG0037-133 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 124 | Damaged knife edge. See attachment for dimensional non-conforming. | OGLE | 2011-08-29 |  |  |  | HKPARK | 2011-08-31 |  | Return To Vendor | [14 empty] | FLG0037-133 | OGLE | 2011-08-29 | E720253C-B38A-42E6-00BF62860A8EFB6E |  | 2011-08-31 | Return To Vendor
720 | FLG0037-134 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 125 | Damaged knife edge. | OGLE | 2011-08-29 |  |  |  | HKPARK | 2011-09-02 |  | Return To Vendor | [14 empty] | FLG0037-134 | OGLE | 2011-08-29 | E761E74C-91F8-3486-99DB91EDA21101B4 |  | 2011-09-02 | Return To Vendor
721 | FLG0037-135 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 126 | Damaged knife edge. | OGLE | 2011-08-29 |  |  |  | HKPARK | 2011-09-02 |  | Return To Vendor | [14 empty] | FLG0037-135 | OGLE | 2011-08-29 | E7901908-D067-0DD4-0488FC297C0A8628 |  | 2011-09-02 | Return To Vendor
722 | FLG0037-136 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 127 | Knife edge damage. See attachment for dimensional non-conforming. | OGLE | 2011-08-29 |  |  |  | HKPARK | 2011-08-31 |  | Return To Vendor | [14 empty] | FLG0037-136 | OGLE | 2011-08-29 | E7FF58B8-F795-34D5-D12F22F2B77977FB |  | 2011-08-31 | Return To Vendor
723 | FLG0037-137 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 128 | Damaged knife edge. | OGLE | 2011-08-29 |  |  |  | HKPARK | 2011-08-31 |  | Return To Vendor | [14 empty] | FLG0037-137 | OGLE | 2011-08-29 | E8F08F7A-A925-6A90-D7DFF41A1614ABC6 |  | 2011-08-31 | Return To Vendor
724 | FLG0037-139 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 130 | See attachment for dimensional non-conforming. | OGLE | 2011-08-29 |  |  |  | HKPARK | 2011-08-31 |  | Use As Is | [14 empty] | FLG0037-139 | OGLE | 2011-08-29 | F1349736-F4E9-463D-3C104C2C754E9DB5 |  | 2011-08-31 | Use As Is
725 | FLG0037-140 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 131 | See attachment for dimensional non-conforming. | OGLE | 2011-08-29 |  |  |  | HKPARK | 2011-08-31 |  | Use As Is | [14 empty] | FLG0037-140 | OGLE | 2011-08-29 | F1C94211-DAF4-7256-ED0DA69966420DA5 |  | 2011-08-31 | Use As Is
726 | F892522-06 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R6 | 134 | Deep scratch present on the small flange. Flatness measured .006" and surface fi | GEORGED | 2011-08-31 |  |  |  | KWILSON |  |  | Use As Is | [14 empty] | F892522-06 | GEORGED | 2011-08-31 |  |  |  | Use As Is
727 | F356224-06 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-RWRK-WGD | R4 | 75 | Copper flaking on bellows inside waveguide. Please see attached pictures | GEORGED | 2011-09-02 |  |  |  | KWILSON |  |  | Return To Vendor | [14 empty] | F356224-06 | GEORGED | 2011-09-02 | B5BEA3FA-DEA4-9AB1-D5B1361D7B42438C |  |  | Return To Vendor
728 | C100-RI-041 |  | 12070200000-INSP | HELIUM VESSEL FINAL WELD |  | C100-CAV-INST-HELV | R2 | 47 | See attachment for dimensional non-conforming. | OGLE | 2011-09-07 |  |  |  | FISCHER |  | test fit of tuning plate was performed. Rotation of axis OK. | Use As Is | [14 empty] | C100-RI-041 | OGLE | 2011-09-07 | B4767783-FEB3-B66D-ED4BAB1CAF47E0E1 |  |  | Use As Is
729 | F067763-03 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R6 | 141 | Scratches on the small flange. Many stains inside waveguide. Acetone and isoprop | GEORGED | 2011-09-08 |  |  |  | KWILSON |  | Stains ok. Please polish scratches. | Modify | [14 empty] | F067763-03 | GEORGED | 2011-09-08 |  |  |  | Modify
730 | C100-RI-051 |  | 12070130001 | CAVITY |  | C100-CAV-INSP2 | R3 | 59 | See attachment for dimensional non-conforming. | OGLE | 2011-09-09 |  |  |  | HOGAN | 2011-09-09 | The dimensional non-conformance of 0.01" from the FPC to the FP is usable as is. | Use As Is | [14 empty] | C100-RI-051 | OGLE | 2011-09-09 | 147F4F82-ED46-B3B9-A2FDBB1DA02A6AEC |  | 2011-09-09 | Use As Is
731 | F890195-07 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R6 | 137 | A large number of stains on copper plating inside waveguide. Flatness out of tol | GEORGED | 2011-09-09 |  |  |  | KWILSON |  |  | Use As Is | [14 empty] | F890195-07 | GEORGED | 2011-09-09 |  |  |  | Use As Is
732 | C100-RI-048 |  |  | CAVITY | Cavity does no meet field flatness specification in step 4 | C100-CAV-TUNE | R4 | 64 | Field flatness specification is +-5% or 10% total---cavity measures at +7.5%, - | FOREHAND | 2011-09-09 |  |  | Danny, I have started investigation of this NCR report including previous r | HAIPENG | 2011-09-13 | Danny, I have started investigation of this NCR report including previous r | Remeasure | [14 empty] | C100-RI-048 | FOREHAND | 2011-09-09 |  |  |  | Use As Is | Please read TN-10-021 technote. Handing this cavity with care and continue the r
733 | F067763_07 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-RWRK-WGD | R3 | 19 | Waveguide leaked in the cleanroom. The knife edge needs rework done. Please see attached PDF for history of leaks. | GEORGED | 2011-09-12 |  |  |  | GEORGED | 2012-01-13 | 1-13-12 Please rework as needed. | Modify | [14 empty] | F067763_07 | GEORGED | 2011-09-12 | B2C239D2-B3F5-20E1-CB213C9F2216C408 |  |  | Modify
734 | C100-RI-048 |  | 12070200000-INSP | HELIUM VESSEL FINAL WELD |  | C100-CAV-INST-HELV | R2 | 49 | See attachment for dimensional non-conforming. | OGLE | 2011-09-12 |  |  |  | FISCHER | 2011-09-13 |  | Use As Is | [14 empty] | C100-RI-048 | OGLE | 2011-09-12 | B785C3F9-0282-0737-097A5DAD70F4CE1C |  | 2011-09-13 | Use As Is
735 | C100_RI_045 |  |  |  |  | C100-CAV-VTRF | R3 | 89 | No S21 signla between FPC and FP connectors while at 2.07 K. FPC cable not conne |  | 2011-09-12 |  |  |  | HAIPENG | 2011-09-13 | Check FPC to other HOM ports if FP and HOM cables swapped. If the HOM signals a | Remeasure | [14 empty] | C100_RI_045 |  | 2011-09-12 |  |  | 2011-09-13 | Remeasure
736 | C100-RI-027 |  |  | Helium Vessel | Bellows assembly was repaired. Assembled into cavity string, bellows on HOM end | C100-CAV-INST-HELV | R1 | 24 | Bellows leaks, string will be disassembled and returned to cavity group for rebu | FISCHER | 2011-09-13 |  |  | Leaky helium vessel bellows. | FISCHER | 2011-09-13 | unit must be repaired for use. | Modify | [14 empty] | C100-RI-027 | FISCHER | 2011-09-13 |  |  | 2011-09-13 | Modify
737 | C100-RI-052 |  | 12070130001 | CAVITY |  | C100-CAV-INSP2 | R3 | 60 | See attachment for dimensional non conforming. | OGLE | 2011-09-13 |  |  |  | HOGAN | 2011-09-13 | The .01" out of tolerance for the overall length of the cavity is acceptable for | Use As Is | [14 empty] | C100-RI-052 | OGLE | 2011-09-13 | E0B80C24-AB49-53BF-FED51ADFF2F801FA |  | 2011-09-13 | Use As Is
738 | F892523-03 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R6 | 132 | Surface finish out of tolerance. Scratches on the small flange. | GEORGED | 2011-09-13 |  |  |  | KWILSON |  | Please polish to repair scratches and surface finish. | Modify | [14 empty] | F892523-03 | GEORGED | 2011-09-13 |  |  |  | Modify
739 | F892522-08 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R6 | 135 | Copper flaking inside the bellows. Please see pictures attached. | GEORGED | 2011-09-13 |  |  |  | KWILSON |  |  | Return To Vendor | [14 empty] | F892522-08 | GEORGED | 2011-09-13 | E8CE0721-C942-38F3-01CDEB21E0073FCB |  |  | Return To Vendor
740 | FLG0037-149 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 140 | Damaged knife edge. See attachment for dimensional non conforming. | OGLE | 2011-09-14 |  |  |  | HKPARK | 2011-09-14 |  | Return To Vendor | [14 empty] | FLG0037-149 | OGLE | 2011-09-14 | 0E0D3F5D-BB53-97FD-0E694BD0EBDDB4A2 |  | 2011-09-14 | Return To Vendor
741 | FLG0037-150 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 141 | Damaged knife edge. | OGLE | 2011-09-14 |  |  |  | HKPARK | 2011-09-14 |  | Return To Vendor | [14 empty] | FLG0037-150 | OGLE | 2011-09-14 | 0E6BE25A-C73B-8448-5C2C1E4C1D024D51 |  | 2011-09-14 | Return To Vendor
742 | FLG0037-151 |  | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 142 | Damaged knife edge. | OGLE | 2011-09-14 |  |  |  | HKPARK | 2011-09-14 |  | Return To Vendor | [14 empty] | FLG0037-151 | OGLE | 2011-09-14 | 0E974567-D51E-A323-EE484B828C34CD75 |  | 2011-09-14 | Return To Vendor
743 | C100-RI-049 |  | 12070200000-INSP | HELIUM VESSEL FINAL WELD |  | C100-CAV-INST-HELV | R2 | 50 | See attachment for dimensional non-conforming. | OGLE | 2011-09-19 | 89607B1A-B87C-40C5-3CF5FE59AC59179B |  |  | ABURRILL | 2011-10-07 | This is a welded assembly, CM installation will take care of discrepancy. | Use As Is | [14 empty] | C100-RI-049 | OGLE | 2011-09-19 | 16DF3D58-91F3-7847-FF2A2B4767E98FC1
744 | C100-RI-054 |  | 12070130001 | CAVITY |  | C100-CAV-INSP2 | R3 | 62 | See attachment for dimensional non-conforming. | OGLE | 2011-09-19 |  |  |  | HOGAN | 2011-09-19 | The .014" out of tolerance on the FPC flange position and the .02" on the end fl | Use As Is | [14 empty] | C100-RI-054 | OGLE | 2011-09-19 | 2004AA01-F734-4ED4-B5478C421E062148 |  | 2011-09-19 | Use As Is
745 | F890195-07 |  | CRM12070700000 | C100 Waveguide |  | C100-CAV-RWRK-WGD | R4 | 81 | Scratches on the small flange | GEORGED | 2011-09-20 |  |  |  | KWILSON |  | Please polish flange face to remove scratches. | Modify | [14 empty] | F890195-07 | GEORGED | 2011-09-20 |  |  |  | Modify
746 | 48310-0052 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 183 | See attachment for dimensional non conforming. | OGLE | 2011-09-20 |  |  |  | HKPARK | 2011-09-21 | It was decided to use nominal +/-.010. | Use As Is | [14 empty] | 48310-0052 | OGLE | 2011-09-20 | 4DD8FE4E-AFDC-62A5-7A870EE86D2C05FF |  | 2011-09-21 | Use As Is
747 | 48310-0058 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 189 | See attachment for dimensional non conforming. | OGLE | 2011-09-20 |  |  |  | HKPARK | 2011-09-21 | It was decided to use nominal +/-.010. | Use As Is | [14 empty] | 48310-0058 | OGLE | 2011-09-20 | 4F5BB053-F467-215A-01B7CBD6EB7F6298 |  | 2011-09-21 | Use As Is
748 | 48310-0059 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 190 | See attachment for dimensional non conforming. | OGLE | 2011-09-20 |  |  |  | HKPARK | 2011-09-21 | It was decided to use nominal +/-.010. | Use As Is | [14 empty] | 48310-0059 | OGLE | 2011-09-20 | 5009A146-0158-2510-50FD13FECC44D4EC |  | 2011-09-21 | Use As Is
749 | 48310-0061 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 192 | See attachment for dimensional non conforming. | OGLE | 2011-09-20 |  |  |  | HKPARK | 2011-09-21 | It was decided to use nominal +/-.010. | Use As Is | [14 empty] | 48310-0061 | OGLE | 2011-09-20 | 513D339B-94B3-0FC7-E2307A34A383780C |  | 2011-09-21 | Use As Is
750 | 48310-0063 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 194 | See attachment for dimensional non conforming. | OGLE | 2011-09-20 |  |  |  | HKPARK | 2011-09-21 | It was decided to use nominal +/-.010. | Use As Is | [14 empty] | 48310-0063 | OGLE | 2011-09-20 | 543CCBC6-F962-2646-8C584015175F11BD |  | 2011-09-21 | Use As Is
751 | 48310-0064 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 195 | See attachment for dimensional nonconforming. | OGLE | 2011-09-21 |  |  |  | HKPARK | 2011-09-21 | It was decided to use nominal +/-.010. | Use As Is | [14 empty] | 48310-0064 | OGLE | 2011-09-21 | 775E058E-CDE8-73C4-A49BC927BE28C61E |  | 2011-09-21 | Use As Is
752 | [5 empty] | C100-CAV-VTRF | R3 | 95 | Cavity had low Q, at limit of specification. Strange behavior in field probe ch |  | 2011-09-21 |  |  |  | HAIPENG | 2011-09-21 | After talking to Mircea, Ari, Pete, Rongli, Andrew, and looking at the traveler | Remeasure | [16 empty] | 2011-09-21 |  |  | 2011-09-21 | Remeasure
753 | 48310-0065 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 196 | See attachment for dimensional non conforming. | OGLE | 2011-09-21 |  |  |  | HKPARK | 2011-09-21 | It was decided to use nominal +/-.010. | Use As Is | [14 empty] | 48310-0065 | OGLE | 2011-09-21 | 77E10DC6-AB0C-F5DF-064D1A4E45908077 |  | 2011-09-21 | Use As Is
754 | 48310-0066 |  | 12070140020 | HOM FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 197 | See attachment for dimensional non conforming. | OGLE | 2011-09-21 |  |  |  | HKPARK | 2011-09-21 | It was decided to use nominal +/-.010. | Use As Is | [14 empty] | 48310-0066 | OGLE | 2011-09-21 | 78224C5E-BD37-489F-AE8BC1232113ABE9 |  | 2011-09-21 | Use As Is
755 | 123 | MYSN |  |  | TESTING ONLY | C100-CAV-ASSY | R1 | 1 | TESTING | BOOKWALT | 2011-09-22 |  |  | TESTING | BOOKWALT | 2011-09-22 | TESTING | Proceed |  | TEST | BOOKWALT | 2011-09-22 |  | TEST | BOOKWALT | 2011-09-22 |  | BOOKWALT | 2011-09-22 | TEST CONPLETE | Reject
756 | C100-RI-043 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY | THIS IS ONLY A TEST OF THE NEW NCR REVISION | C100-CAV-INST-HELV | R1 | 36 | THIS IS ONLY A TEST. | BOOKWALT | 2011-09-22 |  |  | PLEASE DO NOT RESPOND. THIS IS ONLY A TEST | BOOKWALT | 2011-09-22 | NO RESPONSE PLEASE... THIS IS ONLY A TEST | Re-process |  | TESTING OF NCR SYSTEM | BOOKWALT | 2011-09-22 |  | TESTING ONLY. NO RE-PROCESSING NECESSARY | BOOKWALT | 2011-09-22 |  | BOOKWALT | 2011-09-22 | TESTING HAS BEEN COMPLETED. THANKS FOR IGNORING. | Reject
757 | C100-RI-051 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INST-HELV | R2 | 52 | See attachment for dimensional non-conforming. | OGLE | 2011-09-22 | C0644CDC-0B1C-397D-D2B1F077D7CB4021 |  |  | ABURRILL | 2011-09-26 | Dimensional deviation will be made up in cryomodule assembly | Use As Is
758 | C100-RI-056 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INSP2 | R3 | 63 | See attachment for dimensional non conforming. | OGLE | 2011-09-23 | DF735C09-B351-4799-10B8304A9A44224A |  |  | HOGAN | 2011-09-23 | The .003" out of tolerance on the FPC flange location can be accommodated at the next assembly. | Use As Is
759 | C100-RI-055 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INSP2 | R3 | 64 | See attachment for dimensional non-conforming. | OGLE | 2011-09-23 | E1F30A4E-EA0C-AE48-BFA7E1406BD1CC09 |  |  | ABURRILL | 2011-09-23 | Please proceed, dimensional issue will be addressed during assembly step. | Use As Is
760 | 48310-0067 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 198 | See attachment for dimensional non conforming. | OGLE | 2011-09-23 | E36600CF-0BE1-7F80-ED17F372C35AF705 |  |  | HKPARK | 2011-09-26 | The tip length acceptance tolerance is nominal +/- .010. | Use As Is
761 | 48310-0069 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 200 | See attachment for dimensional non conforming. | OGLE | 2011-09-23 | E37AFFBC-0EF1-982B-D65B06E9703709AF |  |  | HKPARK | 2011-09-26 | The tip length acceptance tolerance is nominal +/- .010. | Use As Is
762 | 48310-0072 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 202 | See attachment for dimensional non conforming. | OGLE | 2011-09-23 | E38D0E17-FA12-5D20-AD0B2B64681839AB |  |  | HKPARK | 2011-09-26 | The tip length acceptance tolerance is nominal +/- .010. | Use As Is
763 | 48310-0073 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 203 | See attachment for dimensional non conforming. | OGLE | 2011-09-23 | E39BC42B-D8FC-4599-04BF4856EF782A80 |  |  | HKPARK | 2011-09-26 | The tip length acceptance tolerance is nominal +/- .010. | Use As Is
764 | 48310-0075 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 205 | See attachment for dimensional non conforming. | OGLE | 2011-09-23 | E3AF75A9-F5E2-3385-603A9CAB48EEA56B |  |  | HKPARK | 2011-09-26 | The tip length acceptance tolerance is nominal +/- .010. | Use As Is
765 | 48310-0077 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 207 | Scratch in seal groove. See attachment for dimensional non conforming. | OGLE | 2011-09-23 | E3C72D1D-B08A-E948-FC0F7AB56D783084 |  |  | HKPARK | 2011-09-26 | The tip length acceptance tolerance is nominal +/- .010. Polish the seal surface and notify the responsible engineer the result. | Modify |  |  | OGLE | 2011-10-19 |  | Groove polished 10/19/11. | CARPENTE | 2011-10-19 |  | HKPARK | 2011-11-03 |  | Use As Is
766 | 48310-0078 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 208 | Braze material in seal groove. | OGLE | 2011-09-23 | E3EA6AA4-0029-A83A-50006D45CCD54C83 |  |  | HKPARK | 2011-09-26 | Remove the braze material by polishing or equivalent method. Notufy the responsible engineer of the result. | Modify |  |  | OGLE | 2011-10-19 |  | Groove re polished 10/19/11 JO. | OGLE | 2011-10-19 |  | HKPARK | 2011-11-03 |  | Use As Is
767 | 48310-0079 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 209 | See attachment for dimensional non conforming. | OGLE | 2011-09-23 | E3FE2CA4-A28C-F5DD-3AB91F43CC0ECF6C |  |  | HKPARK | 2011-09-26 | The tip length acceptance tolerance is nominal +/- .010. | Use As Is
768 | 48310-0081 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 211 | See attachment for dimensional non conforming. | OGLE | 2011-09-23 | E40C5EEC-C05F-3C31-1652022CF7E23E30 |  |  | HKPARK | 2011-09-26 | The tip length acceptance tolerance is nominal +/- .010. | Use As Is
769 | 48310-0082 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 212 | See attachment for dimensional non conforming. | OGLE | 2011-09-23 | E4181877-A95E-309F-4D368EFA53B94C3D |  |  | HKPARK | 2011-09-26 | The tip length acceptance tolerance is nominal +/- .010. | Use As Is
770 | C100-RI-052 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INST-HELV | R2 | 53 | See attachment for dimensional non conforming. | OGLE | 2011-09-23 | ED1B8290-0854-AFF6-9D27CEF7A2A818AA |  |  | ABURRILL | 2011-09-26 | dimensional issue will be made up in welding stage. | Use As Is
771 | C100_RI_048 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-VTRF | R3 | 99 | Cavity has a superfluid He leak. Radiation burst (7 R/h) at about 26 MV/m. He discharge when try to lock the cavity at low field. | GCIOVATI | 2011-09-23 |  |  |  | HAIPENG | 2011-09-23 | Warm up to do a careful leak check. If the leak is found at gasket location, tight up by torque wrench. After the leak proof check, remeasure it at VTA. If the leak is found at the weld location, undress the cavity for the EBW repair. | Re-process |  | Cavity sent back to lapping stage to remove scratches from FPC. | ABURRILL | 2011-10-04 |  | Cavity leak tight before VTA testing | ABURRILL | 2011-10-04 |  | ABURRILL | 2011-10-04 | Cavity ready for testing. | Use As Is
772 | C100-RI-057 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INSP2 | R3 | 65 | Discoloration on Field Probe Flange. See attachment for dimensional non conforming. | OGLE | 2011-09-26 | 7DFE36F0-EEEA-DB39-79BACF65768692EF |  |  | ABURRILL | 2011-09-26 | Dimensional tolerance is acceptable to use as is. Please scotchbrite the discoloration and re-assess prior to welding. | Modify |  |  | OGLE | 2011-09-26 |  | Rework complete. JO See attached photo. | OGLE | 2011-09-26 | 7F630641-081A-AEFE-A8BCE243387AB029 | ABURRILL | 2011-09-26 | Progress Cavity to next workcenter. | Use As Is
773 | C100-RI-053 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INST-HELV | R2 | 54 | See attachment for dimensional non conforming. | OGLE | 2011-09-26 | 851EA3C4-CE0E-49AC-AEDD1F9AD3D577E9 |  |  | ABURRILL | 2011-09-26 | Please proceed. Dimensional error will be accounted for in cryomodule assembly | Use As Is
775 | 48310-0075 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 205 | Unit it has a -7 range leak. | WILCOX | 2011-09-26 |  |  |  | HKPARK | 2011-09-27 |  | Return To Vendor
776 | C100-RI-058 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INSP2 | R3 | 66 | See attachment for dimensional non conforming. | OGLE | 2011-09-27 | B19E1F24-9B03-DCC4-A4F821ED660369E2 |  |  | ABURRILL | 2011-09-27 | Dimensional deviations will be made up for during cryomodule assembly step. | Use As Is
777 | C100-RI-054 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INST-HELV | R2 | 55 | See attachment for dimensional non conforming. | OGLE | 2011-09-27 | B8A1D39D-9857-14AB-627ADB88A654B882 |  |  | ABURRILL | 2011-10-07 | Dimensional non-conformance will be addressed at the cryomodule assembly level. | Use As Is
778 | F893562/01 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R6 | 130 | Flatness measured .006" | GEORGED | 2011-09-27 |  |  |  |  | 2011-09-28 |  | Use As Is
779 | ISI-0015 |  | CRM120-7014-0020 | HOM Feedthru |  |  |  |  | Antenne/center pin assembly broken away from sapphire. Discovered during cleaning for cavity assembly. | CARPENTE | 2011-09-28 |  |  |  | HKPARK | 2011-11-02 |  | Reject
780 | C100-RI-049 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY | Cavity developed a superfluid cold leak, high power RF test aborted. | C100-CAV-VTRF | R3 | 102 | Cavity developed a superfluid cold leak, high power RF test aborted. | ABURRILL | 2011-09-29 |  |  |  | ABURRILL | 2011-09-29 | Remove from dewar, attempt to locate cold leak, remove component that leaks, inspect flange surface, determine path forward. | Re-process | [9 empty] | ABURRILL | 2011-11-16 | cavity meets spec | Use As Is
781 | 0 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY | C100-CAv-VTRF-R3-103 C100-RI-051 | C100-CAV-VTRF | R3 | 103 | Since C100-RI-051 was limited below 25 MV/m, quench fields in other modes were measured. | HAIPENG | 2011-10-01 | 8387F0A3-9F60-A6EE-B113F258A7865A14 |  | warm it up and undress the cavity to do optical inspection on 1st and 7th cells. Try to find any defect related to the end-cell early quench. Don't need to do HOM measurement next time. | GRIGORY | 2011-10-01 | Oct 1 post processing data shows a passing cavity | Use As Is | [9 empty] | GRIGORY | 2011-10-01
782 | C100-RI-061 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INSP2 | R3 | 69 | See attachment for dimensional non conforming. | OGLE | 2011-10-03 | F7D11FC0-94E3-A29B-5FC2FD41CA6CB4C7 |  |  | ABURRILL | 2011-10-04 | dimensional deviation will be made up in cryomodule assembly. | Use As Is
783 | F063134/04 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R6 | 139 | Scratches on the small flange. | GEORGED | 2011-10-04 |  |  |  | KWILSON | 2011-10-20 | Please polish scratches. | Modify |  | Small flange re-machined. | GEORGED | 2011-12-09 |  | Visual done, flange okay. | GEORGED | 2011-12-09 |  | GEORGED | 2012-01-05 |  | Use As Is
784 | test | test | unkn | unkn |  | test |  |  |  | ABURRILL | 2011-10-05 |  |  |  | ABURRILL | 2011-11-16 |  | Reject
785 | FLG0037-153 | DSRFSN | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 144 | See attachment for dimensional non conforming. | OGLE | 2011-10-05 | 5E28ADC5-D15E-7B7D-533D15DA867DB383 |  |  | HKPARK | 2011-10-11 | The out tol. value won't affect assembly. | Use As Is
786 | FLG0037-156 | DSRFSN | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 147 | See attachment for dimensional non conforming. | OGLE | 2011-10-05 | 63D42FFD-EDC7-CC04-0B4FCF1C9A1AE042 |  |  | HKPARK | 2011-10-11 | Out tol. value is minor. it won't affect assembly. | Use As Is
787 | C100-RI-055 | CAVSN | 12070200000-INSP | HELIUM VESSEL ASSEMBLY |  | C100-CAV-INST-HELV | R2 | 57 | See attachment for dimensional non conforming. | OGLE | 2011-10-05 | 650E9181-C7F2-595C-72CD39EC32E275F7 |  |  | ABURRILL | 2011-10-06 | Dimensional deviation will be made up in CM assembly step. | Use As Is
788 | C100-RI-063 | CAVSN | unkn | unkn |  | C100-CAV-INSP2 | R3 | 70 | Field Probe flange heavily oxidized. | CARPENTE | 2011-10-06 | 7B6EF16F-A2A2-5CCE-0B92FDD0CE9516B7 |  |  | ABURRILL | 2011-10-06 | scotch-brite the flange and attempt to remove oxidation. Review post-cleaning to determine if further action is necessary. | Modify | [9 empty] | ABURRILL | 2011-10-24 |  | Use As Is
789 | C100-RI-064 | CAVSN | 12070130001 | CAVITY |  | C100-CAV-INSP2 | R3 | 71 | See attachment for dimensional non conforming. Oxidation on field probe. See attached photo. | OGLE | 2011-10-06 | 94A6F41B-C78A-B542-D6BC40604169A8A6 |  |  | ABURRILL | 2011-10-07 | Dimensional issue is ok and will be accounted for during assembly. Please have flange face cleaned with scotch-brite and re-inspect to ensure seal surface integrity | Modify |  |  | OGLE | 2011-10-10 |  | Field probe rework complete and accepted JO. | OGLE | 2011-10-10 | 58B2D2D8-CC09-93EA-FD2BED20CEA7046D | ABURRILL | 2011-10-20 |  | Use As Is
791 | C100-RI-053 | CMSN | unkn | unkn |  | C100-CAV-VTRF | R3 | 105 | Cavity developed a leak during cooldown. Vacuum level had reached 2e-6 Torr at ~70 cm of liquid in dewar. Fill and test aborted. | ABURRILL | 2011-10-07 |  |  |  | ABURRILL | 2011-10-24 | Cavity has been re-processed and meets spec. | Use As Is
792 | C100-RI-054 | CAVSN | unkn | unkn |  | C100-CAV-VTRF | R3 | 107 | Cavity developed a leak during cooldown. Vacuum level reached 2e-6 Torr with ~ 70 cm of liquid in the dewar. Fill and test aborted. | ABURRILL | 2011-10-07 |  |  |  | ABURRILL | 2011-10-07 | ID leak, inspect, replace defective component, re-rinse and return to dewar | Modify |  |  | ABURRILL | 2011-10-07 |  | Leak found in FPC during cold shock | ABURRILL | 2011-10-07 |  | ABURRILL | 2011-10-07 | visually inspect FPC flange, if no damage noted, re-install tophat, re-rinse and re-test. If damage found pass back to lapping step and begin again. | Re-process
793 | C100-RI-051 | CAVSN | unkn | unkn |  | C100-CAV-VTRF | R3 | 103 | Superfliud leak during before testing at 2 K. The cavity pressure was 2e-06 mbar at the beginning of the test. During the test the pressure reached high e-06 mabr. After the test the pressure was improved to low E-08. | GRIGORY | 2011-10-07 |  |  |  | ABURRILL | 2011-10-07 | Cavity meets specification, please place on shelf for use in string assembly. Leak ID at FPC during cold shock with LN2 | Use As Is
794 | C100-RI-056 | CAVSN | unkn | unkn |  | C100-CAV-INST-HELV | R2 | 56 | See attachment for dimensional non conforming. | OGLE | 2011-10-07 | C1C35737-EFC9-19A6-21AEE3ABA8B2A1BE |  |  | ABURRILL | 2011-10-07 | dimensional deviation will be made up in assembly process | Use As Is
795 | C100-RI-065 | CAVSN | 12070130001 | CAVITY |  | C100-CAV-INSP2 | R3 | 72 | See attachment for dimensional non conforming. | OGLE | 2011-10-07 | C89735B7-C8A1-6614-BB39D391FC46A39D |  |  | ABURRILL | 2011-10-07 | will be made up in assembly | Use As Is
796 | F897109/01 | WGDSN | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R6 | 143 | Scratches on the small flange. One bolt hole is bad on the large flange (outer pattern). NoGo gauge screws into the entire length of the threads. | GEORGED | 2011-10-10 |  |  |  | KWILSON | 2011-10-20 | Return to vendor - oversized bolt hole. REVISED: WG passed leak check, use as is. | Use As Is
797 | ASSEM0035-093 | WINSN | unkn | unkn |  | C100-CAV-INSP-ASSY-WIN | R3 | 95 | Damaged knife edge. Repaired by C. Wilcox under Mike McCrea's supervision 10-21-11 | WILCOX | 2011-10-10 | 688FFE1A-B234-0417-B5A72198D9469256 |  |  | HKPARK | 2011-10-24 |  | Use As Is
798 | C100-RI-066 | CAVSN | 12070130001 | CAVITY |  | C100-CAV-INSP2 | R3 | 73 | See attachment for dimensional non conforming. | OGLE | 2011-10-11 | 928DF69E-0FC4-62C8-61186DD53B172226 |  |  | ABURRILL | 2011-10-11 | not a problem | Use As Is
799 | FLG0037-167 | DSRFSN | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 158 | See attachment for dimensional non conforming. | OGLE | 2011-10-11 | 94162624-ABC0-4EFC-D9D801D9206E2C0A |  |  | HKPARK | 2011-10-12 | Out tol. value is small enough it won't affect the assembly. | Use As Is
800 | FLG0037-170 | DSRFSN | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 161 | See attachment for dimensional non conforming. | OGLE | 2011-10-11 | 9CA794DE-B704-A23A-B6D39C8B763E79B2 |  |  | HKPARK | 2011-10-12 |  | Return To Vendor
801 | C100-RI-009 | CAVSN | unkn | unkn |  | C100-CAV-INSP | R2 | 12 | Both HOM flanges need re-machining and polishing. Both flanges had un-acceptable seal surfaces at incoming inspection. The seal surfaces have had an attempted re-work by using an abrasive stone, and now need to be machined. | CARPENTE | 2011-10-12 | B0CFFAD1-99A2-1FC0-72477134159B85AD |  |  | ABURRILL | 2011-10-12 | Send cavity to machine shop along with C100-62 for re-machining. Contact Brian Carpenter or Scott Williams for details. | Modify |  | Both HOM flanges have been re-surfaced by the Jlab machine shop by lapping with a brass tool, using 30 uin. diamond lapping paste. | CARPENTE | 2011-12-05 |  | Flatness,depth, and surface finish of sealing surface inspected, all parameters ok. | CARPENTE | 2011-12-05 | 95A7EF43-F773-70C1-60B339B335D435D3 | ABURRILL | 2011-12-05 |  | Use As Is
802 | F893561/01 | WGDSN | CRM-120-7070-0000 | C100 Waveguide |  | C100-CAV-INSP-WGD | R6 | 144 | Nick on the knife edge. Scratches on the small flange. Flatness out of tolerance, measured .005" | GEORGED | 2011-10-12 |  |  |  | KWILSON | 2011-10-12 | Please polish scratches and repair knife edge. | Modify |  | Small flange polished. | CARPENTE | 2011-12-02 |  | Visual done, part okay. | GEORGED | 2011-12-02 | 967F3D4D-EC0A-D14E-5123AD2A692AA0A0 | KWILSON | 2011-12-05 |  | Use As Is
803 | FLG0037-176 | DSRFSN | 12070750037 | DOUBLE SIDED FLANGE |  | C100-CAV-INSP-FLNG | R2 | 167 | See attachment for dimensional non-conforming. | OGLE | 2011-10-12 | CB6CF18F-ACB1-1377-DA47684352CD84D2 |  |  | HKPARK | 2011-10-13 | The dimension out of tolerance is not critical. | Use As Is
804 | F892522/02 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R6 | 138 | Scratches on the small flange. Surface finish out of tolerance, meas. 26 micro inches. | GEORGED | 2011-10-14 |  |  |  | KWILSON | 2011-10-20 | Please repair scratches. | Modify |  | Used 268L 40 micron paper on entire flange with Dremel tool. Then used ScotchBrite 7447, and then 7448 to finish. | CASTAGNO | 2011-11-03 |  | Visual done, part okay. Flatness .004" Finish 12 micro inches. Waveguide Rework traveler done, seq num 85 | GEORGED | 2011-11-08 | 579F1DDD-018B-7268-2542BDA2D0001200 | GEORGED | 2011-11-09 | Small flange okay. | Use As Is
805 | C100-RI-060 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INST-HELV | R2 | 59 | See attachment | CARPENTE | 2011-10-14 | 22572565-B0E8-847B-4DEA919165C766B8 |  |  | CARPENTE | 2011-10-14 | Dimensional out of tolerance condition will be accommodated at cryomodule assembly. | Use As Is
806 | ASSEM0035-100 | WINSN | CRM1207075-0036 | WINDOW INSTALLATION |  | C100-CAV-INSP-ASSY-WIN | R3 | 87 | Knife edges are damaged. Repaired by C. Wilcox under Mike McCrea's supervision 10-21-11 | WILCOX | 2011-10-14 | 2441A016-A3ED-763D-8A90209486485747 |  |  | HKPARK | 2011-10-24 |  | Use As Is
807 | ASSEM0035-099 | WINSN | CRM1207075-0036 | WINDOW INSTALLATION |  | C100-CAV-INSP-ASSY-WIN | R3 | 96 | Knife edges are damaged. Repaired by C. Wilcox under Mike McCrea's supervision 10-21-11 | WILCOX | 2011-10-14 | 24545907-0AA7-6D64-6F942FBDC054E5FC |  |  | HKPARK | 2011-11-03 |  | Use As Is
808 | C100-RI-067 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INSP2 | R3 | 74 | See attachment for dimensional non conforming. | OGLE | 2011-10-14 | 2A88DB3D-05C7-5CFB-83455DB5BD846962 |  |  | ABURRILL | 2011-10-14 | dimensional deviation will be made up in cryomodule assembly. | Use As Is
809 | F890195/06 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R6 | 142 | A large amount of copper is flaking off face of the large flange. Raised edge on corner of large flange face. Scratches on the small flange. | GEORGED | 2011-10-18 | EC3375B3-90D5-A170-43AA1657B5374EB6 |  |  | KWILSON | 2011-11-01 | Please polish scratches. Other items ok. | Modify |  | Small flange polished. | CARPENTE | 2011-11-09 |  | Visual done, flange okay. Finish 6,8,6,8 Flatness .001" | GEORGED | 2011-11-09 |  | GEORGED | 2011-11-09 | Small flange okay. | Use As Is
810 | C100-RI-069 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INSP2 | R3 | 76 | See attachment for dimensional non conforming. Oxidation on field probe seal groove. | OGLE | 2011-10-18 | F94A0641-CC8C-89A3-40B54716873DB88D |  |  | ABURRILL | 2011-10-18 | Dimensional issue will be addressed on cryomodule assembly. Oxidation to be removed with scotchbrite and re-inspected | Modify |  |  | ABURRILL | 2011-10-20 |  | flange polished, ok to use | ABURRILL | 2011-10-20 |  | ABURRILL | 2011-10-20 |  | Use As Is
811 | C100-RI-061 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INST-HELV | R2 | 62 | See attachment for dimensional non conforming. | OGLE | 2011-10-18 | 03C49A03-079A-813D-54B7239D3A595700 |  |  | ABURRILL | 2011-10-19 | Dimensional deviation will be made up in cryomodule assembly. | Use As Is
812 | 48310-0083 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 213 | See attachment for dimensional non-conforming. | OGLE | 2011-10-18 | 06338609-9447-9807-7BBB2BD385A86640 |  |  | HKPARK | 2011-10-19 | The length withing the relaxed tolerance +/-.010. | Use As Is
813 | 1 |  | unkn | unkn | testing |  |  |  |  | MAJONES | 2011-10-19 |  |  |  | ABURRILL | 2011-11-14 |  | Use As Is
816 | 47726-0015 | HMFTSN | CRM120014-0020 | HOM Probe Feedthru |  | C100-CAV-INSP-HMFT | R2 | 7 | Center Pin off center. Antenna length out of tolerance. | CARPENTE | 2011-10-20 |  |  |  | CARPENTE | 2011-10-20 | Re-center pin. Antenna length ok. | Modify |  | Center Pin re-centered. Seal surfaced scotch-brighted. | CARPENTE | 2011-10-20 |  | Complete. | CARPENTE | 2011-10-20 |  | CARPENTE | 2011-10-20 |  | Use As Is
817 | 47726-0026 | HMFTSN | CRM120014-0020 | HOM Probe Feedthru |  | C100-CAV-INSP-HMFT | R2 | 20 | Center Pin off center. Antenna length out of tolerance. | CARPENTE | 2011-10-20 |  |  |  | CARPENTE | 2011-10-20 | Re-center pin. Antenna length ok. | Modify |  | Center Pin re-centered. Seal surfaced scotch-brighted. | CARPENTE | 2011-10-20 |  | Complete. | CARPENTE | 2011-10-20 |  | CARPENTE | 2011-10-20 |  | Use As Is
818 | 47726-0044 | HMFTSN | CRM120014-0020 | HOM Probe Feedthru |  | C100-CAV-INSP-HMFT | R3 | 35 | Center Pin off center. Antenna length out of tolerance. | CARPENTE | 2011-10-20 |  |  |  | CARPENTE | 2011-10-20 | Re-center pin. Antenna length ok. | Modify |  | Center Pin re-centered. Seal surfaced scotch-brighted. | CARPENTE | 2011-10-20 |  | Complete | CARPENTE | 2011-10-20 |  | CARPENTE | 2011-10-20 |  | Use As Is
819 | C100-070 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INSP2 | R3 | 77 | FPC Flange location out of spec. | CARPENTE | 2011-10-20 | 5CBA12E5-9D58-F02B-D08A3646487809C6 |  |  | ABURRILL | 2011-10-20 | FPC flange dimension deviation will be taken up in cryomodule assembly step. | Use As Is
820 | FLG0037-180 | DSRFSN | MDC 203046-2000 | DOUBLE SIDED WINDOW FLANGE |  | C100-CAV-INSP-FLNG | R2 | 171 | Damaged knife edge. | OGLE | 2011-10-20 | 5CCED8CA-EA15-AE9A-D5FF78375867D4AF |  |  | HKPARK | 2011-10-21 |  | Return To Vendor
821 | FLG0037-181 | DSRFSN | MDC 203046-2000 | DOUBLE SIDED WINDOW FLANGE |  | C100-CAV-INSP-FLNG | R2 | 172 | Damaged knife edge.See attachment for dimensional non conforming. | OGLE | 2011-10-20 | 5D558684-A5F2-DD53-75D7AA604577F962 |  |  | HKPARK | 2011-10-21 |  | Return To Vendor
822 | C100-RI-064 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INST-HELV | R2 | 64 | See attachment for dimensional non conforming. | OGLE | 2011-10-20 | 6169CD6B-0175-C9FA-420F33467D2178E4 |  |  | ABURRILL | 2011-10-20 | dimensional deviation acceptable. | Use As Is
823 | FLG0037-185 | DSRFSN | MDC 203046-2000 | DOUBLE SIDED WINDOW FLANGE |  | C100-CAV-INSP-FLNG | R2 | 176 | See attachment for dimensional non conforming. | OGLE | 2011-10-20 | 68EAE998-CA76-ADAB-78C24B0D011E53B6 |  |  | HKPARK | 2011-10-21 | Out tol. not critical. | Use As Is
824 | FLG0037-186 | DSRFSN | MDC 203046-2000 | DOUBLE SIDED WINDOW FLANGE |  | C100-CAV-INSP-FLNG | R2 | 177 | See attachment for dimensional non conforming. | OGLE | 2011-10-20 | 6A7403B4-BA1B-ADEE-502772626F7EA3CD |  |  | HKPARK | 2011-10-21 | Out tol. not critical. | Use As Is
825 | C100-RI-071 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INSP2 | R3 | 78 | See attachment for dimensional non conforming. | OGLE | 2011-10-21 | 9638F94F-A961-B8A3-33323FF9EB39906E |  |  | HAIPENG | 2011-10-22 | This discrepancy can be compensated by the cryomodule assembly process where the FPC true position is determined by the beam pipe end flange position. The deviation of the FPC position can be compensated by the waveguide bellows. | Use As Is
826 | 0 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-VTRF | R3 | 114 | C100-RI-058 with Qo at 19.2 MV/m bellow specifications | STIRBET | 2011-10-22 |  |  |  | ABURRILL | 2011-11-14 | follow-up testing by K. Davis shows cavity meets spec | Use As Is
827 | C100-RI-072 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INSP2 | R3 | 79 | See attachment for dimensional non conforming. | OGLE | 2011-10-24 | 34A96F07-DCFA-230B-57BE273776614EC1 |  |  | ABURRILL | 2011-10-24 | dimensional deviations are acceptable and will be made up in assembly process. | Use As Is
828 | C100-RI-065 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INST-HELV | R2 | 65 | See attachment for dimensional non conforming. | OGLE | 2011-10-25 | 61AF5583-DB43-7796-5F9C29AF5AEBB509 |  |  | ABURRILL | 2011-10-25 | dimensional deviation acceptable and will be made up in cryomodule assembly | Use As Is
829 | C100-RI-066 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INST-HELV | R2 | 66 | See attachment for dimensional non conforming. | OGLE | 2011-10-26 | 9742F948-095F-6FFE-E5923480CB9E4F6A |  |  | CARPENTE | 2011-10-27 | Dimensional deviation acceptable and will be made up in cryomodule assembly. | Use As Is
830 | ASSEM0035-122 | WINSN | CRM1207075-0036 | WINDOW INSTALLATION |  | C100-CAV-INSP-ASSY-WIN | R3 | 130 | Nick on knife edge. | GEORGED | 2011-10-27 | B9C13FAE-D0B3-743E-4BFF6374F342983B |  |  | HKPARK | 2011-10-31 |  | Reject
832 | C100-RI-074 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INSP2 | R3 | 82 | See attachment for dimensional non conforming. | OGLE | 2011-10-31 | 9BF2F1E8-9514-DD1B-B3D14B770CA22DCB |  |  | ABURRILL | 2011-11-03 | Dimensional deviations will be made up in cryomodule assembly | Use As Is
833 | ASSEM0035-099 | WINSN | CRM1207075-0036 | WINDOW INSTALLATION |  | C100-CAV-INSP-ASSY-WIN | R3 | 96 | Machining flat on knife edge | GEORGED | 2011-11-01 | C697561C-E110-E326-C342323C47860D6C |  |  | HKPARK | 2011-11-03 | Checked leak tight. | Use As Is
834 | C100-RI-075 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INSP2 | R3 | 83 | See attachment for dimensional non conforming. | OGLE | 2011-11-01 | D38871C0-F5F7-5CF8-8F6B81F80E100338 |  |  | ABURRILL | 2011-11-02 | dimensional deviation will be made up in assembly | Use As Is
835 | FLG0037-182 | DSRFSN | MDC 203046-2000 | DOUBLE SIDED WINDOW FLANGE |  | C100-CAV-INSP-FLNG | R2 | 173 | Damaged knife edge | OGLE | 2011-11-01 | D590BD2C-BB90-5621-547779D5BB4F0769 |  |  | HKPARK | 2011-11-02 |  | Return To Vendor
836 | C100-RI-069 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INST-HELV | R2 | 69 | Dimensional, see attachment | CARPENTE | 2011-11-04 | 60DEC3B4-BC64-B8BD-258E76F40EDF4D1E |  |  | ABURRILL | 2011-11-07 | Dimensional deviation will be made up in cryomodule assembly | Use As Is
837 | ASSEM0035-122 | WINSN | CRM1207075-0036 | WINDOW INSTALLATION |  | C100-CAV-INSP-ASSY-WIN | R3 | 130 | Unit only passed Leak check at 150 in./lbs. with brand new gasket | WILCOX | 2011-11-07 |  |  |  | HKPARK | 2011-11-07 |  | Reject
838 | C100-RI-076 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INSP2 | R3 | 84 | Please see attachment. | GEORGED | 2011-11-07 | F54AFABE-BA60-06A1-70D812FFC5B5939E |  |  | ABURRILL | 2011-11-07 | dimensional deviation will be made up in cryomodule assembly | Use As Is
839 | C100-077 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INSP2 | R3 | 85 | Dimensional, see attachment | CARPENTE | 2011-11-07 | FD061FD3-CEBE-E3C7-FCAC51F7BD40A687 |  |  | ABURRILL | 2011-11-07 | dimensional deviation will be made up in assembly process | Use As Is
840 | F067764-16 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R5 | 100 | Waveguide will only seal at 130 in/lbs | WILCOX | 2011-11-08 |  |  |  | KWILSON | 2011-11-14 |  | Use As Is
841 | C100-RI-086 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INSP | R3 | 91 | Void on Beam tube Flange-HOM End | CARPENTE | 2011-11-08 | 2787C2C3-0263-DABE-828452EA7BF5BABC |  |  | ABURRILL | 2011-11-08 | This type of damage has been noted before. It is good to track, but should not impact cavity performance. | Proceed | [9 empty] | ABURRILL | 2011-11-08 | visual inspection looks like several other successfully tested cavities. | Use As Is
842 | OML-0090 | WBASN | CRM1207075-0024 | BRAZEMENT ADAPTER ASSY |  | C100-CAV-INSP-ADPT | R3 | 117 | This adapter has been returned to the vendor for re-cleaning. RF test does not show a difference after the re-clean. | HKPARK | 2011-11-08 |  |  |  | HKPARK | 2011-11-08 |  | Reject
843 | ISI-0061 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 214 | Light scratch in seal groove. See attachment for dimensional non conforming. | OGLE | 2011-11-08 | 3FBB47A1-0E60-621F-C299C009558BA9EC |  |  | HKPARK | 2011-11-09 | Polish the seal surface if the scratch is under the gasket. The dimension is OK, within +/-.010. | Modify |  | Seal groove polished. | CARPENTE | 2011-11-18 |  | Visual done, seal groove okay. | GEORGED | 2011-11-21 |  | HKPARK | 2011-11-23 |  | Use As Is
844 | ISI-0062 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 215 | Light scratch in seal area.(May not reach the groove center) Spot on surface of sapphire insert. See attachment for dimensional non conforming. | OGLE | 2011-11-08 | 4017BD9A-C4BF-2E63-EA3789031E92F018 |  |  | HKPARK | 2011-11-09 | Polish the seal surface if the scratch is under the gasket. The angle is OK. After polish proceed. The spot on sapphire will be tested at cryocycle. | Modify |  | Seal groove polished. | CARPENTE | 2011-11-18 |  | Visual done, seal groove okay. | GEORGED | 2011-11-21 |  | HKPARK | 2011-11-23 |  | Use As Is
845 | ISI-0064 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 216 | See attachment for dimensional non conforming. | OGLE | 2011-11-08 | 403442CA-C663-C962-47410A0F2A5F17E6 |  |  | HKPARK | 2011-11-09 |  | Use As Is
846 | ISI-0065 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 217 | Scratch in seal groove. | OGLE | 2011-11-08 | 404B01E5-E665-38DB-8A3AF400F1B75F97 |  |  | HKPARK | 2011-11-09 | Polish the seal surface. | Modify |  | Seal groove polished. | CARPENTE | 2011-11-18 |  | Visual done, seal groove acceptable. | GEORGED | 2011-11-21 |  | HKPARK | 2011-11-23 |  | Use As Is
847 | ISI-0066 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 218 | Small scratch/gouge in seal groove. Braze material shows between probe and sapphire surface. | OGLE | 2011-11-08 | 405EA1EB-015A-1356-69F99BA56BA4D30D |  |  | HKPARK | 2011-11-09 | Polish the seal surface if the scratch is under the gasket. The effect of the braze material will be checked at RF test. | Modify |  | Seal groove polished. | CARPENTE | 2011-11-18 |  | Visual done, seal groove acceptable. | GEORGED | 2011-11-21 |  | HKPARK | 2011-11-23 |  | Use As Is
848 | ISI-0067 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 219 | Unusual spot on sapphire surface. Center connector pin very loose and ROTATES,however is attached. Braze material visible between probe and sapphire surface. See attachment for dimensional non conforming. | OGLE | 2011-11-08 | 40A0CC79-F735-F2A0-8444A52C186CC3D9 |  |  | HKPARK | 2011-11-09 |  | Return To Vendor
849 | ISI-0068 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 220 | Small scratch in seal area. Unusual spot on sapphire surface. See attachment for dimensional non conforming. | OGLE | 2011-11-08 | 40D22740-F66C-CAF5-412E1BCEECC2C543 |  |  | HKPARK | 2011-11-09 | The probe tip length is OK. Polish seal surface. The spot on the sapphire will be tested at cryo cycle and RF. | Modify |  | Seal groove polished. | CARPENTE | 2011-11-18 |  | Visual done, seal path okay. | GEORGED | 2011-11-21 |  | HKPARK | 2011-11-23 |  | Use As Is
850 | ISI-0069 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 221 | Small scratch in seal area. Small amount of braze material visible between probe and sapphire surface. | OGLE | 2011-11-08 | 40FF7F5D-D93D-9D44-C3C55578AF0FD4E9 |  |  | HKPARK | 2011-11-09 | Polish the seal surface if the scratch is under the gasket. The braze material is OK if it passes RF test. | Modify |  | Seal groove polished. | CARPENTE | 2011-11-18 |  | Visual done, seal groove okay. | GEORGED | 2011-11-21 |  | HKPARK | 2011-11-23 |  | Use As Is
851 | ISI-0070 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 222 | Braze material seen between probe and sapphire surface. See attachment for dimensional non conforming. | OGLE | 2011-11-08 | 4113BE09-B414-69AD-B79FD26A0DC91571 |  |  | HKPARK | 2011-11-09 | Dimentions acceptable. The braze material is OK if it passes RF test. proceed. | Use As Is
852 | ISI-0072 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 224 | See attachment for dimensional non conforming. | OGLE | 2011-11-08 | 413C5870-F26D-325A-7DF4BE45BB45815A |  |  | HKPARK | 2011-11-11 | The probe .001 short but it can be used. Tag the probe length. | Use As Is
853 | ISI-0073 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 225 | See attachment for dimensional non conforming. | OGLE | 2011-11-08 | 414A475D-F386-DB2B-AAC8908E589E199C |  |  | HKPARK | 2011-11-09 | Dimensional acceptable. | Use As Is
854 | ISI-0074 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 226 | See attachment for dimensional non conforming. | OGLE | 2011-11-08 | 41590013-E45B-AF2A-9894FE4CEC37C68B |  |  | HKPARK | 2011-11-09 | Angle acceptable | Use As Is
855 | ISI-0075 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 227 | See attachment for dimensional non conforming. | OGLE | 2011-11-08 | 4161FE1B-C8FA-2928-038245197C16A0F1 |  |  | HKPARK | 2011-11-09 | Dimensional acceptable | Use As Is
856 | ISI-0076 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 228 | See attachment for dimensional non conforming. | OGLE | 2011-11-08 | 41705DF7-0442-403F-AD416CB0FC9435EA |  |  | HKPARK | 2011-11-09 | Dimensional acceptable | Use As Is
857 | C100-RI-078 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INSP2 | R3 | 86 | See attachment for dimensional non conforming. | OGLE | 2011-11-10 | 98CC105C-D3C5-3DDB-5D5EE80E9C4D09E7 |  |  | ABURRILL | 2011-11-10 | dimensional issues will be dealt with in CM assembly | Use As Is
858 | C100-RI-072 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INST-HELV | R2 | 72 | See attachment for dimensional non conforming. | OGLE | 2011-11-10 | 9BAED944-DDF6-8E2D-9C3F223C9FA9EB78 |  |  | ABURRILL | 2011-11-11 | not a problem | Use As Is
859 | C100-RI-073 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INST-HELV | R2 | 73 | See attachment for dimensional non conforming. | OGLE | 2011-11-11 | D0E2B68E-C92D-CCE8-84DD0803E57A52EB |  |  | ABURRILL | 2011-11-14 | dimensional deviation will be made up in assembly | Use As Is
860 | OML-0125 | WBASN | CRM1207075-0024 | BRAZEMENT ADAPTER ASSY |  | C100-CAV-INSP-ADPT | R4 | 171 | See attachment for dimensional non conforming. | OGLE | 2011-11-14 | 6687AB65-9C7B-D5B5-67B60CF38D08ED98 |  |  | HKPARK | 2011-11-15 |  | Use As Is
861 | OML-0126 | WBASN | CRM1207075-0024 | BRAZEMENT ADAPTER ASSY |  | C100-CAV-INSP-ADPT | R4 | 172 | See attachment for dimensional non conforming. | OGLE | 2011-11-14 | 67750F2A-CF08-FA2E-3C8CA3BDEBA707B0 |  |  | HKPARK | 2011-11-15 | The radius has been changed and these dimensions are within spec. Ceramic thickness OK. | Use As Is
862 | OML-0127 | WBASN | CRM1207075-0024 | BRAZEMENT ADAPTER ASSY |  | C100-CAV-INSP-ADPT | R4 | 173 | See attachment for dimensional non conforming. | OGLE | 2011-11-14 | 67810F91-DDA5-2AF6-B0BBE5BCFEA58B8A |  |  | HKPARK | 2011-11-15 |  | Use As Is
863 | C100-RI-074 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INST-HELV | R2 | 74 | See attachment for dimensional non conforming. | OGLE | 2011-11-14 | 71A3DBBE-0698-B2C4-5E344E8D00D85963 |  |  | ABURRILL | 2011-11-15 | dimensional deviation will be made up in CM assembly | Use As Is
865 | ISI-0064 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 216 | Probe tip very loose, appears to be broken. | GEORGED | 2011-11-15 |  |  |  | HKPARK | 2011-11-15 |  | Return To Vendor
866 | C100-RI-080 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INSP2 | R3 | 88 | See attachment for dimensional non conforming. | OGLE | 2011-11-15 | A28E7356-D23C-4202-BA7B1F011DD29449 |  |  | ABURRILL | 2011-11-15 | dimensional deviation acceptable. | Use As Is
867 | C100-RI-081 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INSP2 | R3 | 89 | See attachment for dimensional non conforming. | OGLE | 2011-11-16 | CFA5328C-A87D-E710-BCF35876B830D476 |  |  | ABURRILL | 2011-11-16 | dimensional deviation will be made up in CM assembly | Use As Is
868 | C100-RI-076 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INST-HELV | R2 | 76 | See attachment for dimensional non conforming. | OGLE | 2011-11-16 | D29DF248-CF24-2567-7BBF8EF66547DE75 |  |  | ABURRILL | 2011-11-16 | dimensional deviation will be made up in CM assembly | Use As Is
869 | F063950/01 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-RWRK-WGD | R2 | 13 | The small flange needs to be re-machined. Small flange flatness .006" | GEORGED | 2011-11-17 |  |  |  | KWILSON | 2011-11-29 | Please polish flange face. | Modify |  | Work completed by the Machine Shop. | GEORGED | 2012-01-06 |  | Flatness is no longer out of tolerance, meas. .002" | GEORGED | 2012-01-06 |  | GEORGED | 2012-01-06 |  | Use As Is
870 | F058824/01 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R1 | 2 | Small ding on the bellows. Serpentine seal imprint on the small flange, it needs to be re-machined. | GEORGED | 2011-11-18 | 2AA0BFF1-AE95-6856-D5986119D7659E38 |  |  | KWILSON | 2012-04-18 |  | Modify |  | Work completed by the Machine Shop. | GEORGED | 2012-01-10 |  | Visual done, small flange is okay. | GEORGED | 2012-01-10 | CFAE95E4-02A0-CCB7-61809F004D7361B7 | KWILSON | 2012-04-18 |  | Reject
871 | F061359/13 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R2 | 12 | Serpentine seal imprint present on the small flange, it needs to be re-machined. | GEORGED | 2011-11-18 |  |  |  | KWILSON | 2011-11-18 | Please repair imprint. Update: Part not usable. Consider scrap. Please do not throw away. Might be useful for testing. | Reject |  | Work completed by the Machine Shop. | GEORGED | 2012-01-10 |  | Visual done, small flange is okay. | GEORGED | 2012-01-10 | CF8FACAB-9E2D-C240-BB8114491575D07C | GEORGED | 2012-01-10 |  | Use As Is
872 | ISI-0077 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 229 | Small scratch in seal groove. | OGLE | 2011-11-21 | CD891FAE-B7C7-7F9D-AC13282764790556 |  |  | HKPARK | 2011-11-21 | Polish if the scratch is located under the gasket. | Modify |  | Seal groove polished. | CARPENTE | 2011-12-16 |  | Visual done, seal groove okay. | GEORGED | 2011-12-16 |  | HKPARK | 2011-12-16 |  | Use As Is
873 | ISI-0078 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 230 | No visible gap between probe and sapphire surface. | OGLE | 2011-11-21 | CD985469-DB11-FBEB-8BFE380CCBC4E7E2 |  |  | HKPARK | 2011-11-21 | Cryo cycle test will determine if it stresses too much. | Proceed |  |  | HKPARK | 2011-12-14 |  | The unit passed cryo cycle test. | HKPARK | 2011-12-14 |  | OGLE | 2011-12-13 | Rejected for no gap between sapphire and probe. Disposition = Cryo cycle test will determine if it stresses too much. Part has passed Cryo cycle test without damaging sapphire | Use As Is
874 | ISI-0080 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 232 | Small scratch in seal groove. No visible gap between probe and sapphire surface. See attachment for dimensional non conforming. | OGLE | 2011-11-21 | CE17B6CB-C7C4-ECEA-1D39EAB757E41A48 |  |  | HKPARK | 2011-11-21 | Polish seal surface and proceed to the next step. Angle is acceptable as long as the probe tip is within the limit. The gap will be determined at cryo cycle test. | Modify |  | Seal groove polished. | CARPENTE | 2011-12-16 |  | Visual done, seal groove okay. | GEORGED | 2011-12-16 |  | HKPARK | 2011-12-16 |  | Use As Is
875 | ISI-0081 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 233 | Small scratch in seal area. No gap between probe and sapphire surface. Braze material seen inside and under sapphire surface. | OGLE | 2011-11-21 | CE444595-FBDC-AB71-0707D33BA6A40C02 |  |  | HKPARK | 2011-11-23 | The sapphire appears chipped. Not worth proceeding test. | Return To Vendor
876 | ISI-0082 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 234 | No visible gap between probe and sapphire surface. See attachment for dimensional non conforming. | OGLE | 2011-11-21 | CE64D660-A602-893B-6D861C39660062BE |  |  | HKPARK | 2011-11-21 | Dimention is acceptable. Cryocycle is to determine the adequacy of the gap. | Use As Is
877 | ISI-0083 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 235 | Braze spot in seal groove. See attachment for dimensional non conforming. | OGLE | 2011-11-21 | CE7DCD2B-AF86-4C5F-6465BECA9EF23524 |  |  | HKPARK | 2011-11-23 | Braze spot is OK if it is not under the gasket. Angle is acceptable. | Use As Is
878 | ISI-0085 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 236 | Small scratch in the seal groove. No gap between probe and sapphire surface. Bent probe. See attachment for dimensional non conforming. | OGLE | 2011-11-21 | CE914C8B-EC4E-01D5-3B5F10859A452C76 |  |  | HKPARK | 2011-11-21 |  | Return To Vendor
879 | FLG0037-195 | DSRFSN | MDC 203046-2000 | DOUBLE SIDED WINDOW FLANGE |  | C100-CAV-INSP-FLNG | R2 | 186 | Damaged knife edge. | OGLE | 2011-11-23 | 2ADAFC70-D8A2-4E64-4F56AEC966282399 |  |  | HKPARK | 2011-11-23 |  | Return To Vendor
880 | FLG0037-196 | DSRFSN | MDC 203046-2000 | DOUBLE SIDED WINDOW FLANGE |  | C100-CAV-INSP-FLNG | R2 | 187 | Damaged knife edge. | OGLE | 2011-11-23 | 2AE449BF-01EA-9CD7-AE511225675484C6 |  |  | HKPARK | 2011-11-23 |  | Return To Vendor
881 | F061359/11 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R2 | 14 | Waveguide leaked inside the cleanroom. The knife edge will need to be reworked. | GEORGED | 2011-11-28 |  |  |  | GEORGED | 2012-02-06 |  | Modify |  | Knife edge reworked by the machine shop. | GEORGED | 2012-02-06 |  | Visual okay. Knife edge height measured .016" | GEORGED | 2012-02-06 | 3ED2D23C-A9CF-6B8F-C7575F89E65870DB | GEORGED | 2012-02-06 |  | Use As Is
882 | F067764/08 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R6 | 111 | Waveguide leaked inside the cleanroom. The knife edge will need to be reworked. | GEORGED | 2011-11-28 |  |  |  | GEORGED | 2012-01-30 |  | Modify |  | Knife edge reworked by the machine shop. | GEORGED | 2012-01-30 |  | Visual okay, knife edge height measured .016" | GEORGED | 2012-01-30 | D5A34288-0B60-C7A7-0F45DD7F00037996 | GEORGED | 2012-01-30 |  | Use As Is
883 | F356223/04 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R6 | 129 | Waveguide leaked inside the cleanroom. The knife edge will need to be reworked. | GEORGED | 2011-11-28 |  |  |  | GEORGED | 2012-02-01 |  | Modify |  | Knife edge reworked by the machine shop. | GEORGED | 2012-02-01 |  | Rough edges and dings present on the knife edge. Height measured .015" UPDATE: knife edge was repaired with an abrasive stone and scotchbrite, see pics 6983 and 6984. BC 2/08/12 | GEORGED | 2012-02-13 | 45D005F4-0117-9434-E8E269A72A637A04 | GEORGED | 2012-02-13 |  | Use As Is
884 | C100-084 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INSP2 | R3 | 92 | FPC flange location out of tolerance | CARPENTE | 2011-11-29 | 5F047B6D-9CC5-EC4D-7805E6A652C0286D |  |  | ABURRILL | 2011-11-29 | dimensional deviation will be made up in CM assembly | Use As Is
885 | F067763/08 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R4 | 48 | Scratches on the small flange | GEORGED | 2011-11-29 |  |  |  | KWILSON | 2011-11-29 | Please polish flange face. | Modify |  | Small flange polished. | CASTAGNO | 2011-12-05 |  | Visual done, part okay. | GEORGED | 2011-12-05 | 97519A51-97A4-C60D-CE7B6E126C0AAF95 | KWILSON | 2011-12-05 |  | Use As Is
886 | C100-RI-077 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY | Helium port angle out of tolerance. | C100-CAV-INST-HELV | R2 | 77 |  | CARPENTE | 2011-11-29 | 69F7BCE1-F978-DB99-0585F3406AA81BA2 |  |  | ABURRILL | 2011-11-29 | dimensional deviation will be made up in CM assembly | Use As Is
887 | F892523/01 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R6 | 146 | Scratches on the small flange. | GEORGED | 2011-11-30 |  |  |  | KWILSON | 2011-11-30 | Please polish small flange | Modify |  | Small flange polished. | CASTAGNO | 2011-12-05 |  | Visual done, part okay. | GEORGED | 2011-12-05 | 9735FCC8-0AB5-2B60-6B3D2A3102608670 | KWILSON | 2011-12-05 |  | Use As Is
888 | F892523/04 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R6 | 145 | Scratches on the small flange. | GEORGED | 2011-11-30 |  |  |  | KWILSON | 2011-11-30 | Please polish! | Modify |  | Small flange polished. | CASTAGNO | 2011-12-05 |  | Visual done, part okay. | GEORGED | 2011-12-05 | 9764F19C-DFE5-E1A2-308A9DDCF7E13531 | KWILSON | 2011-12-05 |  | Use As Is
889 | C100-RI-078 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INST-HELV | R2 | 78 | See attachment for dimensional non conforming. | OGLE | 2011-11-30 | A3426F39-D638-9159-10CDB449E1A5439D |  |  | ABURRILL | 2011-11-30 | dimensional deviations will be made up in assembly | Use As Is
890 | C100-RI-062 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INSP2 | R3 | 80 | Please see attachment. | GEORGED | 2011-12-01 | D135E1E6-A6BE-8EC4-70F5DCD46F596D5F |  |  | ABURRILL | 2011-12-02 | dimensional issue will be made up in CM assembly | Use As Is
891 | C100-RI-080 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-TUNE | R4 | 98 | C100-RI-080 does not pass field flatness specification after HV was welded. The spec. is +-5% with a normalized field. The cavity measures -4.8%, +5.1%----total 9.9%(90.1). | FOREHAND | 2011-12-01 | D81E84F5-B796-D8C2-DE8E3C1B08F654A4 |  |  | ABURRILL | 2011-12-02 | Please proceed | Use As Is
892 | ASSEM0035-089 | WINSN | CRM1207075-0035 | ADAPTER WINDOW ASSEMBLY |  | C100-CAV-INSP-ASSY-WIN | R3 | 94 | Window stabilized at 194 Deg.C holding on to Window until further notice. | CARPENTE | 2011-12-02 |  |  |  | HKPARK | 2011-12-02 | Tag and qurantine. | Reject
893 | C100-RI-081 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INST-HELV | R2 | 81 |  | CARPENTE | 2011-12-06 | C7F1D5AD-AFB6-D94C-231E669B6EA89E0B |  |  | ABURRILL | 2011-12-06 | deviations will be made up in CM assembly | Use As Is
894 | C100-RI-082 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INST-HELV | R2 | 82 | See attachment for dimensional non conforming. | OGLE | 2011-12-07 | 0A73964D-AA71-E622-01DA893049CE2FB6 |  |  | CARPENTE | 2011-12-08 | Dimensional deviation will be accommodated at cryomodule assembly. | Use As Is
895 | FLG0037-200 | DSRFSN | MDC 203046-2000 | DOUBLE SIDED WINDOW FLANGE |  | C100-CAV-INSP-FLNG | R2 | 191 | See attachment for dimensional non conforming. | OGLE | 2011-12-09 | 7C4988AC-04E3-B6B8-C7EF104A27859008 |  |  | HKPARK | 2011-12-13 | Bore width is OK as long as it mates with the adapter. Please check with adapter and also check with wave guide flange to see if the hole location is OK. If it does not show a problem on assembly it will be used as is. | Proceed |  |  | OGLE | 2011-12-13 |  | Checked window opening and bolt pattern and will assemble with mating parts. JO | OGLE | 2011-12-13 |  | HKPARK | 2011-12-14 |  | Use As Is
896 | F897109/01 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R6 | 143 | Waveguide leaked in the cleanroom. The knife edge will need rework done. | GEORGED | 2011-12-12 |  |  |  | KWILSON | 2011-12-12 | Please repair knife edge. | Modify |  | Knife edge reworked by the machine shop. | GEORGED | 2012-02-10 |  | Visual inspection okay. Knife edge height measures .018" | GEORGED | 2012-02-10 | 0C799594-F78D-45BB-7989C9E9217C76DD | GEORGED | 2012-02-10 |  | Use As Is
897 | F892523/01 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R6 | 146 | Waveguide leaked in the cleanroom. The knife edge will need rework done. | GEORGED | 2011-12-12 |  |  |  | KWILSON | 2011-12-12 | Please repair knife edge | Modify |  | Knife edge reworked by the machine shop. | GEORGED | 2012-02-08 |  | Visual acceptable. Knife edge height measured .018" | GEORGED | 2012-02-08 | A2FAC495-E216-EE25-ACA617A441190FEE | GEORGED | 2012-02-08 |  | Use As Is
898 | C100-RI-083 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INST-HELV | R2 | 83 | See attachment for dimensional non conforming. | OGLE | 2011-12-12 | 0F37278F-01EB-EEC3-A0A295566174B782 |  |  | HOGAN | 2011-12-13 | Reworking this assembly would do more harm than good. We will need to accommodate this offset during the cold mass assembly. | Use As Is
899 | F061584/03 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R3 | 35 | Nick on the knife edge. | GEORGED | 2011-12-13 | 39040B98-A4EE-DC06-0C75AFD04FA55DFC |  |  | CARPENTE | 2012-01-11 | repair knife edge | Modify |  | Nick on knife edge repaired using an abrasive stone. | CARPENTE | 2012-01-11 |  | Successful | CARPENTE | 2012-01-11 |  | CARPENTE | 2012-01-11 |  | Use As Is
900 | F892522/04 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R6 | 128 | Copper flaking inside the bellows. | GEORGED | 2011-12-13 | 3C067DDE-F424-72CC-F710D9A03386543E |  |  | KWILSON | 2012-04-18 | ** Please put note on tag that part may be used if we run out of better parts, so do not dispose of it. Thanks. | Reject
901 | FLG0037-201 | DSRFSN | MDC 203046-2000 | DOUBLE SIDED WINDOW FLANGE |  | C100-CAV-INSP-FLNG | R2 | 192 | See attachment for dimensional non conforming. | OGLE | 2011-12-13 | 428B9DED-F148-D709-6952455B79BA9691 |  |  | HKPARK | 2011-12-14 | Check with the mating flange. | Proceed |  |  | OGLE | 2011-12-14 |  | Checks ok with mating flange. | OGLE | 2011-12-14 |  | HKPARK | 2011-12-15 |  | Use As Is
902 | C100-RI-084 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INST-HELV | R2 | 84 | See attachment for dimensional non conforming. | OGLE | 2011-12-14 | 715E4037-BA7B-8D38-538D1DF009F523B3 |  |  | CARPENTE | 2011-12-15 | Dimensional deviation will be made up at cryomodule assembly. | Use As Is
903 | C100-RI-079 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-VTRF | R4 | 136 | Vertical test failed due to FE loading, quenching at 21.5 MV/m. RF Processing was performed with no improvement. | KDAVIS | 2011-12-10 |  |  |  | ABURRILL | 2012-01-20 | re-rinsed and passed | Use As Is
904 | OML-0117 | WBASN | CRM1207075-0024 | BRAZEMENT ADAPTER ASSY |  | C100-CAV-INSP-ADPT | R3 | 143 | RF test temperature too high 164 deg C. | HKPARK | 2011-12-16 |  |  |  | HKPARK | 2011-12-16 | Quarantine in inventory area | Reject
905 | C100-RI-078 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-VTRF | R4 | 142 | Qualification test aborted due to strong Fe and Qo at low cavity fields (bellow 10 MV/m). Seems this cavity had vacuum issues in previous attempted tests. On 17-Dec during HOM survey at 2.07K I notice cavity vacuum fluctuations 5.5 e-8 mbar to 6.3 e-8 mbar | STIRBET | 2011-12-19 |  |  |  | ABURRILL | 2012-01-20 | re-rinsed and passed | Use As Is
906 | ASSEM0035-151 | WINSN | CRM1207075-0035 | ADAPTER WINDOW ASSEMBLY |  | C100-CAV-INSP-ASSY-WIN | R3 | 139 | Ceramic temperature exceeded 190 C which is the maximum temp limit for completed window assembly. | DRURY | 2012-01-02 |  |  |  | HKPARK | 2012-12-06 | Hold but keep it.-original response Jan 10, 2012 rejected. Decided to use as is after the window temperature monitoring at the cryomodule. | Use As Is
907 | ASSEM0035-152 | WINSN | CRM1207075-0035 | ADAPTER WINDOW ASSEMBLY |  | C100-CAV-INSP-ASSY-WIN | R3 | 137 | Ceramic Temperature exceeded maximum temperature limit for completed window assemblies. | DRURY | 2012-01-02 |  |  |  | HKPARK | 2012-01-10 | hold but keep it.-1/10/2012 use as is since the window temerature measurement at cryomodule(test cave) showed the RF windows run much lower temperature than the window only RF test. | Use As Is
908 | ASSEM0035-150 | WINSN | CRM1207075-0035 | ADAPTER WINDOW ASSEMBLY |  | C100-CAV-INSP-ASSY-WIN | R3 | 145 | Ceramic Temperature exceeded maximum limit during RF test. | DRURY | 2012-01-02 |  |  |  | HKPARK | 2012-01-10 | Hold but keep it.-1/10/2012 use as is since the window temerature measurement at cryomodule(test cave) showed the RF windows run much lower temperature than the window only RF test. | Use As Is
909 | FLG0037-206 | DSRFSN | MDC 203046-2000 | DOUBLE SIDED WINDOW FLANGE |  | C100-CAV-INSP-FLNG | R2 | 197 | See attachment for dimensional non conforming. | OGLE | 2012-01-02 | 50741362-E143-B49A-6B622B27C59BBA05 |  |  | HKPARK | 2012-01-10 | Out-tol values are acceptable. | Use As Is
910 | C100-RI-086 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INSP2 | R3 | 94 | See attachment for dimensional non conforming. | OGLE | 2012-01-03 | 7DAA092F-B49D-B68A-15D200A3BE53CBDD |  |  | HOGAN | 2012-01-04 | The end flanges are out of tolerance but can be accommodated during the cryomodule assembly process. | Use As Is
911 | FLG0037-210 | DSRFSN | MDC 203046-2000 | DOUBLE SIDED WINDOW FLANGE |  | C100-CAV-INSP-FLNG | R2 | 201 | See attachment for dimensional non conforming. | OGLE | 2012-01-03 | 83113D15-BC33-C0D5-CDCEEC1B82205002 |  |  | HKPARK | 2012-01-10 | Out-told values are acceptable but verify with mating flange. | Use As Is
912 | FLG0037-212 | DSRFSN | MDC 203046-2000 | DOUBLE SIDED WINDOW FLANGE |  | C100-CAV-INSP-FLNG | R2 | 203 | See attachment for dimensional non conforming. | OGLE | 2012-01-03 | 84FF807D-A20D-A895-940C316FD2D21A5D |  |  | HKPARK | 2012-01-10 |  | Use As Is
913 | JLAB-69 | WBASN | CRM1207075-0024 | BRAZEMENT ADAPTER ASSY |  | C100-CAV-INSP-ADPT | R4 | 182 | Unit failed leak check. Verified by Scott Williams | WILCOX | 2012-01-04 |  |  |  | HKPARK | 2012-01-10 |  | Reject
914 | JLAB-64 | WBASN | CRM1207075-0024 | BRAZEMENT ADAPTER ASSY |  | C100-CAV-INSP-ADPT | R4 | 177 | Unit failed leak check. verified by scott williams | WILCOX | 2012-01-04 |  |  |  | HKPARK | 2012-01-10 |  | Reject
915 | C100-RI-085 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INSP2 | R3 | 93 | See attachment for dimensional non conforming. | OGLE | 2012-01-04 | A97D8BAF-074C-0797-D93D8F58C1CE9F15 |  |  | HOGAN | 2012-01-10 | The end flanges are out of tolerance but can be accommodated during the cryomodule assembly process. | Use As Is
916 | FLG0037-216 | DSRFSN | MDC 203046-2000 | DOUBLE SIDED WINDOW FLANGE |  | C100-CAV-INSP-FLNG | R2 | 207 | See attachment for dimensional non conforming. Will assemble with mating part although some positions snug. Pocket accepts sample adapter. JO | OGLE | 2012-01-04 | B7AC349E-94C8-38F1-E27123F7FD3A2DFC |  |  | HKPARK | 2012-01-11 |  | Use As Is
917 | F061584/02 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R3 | 39 | Small flange needs to be re-machined. | GEORGED | 2012-01-05 |  |  |  | GEORGED | 2012-01-05 |  | Modify |  | Work completed by Machine Shop. | GEORGED | 2012-01-05 |  | Visual done. Flange okay. | GEORGED | 2012-01-05 | CDBDE9A8-C2E7-D86E-659842D5B20A4CD9 | GEORGED | 2012-01-05 |  | Use As Is
918 | F058824/03 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R1 | 4 | Small flange needs to be re-machined. | GEORGED | 2012-01-05 |  |  |  | GEORGED | 2012-01-05 |  | Modify |  | Machine Shop work complete. | GEORGED | 2012-01-05 |  | Visual done. Small flange is okay. | GEORGED | 2012-01-05 | D9FEF3C5-D2D5-B4FF-9DF80027A81D0F60 | GEORGED | 2012-01-05 |  | Use As Is
919 | F063951/02 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R3 | 28 | Small flange needs to be re-machined. | GEORGED | 2012-01-06 |  |  |  | GEORGED | 2012-01-06 |  | Modify |  | Work completed by the Machine Shop. | GEORGED | 2012-01-06 |  | Visual done and small flange is okay. | GEORGED | 2012-01-06 | 01539F9A-FE01-EB51-8C35308C66054566 | GEORGED | 2012-01-06 |  | Use As Is
920 | F067764/02 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R4 | 57 | Small flange needs to be re-machined. | GEORGED | 2012-01-09 |  |  |  | GEORGED | 2012-01-09 |  | Modify |  | Work completed by Machine Shop. | GEORGED | 2012-01-09 |  | Visual done, small flange okay. | GEORGED | 2012-01-09 |  | GEORGED | 2012-01-09 |  | Use As Is
921 | F063951/01 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R3 | 20 | Small flange needs to be re-machined. | GEORGED | 2012-01-09 |  |  |  | GEORGED | 2012-01-09 |  | Modify |  | Work completed by the Machine Shop. | GEORGED | 2012-01-09 |  | Visual done, small flange okay. | GEORGED | 2012-01-09 |  | GEORGED | 2012-01-09 |  | Use As Is
922 | ISI-0084 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 237 | See attachment for dimensional non conforming. | OGLE | 2012-01-09 | B436E68F-96F6-AC8E-15C5DEF09E5B7E5B |  |  | HKPARK | 2012-01-10 | Before returning to vendor, contact Danny Forehand to see if the long probe can be used. | Return To Vendor
923 | ISI-0086 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 238 | See attachment for dimensional non conforming. Probe slightly bent and has tool marks on tip. | OGLE | 2012-01-09 | B65BF5D6-EC85-0D87-F926EF8B9232AB96 |  |  | HKPARK | 2012-01-10 |  | Return To Vendor
924 | F063951/07 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R3 | 24 | Small flange needs to be re-machined. | GEORGED | 2012-01-10 |  |  |  | GEORGED | 2012-01-10 |  | Modify |  | Work completed by the Machine Shop. | GEORGED | 2012-01-10 |  | Visual done, small flange okay. | GEORGED | 2012-01-10 | DB274EC2-EF77-E851-80A828B4BFD7920D | GEORGED | 2012-01-10 |  | Use As Is
925 | C100-RI-009 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INSP2 | R3 | 95 | See attachment for dimensional non conforming. | OGLE | 2012-01-11 | 1CD4F6E9-C616-9736-060689EA30025404 |  |  | HOGAN | 2012-01-12 | The end flanges are out of tolerance but can be accommodated during the cryomodule assembly process. | Use As Is
926 | JLAB-65 | WBASN | CRM1207075-0024 | BRAZEMENT ADAPTER ASSY |  | C100-CAV-INSP-ADPT | R4 | 178 | See attachment for dimensional non conforming. Note: For radius out of tolerance condition part tested with sample flange and will fit together. | OGLE | 2012-01-13 | 86393DE6-07D8-DC81-755E85A40A2A4594 |  |  | HKPARK | 2012-01-17 | Drawing has been updated to have the mentioned radius .688. The unit satisfies the specification. | Use As Is
927 | JLAB-68 | WBASN | CRM1207075-0024 | BRAZEMENT ADAPTER ASSY |  | C100-CAV-INSP-ADPT | R4 | 181 | See attachment for dimensional non conforming. Note: For radius out of tolerance condition part tested with sample flange and will fit together. | OGLE | 2012-01-13 | 8647BFB7-C34E-89AC-F3B17E3493949D24 |  |  | HKPARK | 2012-01-17 | Drawing has been updated to have the mentioned radius .688. The unit satisfies the specification. | Use As Is
928 | JLAB-62 | WBASN | CRM1207075-0024 | BRAZEMENT ADAPTER ASSY |  | C100-CAV-INSP-ADPT | R4 | 175 | See attachment for dimensional non conforming. Note: For radius out of tolerance condition part tested with sample flange and will fit together. | OGLE | 2012-01-13 | 8651FDB8-AEDB-AA73-11F2F49246E946A6 |  |  | HKPARK | 2012-01-17 | Drawing has been updated to have the mentioned radius .688. The unit satisfies the specification. | Use As Is
929 | JLAB-70 | WBASN | CRM1207075-0024 | BRAZEMENT ADAPTER ASSY |  | C100-CAV-INSP-ADPT | R4 | 183 | See attachment for dimensional non conforming. Note: For radius out of tolerance condition part tested with sample flange and will fit together. | OGLE | 2012-01-13 | 8659AB89-DEA8-D939-2AEDE36305E9CF52 |  |  | HKPARK | 2012-01-17 | Drawing has been updated to have the mentioned radius .688. The unit satisfies the specification. The ceramic thinkness OK. | Use As Is
930 | JLAB-61 | WBASN | CRM1207075-0024 | BRAZEMENT ADAPTER ASSY |  | C100-CAV-INSP-ADPT | R4 | 174 | See attachment for dimensional non conforming. Note: For radius out of tolerance condition part tested with sample flange and will fit together. | OGLE | 2012-01-13 | 8662A9B7-D75D-EA50-E8664F7EE4B194D4 |  |  | HKPARK | 2012-01-17 | Drawing has been updated to have the mentioned radius .688. The unit satisfies the specification. The ceramic thikness OK. | Use As Is
931 | JLAB-67 | WBASN | CRM1207075-0024 | BRAZEMENT ADAPTER ASSY |  | C100-CAV-INSP-ADPT | R4 | 180 | See attachment for dimensional non conforming. Note: For radius out of tolerance condition part tested with sample flange and will fit together. | OGLE | 2012-01-13 | 866B24F9-A5B4-AD75-8B4F15326A21228D |  |  | HKPARK | 2012-01-17 | Drawing has been updated to have the mentioned radius .688. The unit satisfies the specification. Ceramic thickness OK. | Use As Is
932 | JLAB-63 | WBASN | CRM1207075-0024 | BRAZEMENT ADAPTER ASSY |  | C100-CAV-INSP-ADPT | R4 | 176 | See attachment for dimensional non conforming. Note: For radius out of tolerance condition part tested with sample flange and will fit together. | OGLE | 2012-01-13 | 86736AAA-F47F-A724-60F3A0A0A907A3AC |  |  | HKPARK | 2012-01-17 | Drawing has been updated to have the mentioned radius .688. The unit satisfies the specification. Ceramic thickness OK. | Use As Is
933 | JLAB-83 | WBASN | CRM1207075-0024 | BRAZEMENT ADAPTER ASSY |  | C100-CAV-INSP-ADPT | R4 | 196 | See attachment for dimensional non conforming. Note: For radius out of tolerance condition part tested with sample flange and will fit together. | OGLE | 2012-01-13 | 86799491-B1BB-E7DE-EA0ABC3E7F98A086 |  |  | HKPARK | 2012-01-17 | Drawing has been updated to have the mentioned radius .688. The unit satisfies the specification. Ceramic thickness OK. | Use As Is
934 | JLAB-82 | WBASN | CRM1207075-0024 | BRAZEMENT ADAPTER ASSY |  | C100-CAV-INSP-ADPT | R4 | 195 | See attachment for dimensional non conforming. Note: For radius out of tolerance condition part tested with sample flange and will fit together. | OGLE | 2012-01-13 | 868009D5-CC84-A27A-1830E7EC091A4C1F |  |  | HKPARK | 2012-01-17 | Drawing has been updated to have the mentioned radius .688. The unit satisfies the specification. Ceramic thickness OK. | Use As Is
935 | JLAB-66 | WBASN | CRM1207075-0024 | BRAZEMENT ADAPTER ASSY |  | C100-CAV-INSP-ADPT | R4 | 179 | See attachment for dimensional non conforming. Note: For radius out of tolerance condition part tested with sample flange and will fit together. | OGLE | 2012-01-13 | 86935434-F7BF-C4B3-505C62DB89F6ECD6 |  |  | HKPARK | 2012-01-17 | Drawing has been updated to have the mentioned radius .688. The unit satisfies the specification. Ceramic thickness OK. | Use As Is
936 | JLAB-72 | WBASN | CRM1207075-0024 | BRAZEMENT ADAPTER ASSY |  | C100-CAV-INSP-ADPT | R4 | 185 | See attachment for dimensional non conforming. Note: For radius out of tolerance condition part tested with sample flange and will fit together. | OGLE | 2012-01-13 | 8699C261-A843-5BF4-FE28AF096B9B3FB0 |  |  | HKPARK | 2012-01-17 | Drawing has been updated to have the mentioned radius .688. The unit satisfies the specification. Ceramic thickness OK. | Use As Is
937 | JLAB-71 | WBASN | CRM1207075-0024 | BRAZEMENT ADAPTER ASSY |  | C100-CAV-INSP-ADPT | R4 | 184 | See attachment for dimensional non conforming. Note: For radius out of tolerance condition part tested with sample flange and will fit together. | OGLE | 2012-01-13 | 86A1DEDE-B1BE-CD48-43D0660C8E34BCD4 |  |  | HKPARK | 2012-01-17 | Drawing has been updated to have the mentioned radius .688. The unit satisfies the specification. Ceramic thickness OK. | Use As Is
938 | JLAB-75 | WBASN | CRM1207075-0024 | BRAZEMENT ADAPTER ASSY |  | C100-CAV-INSP-ADPT | R4 | 188 | See attachment for dimensional non conforming. Note: For radius out of tolerance condition part tested with sample flange and will fit together. | OGLE | 2012-01-13 | 86A86D74-BD60-6F73-FE5C6D2ABC418723 |  |  | HKPARK | 2012-01-17 | Drawing has been updated to have the mentioned radius .688. The unit satisfies the specification. | Use As Is
939 | JLAB-74 | WBASN | CRM1207075-0024 | BRAZEMENT ADAPTER ASSY |  | C100-CAV-INSP-ADPT | R4 | 187 | See attachment for dimensional non conforming. Note: For radius out of tolerance condition part tested with sample flange and will fit together. | OGLE | 2012-01-13 | 86B01BFA-9D76-F7FB-47903B08C86D0DB7 |  |  | HKPARK | 2012-01-17 | Drawing has been updated to have the mentioned radius .688. The unit satisfies the specification. Ceramic thickness OK. | Use As Is
940 | F067764/16 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R5 | 100 | Waveguide leaked inside the cleanroom. The knife edge will need rework done. | GEORGED | 2012-01-17 |  |  |  | GEORGED | 2012-02-06 |  | Modify |  | Knife edge reworked by the machine shop. | GEORGED | 2012-02-06 |  | Visual okay. Knife edge height measured .016" | GEORGED | 2012-02-06 |  | GEORGED | 2012-02-06 |  | Use As Is
941 | F058824/03 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R1 | 4 | Waveguide leaked inside the cleanroom. The knife edge will need rework done. | GEORGED | 2012-01-17 |  |  |  | GEORGED | 2012-01-27 |  | Modify |  | The machine shop reworked the knife edge. | GEORGED | 2012-01-27 |  | Visual done. Knife edge height measured .018" | GEORGED | 2012-01-27 | 3A764A4C-E7B9-B1FB-B17219FBE735E921 | GEORGED | 2012-01-27 |  | Use As Is
942 | F356223/03 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R6 | 109 | Waveguide leaked inside the cleanroom. The knife edge will need rework done. | GEORGED | 2012-01-17 |  |  |  | GEORGED | 2012-02-03 |  | Modify |  | Knife edge reworked by the machine shop. | GEORGED | 2012-02-03 |  | Visual done, part okay. | GEORGED | 2012-02-03 | A429979A-D0D1-9A55-99464E92FE3E6151 | GEORGED | 2012-02-03 |  | Use As Is
943 | JLAB-73 | WBASN | CRM1207075-0024 | BRAZEMENT ADAPTER ASSY |  | C100-CAV-INSP-ADPT | R4 | 186 | See attachment for dimensional non conforming. Part test good with sample flange. | OGLE | 2012-01-18 | 868FB62E-B72B-8C04-A0F603B5F97812CF |  |  | HKPARK | 2012-01-19 | Drawing has been updated to have the mentioned radius nominal .688. Part is within specification. | Use As Is
944 | JLAB-81 | WBASN | CRM1207075-0024 | BRAZEMENT ADAPTER ASSY |  | C100-CAV-INSP-ADPT | R4 | 194 | See attachment for dimensional non conforming. Part test good with sample flange. | OGLE | 2012-01-18 | 869CA309-C774-61AD-9A8D0BEB92548626 |  |  | HKPARK | 2012-01-19 | Drawing has been updated to have the mentioned radius nominal .688. Part is within specification. | Use As Is
945 | JLAB-84 | WBASN | CRM1207075-0024 | BRAZEMENT ADAPTER ASSY |  | C100-CAV-INSP-ADPT | R4 | 197 | See attachment for dimensional non conforming. Part test good with sample flange. | OGLE | 2012-01-18 | 86A3C444-BA83-77A8-7E14CC8108CC4CB2 |  |  | HKPARK | 2012-01-19 | Drawing has been updated to have the mentioned radius nominal .688. Part is within specification. | Use As Is
946 | JLAB-79 | WBASN | CRM1207075-0024 | BRAZEMENT ADAPTER ASSY |  | C100-CAV-INSP-ADPT | R4 | 192 | See attachment for dimensional non conforming. Part test good with sample flange. | OGLE | 2012-01-18 | 86A9E847-9A90-B62D-2ABA52ECE07DE894 |  |  | HKPARK | 2012-01-19 | Drawing has been updated to have the mentioned radius nominal .688. Part is within specification. | Use As Is
947 | JLAB-80 | WBASN | CRM1207075-0024 | BRAZEMENT ADAPTER ASSY |  | C100-CAV-INSP-ADPT | R4 | 193 | See attachment for dimensional non conforming. Part test good with sample flange. | OGLE | 2012-01-18 | 86AFF7D9-AD02-EA84-D5D0115711462237 |  |  | HKPARK | 2012-01-19 | Drawing has been updated to have the mentioned radius nominal .688. Part is within specification. Ceramic thickness OK. | Use As Is
948 | JLAB-77 | WBASN | CRM1207075-0024 | BRAZEMENT ADAPTER ASSY |  | C100-CAV-INSP-ADPT | R4 | 190 | See attachment for dimensional non conforming. Part test good with sample flange. | OGLE | 2012-01-18 | 86B68782-BCC6-6BB0-11F7E06FF68D0FD7 |  |  | HKPARK | 2012-01-19 | Drawing has been updated to have the mentioned radius nominal .688. Part is within specification. | Use As Is
949 | JLAB-76 | WBASN | CRM1207075-0024 | BRAZEMENT ADAPTER ASSY |  | C100-CAV-INSP-ADPT | R4 | 189 | See attachment for dimensional non conforming. Part test good with sample flange. | OGLE | 2012-01-18 | 86BCDF23-07FC-F642-D22D079F9946C7C0 |  |  | HKPARK | 2012-01-19 | Drawing has been updated to have the mentioned radius nominal .688. Part is within specification. | Use As Is
950 | JLAB-85 | WBASN | CRM1207075-0024 | BRAZEMENT ADAPTER ASSY |  | C100-CAV-INSP-ADPT | R4 | 198 | See attachment for dimensional non conforming. Part test good with sample flange. | OGLE | 2012-01-18 | 86C3314A-04BA-1803-F2A0EF7957E0B278 |  |  | HKPARK | 2012-01-19 | Drawing has been updated to have the mentioned radius nominal .688. Part is within specification. | Use As Is
951 | C100-RI-085 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INST-HELV | R2 | 85 | See attachment for dimensional non conforming. | OGLE | 2012-01-19 | B1444AEC-C298-9897-95C28D5A78E99A62 |  |  | ABURRILL | 2012-01-19 | dimensional deviations will be made up in CM assembly | Use As Is
952 | 48483-0084 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 240 | Unit has 7.0 -9 range leak ran test 3 times to confirm | WILCOX | 2012-01-20 |  |  |  | HKPARK | 2012-01-24 |  | Return To Vendor
953 | 48483-0093 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 241 | -7 RANGE LEAK DO NOT USE | WILCOX | 2012-01-20 |  |  |  | HKPARK | 2012-01-24 |  | Return To Vendor
954 | 48483-0090 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU | 8.0 E-10 RANGE LEAK DO NOT USE | C100-CAV-INSP-HMFT | R6 | 242 |  | WILCOX | 2012-01-20 |  |  |  | HKPARK | 2012-01-24 |  | Return To Vendor
955 | 48483-0094 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU | 9.2 - 4 RANGE LEAK DO NOT USE | C100-CAV-INSP-HMFT | R6 | 243 |  | WILCOX | 2012-01-20 |  |  |  | HKPARK | 2012-01-24 |  | Return To Vendor
956 | C100-RI-062 | CAVSN | CRM1207013-0001 | CAVITY ASSEMBLY |  | C100-CAV-INST-HELV | R2 | 86 | See attachment for dimensional non conforming. | OGLE | 2012-01-20 | E597D1A6-C74E-D12E-9AC0D7AE6068A6AE |  |  | ABURRILL | 2012-01-20 | dimensional deviation ok | Use As Is
957 | 48483-0086 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 247 | See attachment for dimensional non conforming. | OGLE | 2012-01-20 | EDBF826E-0F72-A93E-3066B57BAEF8686C |  |  | HKPARK | 2012-01-24 |  | Use As Is
958 | 48483-0087 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 246 | See attachment for dimensional non conforming. | OGLE | 2012-01-20 | EDD9D7D9-9466-958F-80194220F5E1FFD2 |  |  | HKPARK | 2012-01-24 | +/-.010 is OK. | Use As Is
959 | 48483-0089 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 244 | See attachment for dimensional non conforming. | OGLE | 2012-01-20 | EE0A882D-EC83-10BA-46A67951E6505FEA |  |  | HKPARK | 2012-01-24 | +/-.010 is OK. | Use As Is
960 | 48483-0092 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 245 | See attachment for dimensional non conforming. | OGLE | 2012-01-20 | EE290B01-EAF1-EEED-B212DE4558A443D8 |  |  | HKPARK | 2012-01-24 |  | Use As Is
961 | RI-PRO-014 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 117 | Reject as per HyeKyoung. The sapphire surface failed the visual inspection, the finish is too rough. | GEORGED | 2012-01-23 |  |  |  | HKPARK | 2012-01-23 | Vendor was notified and did not want the parts back. Keep in the inventory area until the PO is officially closed. | Reject
962 | RI-PRO-007 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 239 | Reject as per HyeKyoung. The sapphire surface failed the visual inspection, the finish is too rough. | GEORGED | 2012-01-23 |  |  |  | HKPARK | 2012-01-23 | Vendor was notified and did not want the parts back. Keep in the inventory area until the PO is officially closed. | Reject
963 | RI-PRO-029 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 131 | Reject as per HyeKyoung. The sapphire surface failed the visual inspection, the finish is too rough. | GEORGED | 2012-01-23 |  |  |  | HKPARK | 2012-01-23 | Vendor was notified and did not want the parts back. Keep in the inventory area until the PO is officially closed. | Reject
964 | RI-PRO-032 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 134 | Reject as per HyeKyoung. The sapphire surface failed the visual inspection, the finish is too rough. | GEORGED | 2012-01-23 |  |  |  | HKPARK | 2012-01-23 | Vendor was notified and did not want the parts back. Keep in the inventory area until the PO is officially closed. | Reject
965 | RI-PRO-042 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 142 | Reject as per HyeKyoung. The sapphire surface failed the visual inspection, the finish is too rough. | GEORGED | 2012-01-23 |  |  |  | HKPARK | 2012-01-23 | Vendor was notified and did not want the parts back. Keep in the inventory area until the PO is officially closed. | Reject
966 | RI-PRO-024 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R3 | 127 | Reject as per HyeKyoung. The sapphire surface failed the visual inspection, the finish is too rough. | GEORGED | 2012-01-23 |  |  |  | HKPARK | 2012-01-23 | Vendor was notified and did not want the parts back. Keep in the inventory area until the PO is officially closed. | Reject
967 | RI-PRO-047 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R4 | 144 | Reject as per HyeKyoung. The sapphire surface failed the visual inspection, the finish is too rough. | GEORGED | 2012-01-23 |  |  |  | HKPARK | 2012-01-23 | Vendor was notified and did not want the parts back. Keep in the inventory area until the PO is officially closed. | Reject
968 | 48483-0085 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 249 | 2.2 E-7 leak do not use | WILCOX | 2012-01-23 |  |  |  | HKPARK | 2012-01-24 |  | Return To Vendor
969 | 48483-0085 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 249 | See attachment for dimensional non conforming. | OGLE | 2012-01-23 | 821C1837-9E6B-E7AA-E1A98D25355AD69F |  |  | HKPARK | 2012-01-25 | The unit is .001 longer from the acceptance range. Unit is still functional use as is. | Use As Is
970 | 48483-0091 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 266 | Small burr in corner of seal groove. Not within the seal area. See attachment for dimensional non conforming. | OGLE | 2012-01-23 | 85153E51-9501-D1AE-30E5A9A5CB21C09C |  |  | HKPARK | 2012-01-25 |  | Use As Is
971 | 48483-0095 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 265 | See attachment for dimensional non conforming. | OGLE | 2012-01-23 | 852555A1-DC95-8AFF-9CB067051A54565D |  |  | HKPARK | 2012-01-25 | The unit is .001 longer from the acceptance range. Unit is still functional use as is. | Use As Is
972 | 48483-0096 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 254 | See attachment for dimensional non conforming. | OGLE | 2012-01-23 | 85323187-9046-AC93-E0BD9D5F119DF945 |  |  | HKPARK | 2012-01-25 |  | Use As Is
973 | 48483-0098 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 264 | See attachment for dimensional non conforming. | OGLE | 2012-01-23 | 85540DB4-DE4B-8F16-772DD624E7B235F5 |  |  | HKPARK | 2012-01-25 | The unit is .001 longer from the acceptance range. Unit is still functional use as is. | Use As Is
974 | 48483-0099 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 261 | See attachment for dimensional non conforming. | OGLE | 2012-01-23 | 8978FB00-ABC3-1D13-48C9358AC472FF17 |  |  | HKPARK | 2012-01-24 | +/-.010 is OK. | Use As Is
975 | 48483-0103 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 255 | See attachment for dimensional non conforming. | OGLE | 2012-01-23 | 89958863-D2E3-3CBE-B19CC480181247DA |  |  | HKPARK | 2012-01-25 |  | Return To Vendor
976 | 48483-0108 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 260 | Small burr in seal groove. Not within the seal area. See attachment for dimensional non conforming. | OGLE | 2012-01-23 | 8C88FA5C-D810-9E3E-F6DFF316DF385D9B |  |  | HKPARK | 2012-01-25 | The unit is .002 longer from the acceptance range. Unit is still functional use as is. | Use As Is
977 | 48483-0104 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 256 | See attachment for dimensional non conforming. | OGLE | 2012-01-23 | 8CB75412-A7A6-864F-69543C859C22F000 |  |  | HKPARK | 2012-01-25 | The unit is .001 longer from the acceptance range. Unit is still functional use as is. | Use As Is
978 | 48483-0105 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 257 | See attachment for dimensional non conforming. | OGLE | 2012-01-23 | 8CC53788-EEC7-8613-7F6AFDF5A0257EF1 |  |  | HKPARK | 2012-01-25 |  | Return To Vendor
979 | 48483-0107 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 259 | See attachment for dimensional non conforming. | OGLE | 2012-01-23 | 8CD70885-A903-D61A-C08217A8F45548E0 |  |  | HKPARK | 2012-01-25 |  | Use As Is
980 | FLG0037-221 | DSRFSN | MDC 203046-2000 | DOUBLE SIDED WINDOW FLANGE |  | C100-CAV-INSP-FLNG | R2 | 212 | See attachment for dimensional non conforming. | OGLE | 2012-01-24 | B9913CB1-EE70-01F7-F01F2549299C71BA |  |  | HKPARK | 2012-01-25 |  | Use As Is
981 | FLG0037-222 | DSRFSN | MDC 203046-2000 | DOUBLE SIDED WINDOW FLANGE |  | C100-CAV-INSP-FLNG | R2 | 213 | See attachment for dimensional non conforming. | OGLE | 2012-01-24 | BB875976-003D-522B-30D23C584350610E |  |  | HKPARK | 2012-01-25 |  | Use As Is
982 | FLG0037-223 | DSRFSN | MDC 203046-2000 | DOUBLE SIDED WINDOW FLANGE |  | C100-CAV-INSP-FLNG | R2 | 214 | See attachment for dimensional non conforming. | OGLE | 2012-01-24 | BEF5F354-D5F1-98A0-36A454FD623C3BCC |  |  | HKPARK | 2012-01-25 | Hole locations are slight off which still looks OK for assembly. Please check with the mating flange to make sure. | Use As Is
983 | F890195/04 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R6 | 122 | Waveguide leaked in the cleanroom. | GEORGED | 2012-01-25 |  |  |  | GEORGED | 2012-01-25 |  | Modify |  | The machine shop reworked the knife edge. | GEORGED | 2012-01-25 |  | Visual done. | GEORGED | 2012-01-25 |  | GEORGED | 2012-01-25 |  | Use As Is
984 | FLG0037-224 | DSRFSN | MDC 203046-2000 | DOUBLE SIDED WINDOW FLANGE |  | C100-CAV-INSP-FLNG | R2 | 215 | Damaged corner may affect assembly for gasket compression. | OGLE | 2012-01-30 | ED436A08-BB8C-1CCF-FBE3BF0113FECF50 |  |  | HKPARK | 2012-01-31 | Please repair if it can be done in house. | Modify |  | REPAIRED PIECE AS PER NCR DISPOSITION | WILCOX | 2012-02-02 |  | Verified rework. JO 2/2/12 | OGLE | 2012-02-02 |  | HKPARK | 2012-02-03 |  | Use As Is
985 | FLG0037-225 | DSRFSN | MDC 203046-2000 | DOUBLE SIDED WINDOW FLANGE |  | C100-CAV-INSP-FLNG | R2 | 216 | Damaged knife edge.See attachment for dimensional non conforming. | OGLE | 2012-01-30 | EF90CFEA-0066-4097-39F77A6C4C25B2BE |  |  | HKPARK | 2012-01-31 | Dimensions are acceptable and repair the knife edge if it can be done in house. | Modify |  | REPAIRED KNIFE EDGE TO THE BEST TO MY ABILITY. DOUBTFUL IT WILL SEAL. I THINK FURTHER REVIEW IS NEEDED | WILCOX | 2012-02-02 |  | After repair: knife edge damage still present. May not seal.See photo. JO 2/2/12 | OGLE | 2012-02-02 | 8AF73DD3-0F4A-C9DA-002C3A3FAA535D50 | HKPARK | 2012-02-03 |  | Return To Vendor
986 | FLG0037-226 | DSRFSN | MDC 203046-2000 | DOUBLE SIDED WINDOW FLANGE |  | C100-CAV-INSP-FLNG | R2 | 217 | Damaged knife edge. See attachment for dimensional non conforming. | OGLE | 2012-01-30 | F05C9526-BC22-FA5B-9AC0530DBE021387 |  |  | HKPARK | 2012-01-31 | Dimensions are acceptable. Repair the knife edge if it can be done in house. | Modify |  | Repaired knife edge as per NCR disposition | WILCOX | 2012-02-02 |  |  | WILCOX | 2012-02-02 |  | HKPARK | 2012-02-02 |  | Use As Is
987 | FLG0037-227 | DSRFSN | MDC 203046-2000 | DOUBLE SIDED WINDOW FLANGE |  | C100-CAV-INSP-FLNG | R2 | 218 | Damaged knife edge. | OGLE | 2012-01-30 | F1722573-F4B2-FF5B-A1A82F9CC4D1DDB3 |  |  | HKPARK | 2012-01-31 | Please repair if it can be done in house. | Modify |  | Repaired knife edge as per NCR disposition | WILCOX | 2012-02-02 |  |  | WILCOX | 2012-02-02 |  | HKPARK | 2012-02-02 |  | Use As Is
988 | FLG0037-228 | DSRFSN | MDC 203046-2000 | DOUBLE SIDED WINDOW FLANGE |  | C100-CAV-INSP-FLNG | R2 | 219 | Damaged knife edge. | OGLE | 2012-01-30 | F19E3F01-C107-A088-CC924B387C2FED78 |  |  | HKPARK | 2012-01-31 | Please repair if it can be done in house. | Modify |  | Repaired knife edge as per NCR disposition | WILCOX | 2012-02-02 |  |  | WILCOX | 2012-02-02 |  | HKPARK | 2012-02-02 |  | Use As Is
989 | F067763/07 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R4 | 54 | Waveguide leaked in the cleanroom. | GEORGED | 2012-02-01 |  |  |  | GEORGED | 2012-02-01 |  | Modify |  | Knife edge reworked by the machine shop. | GEORGED | 2012-02-01 |  | Visual okay. Knife edge height measured .016" | GEORGED | 2012-02-01 |  | GEORGED | 2012-02-01 |  | Use As Is
990 | FLG0037-232 | DSRFSN | MDC 203046-2000 | DOUBLE SIDED WINDOW FLANGE |  | C100-CAV-INSP-FLNG | R2 | 223 | See attachment for dimensional non conforming. | OGLE | 2012-02-01 | 58102FFB-A7D8-C2F1-E9231498DB7B40A4 |  |  | HKPARK | 2012-02-02 | Not a critical dimension. use as is. | Use As Is
991 | FLG0037-234 | DSRFSN | MDC 203046-2000 | DOUBLE SIDED WINDOW FLANGE |  | C100-CAV-INSP-FLNG | R2 | 225 | Slight face damage;may affect seal compression. | OGLE | 2012-02-01 | 5A4ED2AF-FE1F-46A8-720824D6AAE3F4E0 |  |  | HKPARK | 2012-02-02 | Try to repair in house. | Modify |  | REPAIRED PART AS PER NCR DISPOSITION | WILCOX | 2012-02-02 |  | Verified rework. 2/2/12 JO | OGLE | 2012-02-02 |  | HKPARK | 2012-02-03 |  | Use As Is
992 | ASSEM0035-140 | WINSN | CRM1207075-0035 | ADAPTER WINDOW ASSEMBLY |  | C100-CAV-INSP-ASSY-WIN | R3 | 142 | Failed RF test. Ceramic has multiple cracks. | CARPENTE | 2012-02-09 | D9BD2E87-B6B2-7CB5-AAF9246C608ABC99 |  |  | HKPARK | 2012-02-13 |  | Reject
993 | ASSEM0035-161 | WINSN | CRM1207075-0035 | ADAPTER WINDOW ASSEMBLY |  | C100-CAV-INSP-ASSY-WIN | R3 | 161 | Slight Knife Edge damage. No photo available. | OGLE | 2012-03-07 |  |  |  | HKPARK | 2012-03-07 | Rework | Modify |  | polished up knife edge at damaged spot. surface is smooth but you can still see SOME visual blemishes under a magnifier. part OK | WILCOX | 2012-03-08 |  | Rework verified. Acceptable JO 3/9/12. | OGLE | 2012-03-09 |  | HKPARK | 2012-03-13 |  | Use As Is
994 | JLAB-44 | WBASN | CRM1207075-0024 | BRAZEMENT ADAPTER ASSY |  | C100-CAV-INSP-ADPT | R4 | 159 | Large temperature gradient (~20 C) across the short axis of the ceramic. | DRURY | 2012-03-28
995 | JLAB-45 | WBASN | CRM1207075-0024 | BRAZEMENT ADAPTER ASSY |  | C100-CAV-INSP-ADPT | R4 | 160 | Large temperature gradient (~20 C) across the short axis of the ceramic. | DRURY | 2012-03-28
996 | FLG0037-010 | DSRFSN | MDC 203046-2000 | DOUBLE SIDED WINDOW FLANGE |  | C100-CAV-INSP-FLNG | R2 | 229 | See attachment for dimensional non conforming. | OGLE | 2012-04-18 | B64DEDF1-FF70-90D0-3C1D09DF67FAF24F |  |  | HKPARK | 2012-04-18 | The deviation values are small and it won't have assembly issue. | Use As Is
997 | FLG0037-006 | DSRFSN | MDC 203046-2000 | DOUBLE SIDED WINDOW FLANGE |  | C100-CAV-INSP-FLNG | R2 | 230 | See attachment for dimensional non conforming. | OGLE | 2012-04-18 | B65AAF05-D4F6-C687-9283316BE092938B |  |  | HKPARK | 2012-04-18 | No assembly issue. | Use As Is
998 | F892522/08 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R6 | 158 | Knife edge rounded off, very dull in one spot. Brian Carpenter has seen this personally. Also sent photos to Brian Carpenter and John Fischer. | WILCOX | 2012-04-25 | 20C3C193-BA35-B481-8BE959B60BD3536B |  |  | KWILSON | 2012-10-01 | Unrepairable damage to knife edge. | Return To Vendor | 20BD40C8-FD8D-D26D-7D33E665AF8B28AF
1001 | F892522/05 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R6 | 153 | Light scratch on FPC flange, and flatness = 0.004" | CARPENTE | 2012-09-12 |  |  |  | KWILSON | 2012-09-12 | Lightly polish scratches, flatness ok. | Modify |  | Polished by Daniels | CARPENTE | 2012-10-15 |  | Complete, looks good. | CARPENTE | 2012-10-15 |  | KWILSON | 2012-10-15 |  | Use As Is
1002 | F890195/03 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R6 | 150 | Light scratches on FPC flange, surface finish = 22 | CARPENTE | 2012-09-13 |  |  |  | KWILSON | 2012-10-01 | Please lightly polish flange face. | Modify |  | Polished by Daniels | CARPENTE | 2012-10-15 |  | Complete, looks good. | CARPENTE | 2012-10-15 |  | KWILSON | 2012-10-15 |  | Use As Is
1003 | F898271/05 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R6 | 155 | Nick on knife edge, flatness = 0.004, FPC bolt pattern out of tolerance. | CARPENTE | 2012-09-13 | 7C7582CA-F704-6710-7CCFDD4F4CDE0C46 |  |  | KWILSON | 2012-10-01 | Flatness ok, bolt pattern ok since it bolted together. Please clean up knife edge. Knife edge could not be cleaned up. Mark as scrap, do not discard. | Modify |  | Repair attempted by Wilcox. | CARPENTE | 2012-10-15 |  | Unable to be repaired, recommend do not use. | CARPENTE | 2012-10-15 |  | KWILSON | 2012-10-15 |  | Reject
1004 | F898271/06 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R6 | 161 | Nick on the knife edge, some areas are rounded off. The small flange has scratches in the seal path. | CARPENTE | 2012-09-13 |  |  |  | KWILSON | 2012-10-01 | Please clean up knife edge, light polish scratches on flange face. | Modify |  | Repaired by Wilcox. | CARPENTE | 2012-10-15 |  | Complete, looks good. | CARPENTE | 2012-10-15 |  | KWILSON | 2012-10-15 |  | Use As Is
1005 | F356225/01 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R6 | 159 | Blisters on copper plating inside waveguide (large flange side). The small flange has a deep scratch in the seal path. Light machining imprint present inside o-ring groove. The knife edge is dull in a few spots. FPC flange flatness = 0.005 | CARPENTE | 2012-09-13 |  |  |  | KWILSON | 2012-10-08 |  | Reject
1006 | F898271/08 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R6 | 147 | Marks on FPC flange in seal area. | CARPENTE | 2012-09-13 |  |  |  | KWILSON | 2012-10-01 | Please polish as needed. Thanks. | Modify |  | Polished by Daniels | CARPENTE | 2012-10-15 |  | Complete, looks good. | CARPENTE | 2012-10-15 |  | KWILSON | 2012-10-15 |  | Use As Is
1007 | F356224/04 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R6 | 148 | Scratch on FPC flange seal area. | CARPENTE | 2012-09-13 |  |  |  | KWILSON | 2012-10-01 | Please polish as needed to remove scratches | Modify |  | Reworked by Daniels. | CARPENTE | 2012-10-15 |  | Complete, looks good. | CARPENTE | 2012-10-15 |  | KWILSON | 2012-10-15 |  | Use As Is
1008 | F898271/03 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R6 | 151 | Mark on FPC flange seal area, flatness = 0.006 | CARPENTE | 2012-09-13 |  |  |  | KWILSON | 2012-10-01 | Please polish to remove scratches. Flatness ok. | Modify |  | Polished by Daniels | CARPENTE | 2012-10-15 |  | Complete, looks good. | CARPENTE | 2012-10-15 |  | KWILSON | 2012-10-15 |  | Use As Is
1009 | F898271/07 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R6 | 154 | Marks on FPC flange seal area. | CARPENTE | 2012-09-14 |  |  |  | KWILSON | 2012-10-01 | Please polish as needed. Thanks! | Modify |  | Polished by Daniels | CARPENTE | 2012-10-15 |  | Complete, looks good. | CARPENTE | 2012-10-15 |  | KWILSON | 2012-10-15 |  | Use As Is
1010 | F898271/04 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R6 | 149 | Nicks on knife edge | CARPENTE | 2012-09-14 |  |  |  | KWILSON | 2012-10-01 | Please clean up knife edge as needed. Thanks. | Modify |  | Knife edge repair was attempted by Chris Wilcox | CARPENTE | 2012-10-12 |  | Unable to repair satisfactory. Recommend to not use. | CARPENTE | 2012-10-12 |  | CARPENTE | 2012-10-12 |  | Modify
1011 | F898271/01 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R6 | 152 | FPC flange flatness = 0.004 | CARPENTE | 2012-09-14 |  |  |  | KWILSON | 2012-10-01 |  | Use As Is
1012 | F892522/08 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | C100-CAV-INSP-WGD | R6 | 158 | FPC flange finish = 0.004 | CARPENTE | 2012-09-14 |  |  |  | KWILSON | 2012-10-01 |  | Use As Is
1013 | F067763/02 |  | unkn | unkn |  | C100-CAV-INSP-WGD | R6 | 157 | Small flange - Deep scratches and pits present. Entire flange has an odd finish. | CARPENTE | 2012-10-02 |  |  |  | KWILSON | 2012-10-15 | Polish flange | Modify |  | Polished by Daniels | CARPENTE | 2012-10-15 |  | Complete, looks good. | CARPENTE | 2012-10-15 |  | KWILSON | 2012-10-15 |  | Use As Is
1014 | F356223/05 |  | unkn | unkn |  | C100-CAV-INSP-WGD | R6 | 162 | Light scratches on the small flange. Flatness = 0.004 | CARPENTE | 2012-10-02 |  |  |  | KWILSON | 2012-10-02 | Please lightly polish to remove scratches. Flatness ok. | Modify |  | Polished by Daniels | CARPENTE | 2012-10-15 |  | Complete, looks good. | CARPENTE | 2012-10-15 |  | KWILSON | 2012-10-15 |  | Use As Is
1015 | F898271/02 |  | unkn | unkn |  | C100-CAV-INSP-WGD | R6 | 160 | Small flange has stains and scratches. Dings present inside o-ring groove. | CARPENTE | 2012-10-03 |  |  |  | KWILSON | 2012-10-03 | Please lightly polish flange face. Ding ok. | Modify |  | Polished by Daniels. | CARPENTE | 2012-10-15 |  | Complete, looks good. | CARPENTE | 2012-10-15 |  | KWILSON | 2012-10-15 |  | Use As Is
1016 |  |  | unkn | unkn |  |  |  |  | Very bad nick on knife edge. | GEORGED | 2013-02-27
1017 | ASSEM0035-172 | WINSN | CRM1207075-0035 | Adapter Window Assy |  | C100-CAV-INSP-ASSY-WIN | R3 | 188 | Very bad nick on knife edge. | GEORGED | 2013-02-27 | 2870D4F8-BC93-DBD1-5056BC4A22EF6F16 |  |  | HKPARK | 2013-03-04 | Please polish and check if it's repaired. | Modify |  | Repaired knife edge. Knife edge is smooth. Had Mike McCrea look at it after I repaired. Leak checked unit using inficon and a new gasket with bolts torqued at 80 in./lbs. using proper sequence. | WILCOX | 2013-03-06 | 8153C060-C533-32D0-89FDFE8F86E0A074 | 3-6-2013 Visual okay. 3-11-2013 Visual okay after 2nd repair to knife edge. | GEORGED | 2013-03-11 |  | HKPARK | 2013-03-21 |  | Use As Is
1018 | 49569-04 |  | unkn | unkn |  | C100-CAV-INSP-HMFT | R6 | 270 | Step depth out of tolerance. | CARPENTE | 2015-08-20 |  |  |  | HKPARK | 2019-09-30 | Feedthru missing | Reject
1019 | 49569-07 |  | unkn | unkn |  | C100-CAV-INSP-HMFT | R6 | 273 | Probe tip length and step depth both out of tolerance. | CARPENTE | 2015-08-20 |  |  |  | HKPARK | 2019-09-30 | Feedthru missing | Reject
1020 | 49569-10 |  | unkn | unkn |  | C100-CAV-INSP-HMFT | R6 | 276 | Step depth out of tolerance. | CARPENTE | 2015-08-20 |  |  |  | HKPARK | 2019-09-30 | Feedthru missing | Reject
1021 | 49569-09 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 275 | Failed leak check. | GEORGED | 2015-09-08 |  |  |  | HKPARK | 2019-09-30 |  | Reject
1022 | 49699-3 | HMFTSN | CRM1207014-0020 | HOM PROBE FEED THRU |  | C100-CAV-INSP-HMFT | R6 | 281 | Failed leak check. Flange seal surface to probe tip dimension out of tolerance (measured 1.473"). | GEORGED | 2015-09-08 |  |  |  | HKPARK | 2019-09-30 |  | Reject

## C100-NCR
Status Report for C100-NCR on 10-Jun-20
TRAV_SEQ_NUM | SN | SNTYPE | REWORKDATE1 | DRAWINGNUMBER | REWORKTECH1 | PARTDESCRIPTION | REWORKCOMMENT1 | SHIPPINGDAMAGE | REWORKFILE1 | DIMENSIONAL | REWORKDATE2 | REWORKTECH2 | LEAKTESTFAILURE | REWORKCOMMENT2 | ROUGHNESS | REWORKFILE2 | FLATNESS | PARALLELISM | REWORKDATE3 | REWORKTECH3 | SCRATCHES | REWORKCOMMENT3 | DINGS | REWORKFILE3 | DELAMINATION | PLATING | REWORKDATE4 | STAINS | REWORKTECH4 | REWORKCOMMENT4 | OXIDATION | RESIDUES | REWORKFILE4 | INSPECTIONFAILURE | REWORKDATE5 | RFTESTFAILURE | REWORKTECH5 | REWORKCOMMENT5 | VACUUM | RFCMTF | REWORKFILE5 | INSTRUMENTATION | VERIFYDATE1 | MECHANICAL | VERIFYTECH1 | CMTFRFPERFORMANCE | VERIFYCOMMENT1 | VERIFYFILE1 | VTARFPERFORMANCE | OTHER | VERIFYDATE2 | TRAVELERID | VERIFYTECH2 | TRAVELERREVISION | VERIFYCOMMENT2 | TRAVSEQNUMBER | VERIFYFILE2 | DESCRIBESTAFF | VERIFYDATE3 | DESCRIBEDATE | VERIFYTECH3 | DESCRIBECOMMENT | VERIFYCOMMENT3 | DESCRIBEDOCS | VERIFYFILE3 | RESPONDINGENGINEER | VERIFYDATE4 | RESPONDINGDATE | VERIFYTECH4 | DISPOSITION | VERIFYCOMMENT4 | DISPOSITIONCOMMENT | VERIFYFILE4 | DISPOSITIONDOCS | VERIFYDATE5 | REWORKSTAFF | VERIFYTECH5 | REWORKDATE | VERIFYCOMMENT5 | REWORKCOMMENT | VERIFYFILE5 | RESPONDINGDATE1 | REWORKDOCS | RESPONDINGENGR1 | VERIFYSTAFF | RESPONDINGCOMMENT1 | VERIFYDATE | RESPONDINGFILE1 | VERIFYCOMMENT | RESPONDINGDATE2 | VERIFYDOCS | FINALENGINEER | RESPONDINGENGR2 | FINALDATE | RESPONDINGCOMMENT2 | FINALDISPOSITION | RESPONDINGFILE2 | FINALCOMMENT | RESPONDINGDATE3 | FINALDOCS | RESPONDINGENGR3 | CONNECTIVITYCOMMENT | RESPONDINGCOMMENT3 | NOTIFICATIONCOMMENT | RESPONDINGFILE3 | RESPONDINGDATE4 | RESPONDINGENGR4 | RESPONDINGCOMMENT4 | RESPONDINGFILE4 | RESPONDINGDATE5 | RESPONDINGENGR5 | RESPONDINGCOMMENT5 | RESPONDINGFILE5
17 | 051810_01 through 051810_20 |  |  | 115120-1013C Rev F |  | Ceramic Windows | [45 empty] | C100-INSR |  | R1 |  | 28 | [5 empty] | Length on windows 051810_01 through 051810_20 are out of tolerance |  |  |  | HKPARK | [5 empty] | Ceramics for in-house adapter fabrication. Brazed adapter were inspected again and used when it passed functional test.
55 | 062310_01 through 062310_20 |  |  | CRM12070750027 |  | Window Adapter | [45 empty] | C100-INSR |  | R1 |  | 30 | [5 empty] | All are dimensional out of tolerance |  |  |  | HKPARK |  |  |  | Use As Is |  | The serial no. 062310_02 and_19 should be returned to the vendor for rework. The
56 | 062310_02 |  |  | CRM12070750027 |  | C100 Window Adapters | [45 empty] | C100-INSR |  | R1 |  | 31 | [5 empty] | Dimensions are out of tolerance |  |  |  | HKPARK |  |  |  | Use As Is |  | The critical dimension (.033/.038) were reworked and within tolerance now.
57 | 5 |  |  | CRM1207080-1000 |  | Helium Headers | [45 empty] | C100-CM-INSP-HEHD |  | R1 |  | 2 | [5 empty] | Please see attached report. Note: welding inspection not complete. |  |  |  | CHENG |  |  |  | Use As Is |  | The liquid level holders (-0901)have some noticeable out of tolerances. The 4.20
58 | MM0025_01 |  |  | CRM1207020-0025 |  | Helium Vessel Hd | [45 empty] | C100-CM-INSP-HVHD |  | R1 |  | 2 | [9 empty] | CHENG |  |  |  | Use As Is |  | The 5.94" dimension is measured to be 5.92". The mating part has oversized holes
59 | MM0025_02 |  |  | CRM1207020-0025 |  | Helium Vessel Hd | [45 empty] | C100-CM-INSP-HVHD |  | R1 |  | 3 | [9 empty] | CHENG |  |  |  | Use As Is |  | The 5.94" dimension is measured to be 5.92". The mating part has oversized holes
60 | MM0025_03 |  |  | CRM1207020-0025 |  | Helium Vessel Hd | [45 empty] | C100-CM-INSP-HVHD |  | R1 |  | 4 | [9 empty] | CHENG |  |  |  | Use As Is |  | The 5.94" dimension is measured to be 5.92". The mating part has oversized holes
61 | MM0025_04 |  |  | CRM1207020-0025 |  | Helium Vessel Hd | [45 empty] | C100-CM-INSP-HVHD |  | R1 |  | 5 | [9 empty] | CHENG |  |  |  | Use As Is |  | The 5.94" dimension is measured to be 5.92". The mating part has oversized holes
62 | MM0025_05 |  |  | CRM1207020-0025 |  | Helium Vessel Hd | [45 empty] | C100-CM-INSP-HVHD |  | R1 |  | 6 | [9 empty] | CHENG |  |  |  | Use As Is |  | The 5.94" dimension is measured to be 5.92". The mating part has oversized holes
63 | MM0025_06 |  |  | CRM1207020-0025 |  | Helium Vessel Hd | [45 empty] | C100-CM-INSP-HVHD |  | R1 |  | 7 | [9 empty] | CHENG |  |  |  | Use As Is |  | The 5.94" dimension is measured to be 5.92". The mating part has oversized holes
64 | MM0025_08 |  |  | CRM1207020-0025 |  | Helium Vessel Hd | [45 empty] | C100-CM-INSP-HVHD |  | R1 |  | 9 | [9 empty] | CHENG |  |  |  | Use As Is |  | The 5.94" dimension is measured to be 5.92". The mating part has oversized holes
65 | MM0025_09 |  |  | CRM1207020-0025 |  | Helium Vessel Hd | [45 empty] | C100-CM-INSP-HVHD |  | R1 |  | 10 | [9 empty] | CHENG |  |  |  | Use As Is |  | The 5.94" dimension is measured to be 5.92". The mating part has oversized holes
66 | MM0025_10 |  |  | CRM1207020-0025 |  | Helium Vessel Hd | [45 empty] | C100-CM-INSP-HVHD |  | R1 |  | 11 | [9 empty] | CHENG |  |  |  | Use As Is |  | The 5.94" dimension is measured to be 5.92". The mating part has oversized holes
67 | MM0025_12 |  |  | CRM-120-7075-0025 |  | He Vessel Head Assy | [45 empty] | C100-CM-INSP-HVHD |  | R1 |  | 13 | [9 empty] | CHENG |  |  |  | Use As Is |  | The positions of tapped and through holes are not critical since the locking rod
68 | MM0025_13 |  |  | CRM-120-7075-0025 |  | He Vessel Head Assy | [45 empty] | C100-CM-INSP-HVHD |  | R1 |  | 14 | [9 empty] | CHENG |  |  |  | Use As Is |  | The position of tapped holes is not critical since the mating parts have oversiz
69 | MM0025_14 |  |  | CRM-120-7075-0025 |  | He Vessel Head Assy | [45 empty] | C100-CM-INSP-HVHD |  | R1 |  | 15 | [9 empty] | CHENG |  |  |  | Use As Is |  | The positions of tapped and through holes are not critical since the locking rod
70 | MM0025_15 |  |  | CRM-120-7075-0025 |  | He Vessel Head Assy | [45 empty] | C100-CM-INSP-HVHD |  | R1 |  | 16 | [9 empty] | CHENG |  |  |  | Use As Is |  | The position of tapped holes is not critical since the parts mating tapped holes
71 | MM0025_16 |  |  | CRM-120-7075-0025 |  | He Vessel Head Assy | [45 empty] | C100-CM-INSP-HVHD |  | R1 |  | 17 | [9 empty] | CHENG |  |  |  | Use As Is |  | The out-of-tolerances are acceptable and this head assembly can be used.
72 | MM0025_17 |  |  | CRM-120-7075-0025 |  | He Vessel Head Assy | [45 empty] | C100-CM-INSP-HVHD |  | R1 |  | 18 | [9 empty] | CHENG |  |  |  | Use As Is |  | The positions of tapped holes are not critical since the mating parts have overs
73 | MM0025_18 |  |  | CRM-120-7075-0025 |  | He Vessel Head Assy | [45 empty] | C100-CM-INSP-HVHD |  | R1 |  | 19 | [9 empty] | CHENG |  |  |  | Use As Is |  | The out-of-tolerances are acceptable and this head assembly can be used.
74 | MM0025_19 |  |  | CRM-120-7075-0025 |  | He Vessel Head Assy | [45 empty] | C100-CM-INSP-HVHD |  | R1 |  | 20 | [9 empty] | CHENG |  |  |  | Use As Is |  | The out-of-tolerance is acceptable and this head assembly can be used.
75 | MM0025_20 |  |  | CRM-120-7075-0025 |  | He Vessel Head Assy | [45 empty] | C100-CM-INSP-HVHD |  | R1 |  | 21 | [9 empty] | CHENG |  |  |  | Use As Is |  | The position of tapped holes is not critical since the mating parts have oversiz
76 | Assy 63 |  |  | CRM-120-7050-0000 |  | Cold Tuner Assembly | [45 empty] | C100-CM-INSP-CTUN2 |  | R1 |  | 4 | [5 empty] | Multiple parts out of tolerance, PDF attached |  |  |  | JOSEPHM |  |  |  | Use As Is
77 | Gaskets /Serpentine Seal |  |  | CRM-100-7014-0028 |  | Gaskets/ Serpentine Seals | [45 empty] | C100-INSR |  | R1 |  | 32 | [5 empty] | Thickness out of tolerance 0.001 on 3 serpentine seals. OD out of tolerance 0.00 |  |  |  | CARPENTE |  |  |  | Use As Is
78 | MM0025_35 |  |  | CRM1207020-0025 |  | Helium Vessel Head | [45 empty] | C100-CM-INSP-HVHD |  | R1 |  | 22 | [9 empty] | CHENG |  |  |  | Use As Is |  | The 5.94" dimension is measured to be 5.92". The mating part has oversized holes
79 | MM0025_36 |  |  | CRM1207020-0025 |  | Helium Vessel Head | [45 empty] | C100-CM-INSP-HVHD |  | R1 |  | 23 | [9 empty] | CHENG |  |  |  | Use As Is |  | The 5.94" dimension is measured to be 5.92". The mating part has oversized holes
80 | Batch/Lot 004 |  |  | CRM-120-7042-0000 |  | Inner Magnetic Shielding | [45 empty] | C100-CM-INSP-IMAG |  | R2 |  | 6 | [5 empty] | Parts dirty and bent up. For part CRM12070420007 Nominal Zone C3 3.55, Measured |  |  |  | CHENG |  |  |  | Return To Vendor |  | Dimension deviations are acceptable and they may in fact facilitate assembly. Ho
81 | 001_4K |  |  | CRM-120-7042-0000 |  | Inner Magnetic Shielding | [45 empty] | C100-CM-INSP-IMAG |  | R2 |  | 8 | [5 empty] | Dimensions out of tolerance. See attachment for data |  |  |  | CHENG |  |  |  | Use As Is |  | This shield will be used in a cryogenic test. The vendor will be notified of the
82 | 10 |  |  | 09-2787-4800 |  | HELIUM VESSEL ASSEMBLY | [45 empty] | C100-CM-INSP-HELV |  | R2 |  | 5 | [5 empty] | BELLOW:WELD ON SHOULDER OF INSIDE DIAMETER.SEE PHOTOS IN TRAVELER. SEE ATTACHME |  |  |  | CHENG |  |  |  | Use As Is |  | The extra weld on bellows ID is found not interfering with helium vessel head. B
83 | C100-RI-019 |  |  | 09-2787-4800 |  | HELIUM VESSEL ASSEMBLY | [45 empty] | C100-CM-INSP-HELV |  | R2 |  | 6 | [5 empty] | SEE ATTACHMENTS FOR DIMENSIONAL NON CONFORMING. |  |  |  | CHENG |  |  |  | Use As Is |  | Ob bellows, the most siginificant out-of-tolerance is on the overall length of 3
84 | 28 |  |  | 09-2787-4800 |  | HELIUM VESSEL ASSEMBLY | [45 empty] | C100-CM-INSP-HELV |  | R2 |  | 7 | [5 empty] | SEE ATTACHMENTS FOR DIMENSIONAL NON CONFORMING |  |  |  | CHENG |  |  |  | Use As Is |  | The bellows cuffs that weld to the shell has ODs nearly 1/4" smaller than requir
85 | 13085-01-001 |  |  | 13085-01 |  | 8 pin cryo feedthru | [45 empty] | C100-CM-INSP-ELFT |  | R1 |  | 2 | [5 empty] | after cold shocking feedthru developed a leak in the ceramic to pin area. sent b |  |  |  | FISCHER |  |  |  | Return To Vendor |  | returned for replacement. | [29 empty] | feedthru leaks
86 | 13085-01-002 |  |  | 13085-01 |  | 8 pin cryo feedthru | [45 empty] | C100-CM-INSP-ELFT |  | R1 |  | 3 | [5 empty] | feedthru developed leak after cold shock in VTA. |  |  |  | FISCHER |  |  |  | Return To Vendor |  | ITEM RETURNED TO VENDOR FOR REPLACEMENT.
87 | 3 |  |  | 13085-01 |  | 8 PIN CRYO FEEDTHRU | [45 empty] | C100-CM-INSP-ELFT |  | R1 |  | 4 | [5 empty] | PART DEVELOPED LEAK AFTER CRYOCYCLE IN VTA. |  |  |  | FISCHER |  |  |  | Return To Vendor |  | FEEDTHRU RETURNED TO VENDOR FOR REPLACEMENT.
88 | 6 |  |  | 13085-01 |  | 8 pin cryo feedthru | [45 empty] | C100-CM-INSP-ELFT |  | R1 |  | 7 | [5 empty] | part failed after vta cryocycle. |  |  |  | FISCHER |  |  |  | Return To Vendor |  | part returned to vendor for replacement.
89 | 10 |  |  | 13085-01 |  | 8 pin cryo feedthru | [45 empty] | C100-CM-INSP-ELFT |  | R1 |  | 11 | [5 empty] | feedthru leaked after cold cycle in VTA |  |  |  | FISCHER |  |  |  | Return To Vendor |  | FEEDTHRU RETURNED FOR REPLACEMENT.
90 | 14 |  |  | 13085-01 |  | 8 pin cryo feedthru | [45 empty] | C100-CM-INSP-ELFT |  | R1 |  | 15 | [5 empty] | feedthru failed during vta cryocycle |  |  |  | FISCHER |  |  |  | Return To Vendor |  | part returned to vendor for replacement
91 | 17 |  |  | 13085-01 |  | 8 pin cryo feedthru | [45 empty] | C100-CM-INSP-ELFT |  | R1 |  | 18 | [5 empty] | part leaked after cryocycle in VTA. |  |  |  | FISCHER |  |  |  | Return To Vendor |  | item returned to vendor for replacement.
92 | 28 |  |  | 13085-01 |  | 8 pin cryo feedthru | [45 empty] | C100-CM-INSP-ELFT |  | R1 |  | 29 | [5 empty] | feedthru leaked affter cold cycle in VTA |  |  |  | FISCHER |  |  |  | Return To Vendor |  | ITEM RETURNED TO VENDOR FOR REPLACEMENT.
93 | 54 |  |  | 13085-01 |  | 8 pin cryo feedthru | [45 empty] | C100-CM-INSP-ELFT |  | R1 |  | 55 | [5 empty] | feedthru leaked after cryocycle in VTA |  |  |  | FISCHER |  |  |  | Return To Vendor |  | item returned to vendor for replacement
94 | 56 |  |  | 13085-01 |  | 8 pin cryo feedthru | [45 empty] | C100-CM-INSP-ELFT |  | R1 |  | 57 | [5 empty] | item leaked after cryocycle in VTA |  |  |  | FISCHER |  |  |  | Return To Vendor |  | ITEM RETURNED FOR REPLACEMENT
95 | 4 |  |  | CRM12070420000 |  | Inner Magnetic Shielding | [45 empty] | C100-CM-INSP-IMAG |  | R3 |  | 11 | [5 empty] | Dimensions out of tolerance. CRM-120-7042-0007 Nominal 2.00" Meas. 1.96" CRM |  |  |  | CHENG |  |  |  | Use As Is |  | We shall try to use these end caps (-0007 and -0008), though the undersized slot
96 | 1 |  |  | CRM12070901000 |  | Vacuum Vessel | [45 empty] | C100-CM-INSP-VV |  | R2 |  | 1 | [5 empty] | See traveler for out of tolerance conditions. |  |  |  | CHENG |  |  |  | Use As Is |  | Vendor has been notified of the non-conformances so that they will improve the q
97 | 2 |  |  | CRM12070901000 |  | Vacuum Vessel | [45 empty] | C100-CM-INSP-VV |  | R3 |  | 3 | [5 empty] | Please see traveler for out of tolerance conditions. |  |  |  | CHENG |  |  |  | Use As Is |  | The vendor visited JLAb to discuss about the possible causes of non-conformances
98 | 1 |  |  | CRM1207061-2001 |  | SUPPLY BEAMPIPE | [45 empty] | C100-CM-INSP-SUBP |  | R1 |  | 3 | [5 empty] | PACKAGING CONCERNS EXIST FOR THIS PART,SEE TRAVELER. SEE ATTACHMENT FOR DIMENSI |  |  |  | JOSEPHM |  |  |  | Use As Is
99 | 2001-02 | [50 empty] | C100-CM-INSP-SUBP |  | R1 |  | 4 | [5 empty] | DAMAGED BELLOW SEE ATTACHED PHOTO. PACKAGING CONCERNS EXIST FOR THIS PART,SEE |  |  |  | JOSEPHM |  |  |  | Return To Vendor |  | Part was sent back to vendor for repair.
100 | CRM12070440000_10 |  |  | CRM12070440000 |  | Thermal Shield | [45 empty] | C100-CM-INSP-THRM |  | R2 |  | 20 | [5 empty] | Length of straps out of tolerance. see attachment. Missing: .25-20 UNC nuts |  |  |  | KSMITH |  |  |  | Use As Is |  | Strap lg ok to use.Missing fasteners have been bought by project engineer. Erron
102 | 9 |  |  | CRM-120-7044-0000 |  | Thermal Shield | [45 empty] | C100-CM-INSP-THRM |  | R2 |  | 19 | [5 empty] | CRM-120-7044-0019 Length of strap out of tolerance. Measured 8.5" Nominal 9" |  |  |  | KSMITH |  |  |  | Use As Is |  | Strap lg is ok to use. Missing fasteners were bought by project engineer. Errone
103 | 8 |  |  | CRM-120-7044-0000 |  | Thermal Shield | [45 empty] | C100-CM-INSP-THRM |  | R2 |  | 21 | [5 empty] | Missing the following: .25-20 UNC nuts .25-20 UNC X .50 LG |  |  |  | KSMITH |  |  |  | Use As Is |  | Missing fasteners were procured by project engineer, and erroneous BOM correctio
104 | 7 |  |  | CRM-120-7044-0000 |  | Thermal Shield | [45 empty] | C100-CM-INSP-THRM |  | R2 |  | 15 | [5 empty] | Missing the following: .25-20 UNC nuts .25-20 UNC X .50 LG |  |  |  | KSMITH |  |  |  | Use As Is |  | Missing fasteners were procured by project engineer, and erroneous BOM correctio
105 | 1 |  |  | CRM1207061-2002 |  | RETURN BEAMPIPE | [45 empty] | C100-CM-INSP-RTBP |  | R1 |  | 2 | [5 empty] | PACKAGING IS NOT ENOUGH TO PROPERLY SUPPORT AND PROTECT UNIT. MOVABLE FLANGE HAS |  |  |  | JOSEPHM |  |  |  | Use As Is
106 | 2 |  |  | CRM1207061-2002 |  | RETURN BEAMPIPE | [45 empty] | C100-CM-INSP-RTBP |  | R1 |  | 3 | [5 empty] | PACKAGING IS NOT ENOUGH TO PROPERLY SUPPORT AND PROTECT UNIT. MOVABLE FLANGE HAS |  |  |  | JOSEPHM |  |  |  | Use As Is
107 | 2 |  |  | CRM1207061-2002 |  | RETURN BEAMPIPE | [45 empty] | C100-CM-INSP-RTBP |  | R1 |  | 3 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING. PACKAGING IS NOT ENOUGH TO PROPE |  |  |  | JOSEPHM |  |  |  | Use As Is
108 | 30 |  |  | 09-2787-4800 |  | HELIUM VESSEL ASSEMBLY | [45 empty] | C100-CM-INSP-HELV |  | R2 |  | 8 | [5 empty] | SEE ATTACHMENTS FOR DIMENSIONAL NON CONFORMING OF EACH COMPONENT. BELLOW 4840-0 |  |  |  | CHENG |  |  |  | Return To Vendor |  | Bellows 4840-061 has nonparallel end surfaces that makes it unusable. It needs t
109 | CRM-1207002-0000_001 |  |  | CRM1207002-0000 | [47 empty] | C100-CM-INSP-TPHTD |  | R1 |  | 2 | [5 empty] | 32 FINISH CHECKS:32,40,44,43 AND:37,39,31,36. SEE ATTACHMENT FOR DIMENSIONAL NO |  |  |  | KWILSON |  |  |  | Use As Is
110 | CRM-1207002-0001_002 |  |  | CRM1207002-0001 |  | TOP HAT SINGLE | [45 empty] | C100-CM-INSP-TPHTS |  | R1 |  | 3 | [5 empty] | 32 SURFACE REQUIREMENT OUT OF TOLERANCE:CHECKS:36,27,25,29. SEE ATTACHMENT FOR |  |  |  | KWILSON |  |  |  | Use As Is
111 | CRM-1207002-0001_004 |  |  | CRM1207002-0001 |  | TOP HAT SINGLE | [45 empty] | C100-CM-INSP-TPHTS |  | R1 |  | 4 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING |  |  |  | KWILSON |  |  |  | Return To Vendor |  | Returned to Meyer for rework
112 | CRM-1207002-0001_007 |  |  | CRM1207002-0001 |  | TOP HAT SINGLE | [45 empty] | C100-CM-INSP-TPHTS |  | R1 |  | 5 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING |  |  |  | KWILSON |  |  |  | Use As Is
113 | CRM-1207002-0001_008 |  |  | CRM1207002-0001 |  | TOP HAT SINGLE | [45 empty] | C100-CM-INSP-TPHTS |  | R1 |  | 6 | [5 empty] | 32 FINISH CHECKS OUT OF TOLERANCE.20,36,22,32. SEE ATTACHMENT FOR DIMENSIONAL N |  |  |  | KWILSON |  |  |  | Use As Is
114 | CRM-1207002-0001_009 |  |  | CRM1207002-0001 |  | TOP HAT SINGLE | [45 empty] | C100-CM-INSP-TPHTS |  | R1 |  | 7 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON CONFORMING |  |  |  | KWILSON |  |  |  | Use As Is
115 | CRM-1207002-0001_010 |  |  | CRM1207002-0001 |  | TOP HAT SINGLE | [45 empty] | C100-CM-INSP-TPHTS |  | R1 |  | 8 | [5 empty] | 32 SURFACE FINISH REQUIREMENT OUT OF TOLERANCE;29,31,30,37. SEE ATTACHMENT FOR |  |  |  | KWILSON |  |  |  | Use As Is
116 | MMT-004 |  |  | FAC-400-5656-1016 |  | WAVEGUIDE SWEEP | [45 empty] | C100-CM-INSP-SWEEP |  | R1 |  | 2 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NONCONFORMING. |  |  |  | KWILSON |  |  |  | Use As Is
117 | MMT-001 |  |  | FAC-400-5656-1016 |  | WAVEGUIDE SWEEP | [45 empty] | C100-CM-INSP-SWEEP |  | R1 |  | 3 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NONCONFORMING |  |  |  | KWILSON |  |  |  | Use As Is
118 | MMT-003 |  |  | FAC-400-5656-1016 |  | WAVEGUIDE SWEEP | [45 empty] | C100-CM-INSP-SWEEP |  | R1 |  | 4 | [5 empty] | TOOL MARKS ON LARGE FLANGE FACE.SEE ATTACHMENT FOR DIMENSIONAL NONCONFORMING. |  |  |  | KWILSON |  |  |  | Use As Is
119 | MMT-005 |  |  | FAC-400-5656-1016 |  | WAVEGUIDE SWEEP | [45 empty] | C100-CM-INSP-SWEEP |  | R1 |  | 5 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NONCONFORMING |  |  |  | KWILSON |  |  |  | Use As Is
120 | MMT-006 |  |  | FAC-400-5656-1016 |  | WAVEGUIDE SWEEP | [45 empty] | C100-CM-INSP-SWEEP |  | R1 |  | 6 | [5 empty] | UNUSUAL DISCOLORATION/SCALE INSIDE ON LARGE FLANGE END.SEE ATTACHMENT FOR DIMENS |  |  |  | KWILSON |  |  |  | Use As Is
121 | MMT-007 |  |  | FAC-400-5656-1016 |  | WAVEGUIDE SWEEP | [45 empty] | C100-CM-INSP-SWEEP |  | R1 |  | 7 | [5 empty] | SMALL AREA ON SMALL FLANGE MISSING ANODIZE COATING. SEE ATTACHMENT FOR DIMENSION |  |  |  | KWILSON |  |  |  | Use As Is
122 | MMT-008 |  |  | FAC-400-5656-1016 |  | WAVEGUIDE SWEEP | [45 empty] | C100-CM-INSP-SWEEP |  | R1 |  | 8 | [5 empty] | SMALL AREA OF DISCOLORATION/SCALE INSIDE AT LARGE FLANGE SIDE.SEE ATTACHMENT FOR |  |  |  | KWILSON |  |  |  | Use As Is
123 | MMT-009 |  |  | FAC-400-5656-1016 |  | WAVEGUIDE SWEEP | [45 empty] | C100-CM-INSP-SWEEP |  | R1 |  | 9 | [5 empty] | ROUGHNESS INTERNAL NEAR PORT HOLE. TOOL MARK ON FLANGE FACE.DISCOLORATION/SCALE |  |  |  | KWILSON |  |  |  | Use As Is
124 | MMT-018 |  |  | FAC-400-5656-1016 |  | WAVEGUIDE SWEEP | [45 empty] | C100-CM-INSP-SWEEP |  | R1 |  | 10 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NONCONFORMING. |  |  |  | KWILSON |  |  |  | Use As Is
125 | MMT-018A |  |  | FAC-400-5656-1016 |  | WAVEGUIDE SWEEP | [45 empty] | C100-CM-INSP-SWEEP |  | R1 |  | 11 | [5 empty] | METAL CHIP ADHERED TO INSIDE SURFACE NEAR PORT HOLE. ROUGHNESS AROUND EDGE OF P |  |  |  | KWILSON |  |  |  | Return To Vendor |  | Reject part due to gouges and spurs on inside surface.
126 | 1 |  |  | CRM-120-7043-0000 |  | Outer Magnetic Shield | [45 empty] | C100-CM-INSP-OMAG |  | R1 |  | 2 | [5 empty] | Please see attached spreadsheet for NCR info. |  |  |  | CHENG |  |  |  | Use As Is |  | These NCRs are found acceptable by doing a mock-up. There are 2 parts, one item
127 | CRM-1207002-0000_002 |  |  | CRM1207002-0000 |  | TOP HAT DOUBLE | [45 empty] | C100-CM-INSP-TPHTD |  | R1 |  | 3 | [5 empty] | SURFACE FINISH EXCEEDS REQUIREMENT ON FACE. MEASURED @ 36,42,46,43,36,31,38, 31 |  |  |  | KWILSON |  |  |  | Modify |  | If scratch is under seal, please polish. If not, ignore. Thanks.
128 | CRM-1207002-0001_001 |  |  | CRM1207002-0001 |  | TOP HAT SINGLE | [45 empty] | C100-CM-INSP-TPHTS |  | R1 |  | 9 | [5 empty] | SURFACE FINISH EXCEEDS REQUIREMENT:41,44,44,29. SCRATCH ON FACE. SEE ATTACHMEN |  |  |  | KWILSON |  |  |  | Use As Is
129 | CRM-1207002-0001_003 |  |  | CRM1207002-0001 |  | TOP HAT SINGLE | [45 empty] | C100-CM-INSP-TPHTS |  | R1 |  | 2 | [5 empty] | SURFACE FINISH EXCEEDS REQUIREMENT;26,32,37, 31. SCRATCH ON FACE. SEE ATTACH |  |  |  | KWILSON |  |  |  | Use As Is
130 | 3 |  |  | CRM-120-7060-2002 |  | RETURN BEAMPIPE | [45 empty] | C100-CM-INSP-RTBP |  | R1 |  | 4 | [5 empty] | SMALL DENT ON LARGE BELLOW. PACKAGING IS NOT ENOUGH TO PROPERLY SUPPORT THE HEA |  |  |  | JOSEPHM |  |  |  | Use As Is
131 | 4 |  |  | CRM-120-7060-2002 |  | RETURN BEAMPIPE | [45 empty] | C100-CM-INSP-RTBP |  | R1 |  | 5 | [5 empty] | SMALL DENT ON LARGE BELLOW.SMALL NICK ON KNIFE EDGE;LARGE FLANGE. PACKAGING IS |  |  |  | JOSEPHM |  |  |  | Modify |  | Knife edge can be repaired.
132 | 4 |  |  | CRM1207061-2001 |  | SUPPLY BEAMPIPE | [45 empty] | C100-CM-INSP-SUBP |  | R1 |  | 7 | [5 empty] | NICK ON KNIFE EDGE OF 2.74 DIAMETER FLANGE. SEE ATTACHMENT FOR DIMENSIONAL NON- |  |  |  | JOSEPHM |  |  |  | Modify |  | Knife edge can be repaired.
133 | 44 |  |  | CRM1207020-1000 |  | HELIUM VESSEL ASSEMBLY | [45 empty] | C100-CM-INSP-HELV |  | R2 |  | 9 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. |  |  |  | CHENG |  |  |  | Use As Is |  | The out of roundness of up to 0.050" on the diameters of bellows cuffs may cause
134 | MPF-0043 |  |  | CRM120-7014-0020 |  | HOM Probe | [45 empty] | C100-CAV-INSP-HMFT |  | R3 |  | 83 | [5 empty] | Center pin found loose during prep for assembly |  |  |  | HKPARK |  |  |  | Scrap
135 | R07 |  |  | CRM1207075_0041 |  | Guard Vacuum Waveguide | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 4 | [5 empty] | See step 11 in inspection traveler. |  |  |  | KWILSON |  |  |  | Use As Is |  | Use because we need part. Pictures sent to vendor.
136 | R06 |  |  | CRM1207075_0041 |  | Guard Vacuum Waveguide | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 3 | [5 empty] | Nick on knife edge. Has passed leak check. |  |  |  | KWILSON |  |  |  | Use As Is
137 | R05 |  |  | CRM1207075-0041 |  | Guard Vacuum Waveguide | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 6 | [5 empty] | See step 4 in traveler, weld inspection. |  |  |  | KWILSON |  |  |  | Use As Is |  | Use because we need part. Not acceptable on future parts.
138 | R04 |  |  | CRM-120-7075-0041 |  | Guard Vacuum Waveguide | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 5 | [5 empty] | Nick on knife edge |  |  |  | CARPENTE |  |  |  | Modify |  | repair knife edge damage. Repair complete.
139 | 6 |  |  | CRM-120-7090-1000 |  | Vacuum Vessel | [45 empty] | C100-CM-INSP-VV |  | R3 |  | 7 | [5 empty] | Loose item 8 & 9 are dirty. Light scratches and stains on tuner mount flanges. L |  |  |  | CHENG |  |  |  | Use As Is |  | We will return the stained items 8&9 subassemblies came as lose items. The fine
140 | L03 |  |  | CRM1207075-0040 |  | Waveguide Spool | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 9 | [5 empty] | Dings on knife edge and (1) hole needs re-tapping. |  |  |  | CARPENTE |  |  |  | Modify |  | Repair ding, re-tap hole and leak check. Complete on 04/08/11
141 | L05 |  |  | CRM1207075-0040 |  | Waveguide Spool | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 10 | [5 empty] | Nick on 2.75" conflat flange |  |  |  | CARPENTE |  |  |  | Modify |  | repaired.
142 | 2001-02 |  |  | 12070612001 |  | BEAM PIPE SUPPLY | [45 empty] | C100-CM-INSP-SUBP |  | R1 |  | 8 | [5 empty] | SMALL DENT ON LARGE BELLOW. NICK ON LARGE FLANGE KNIFE EDGE. SEE ATTACHMENT FO Unit is leaking AT THE DENT in the large bellows...verified by Mike McCrea |  |  |  | OGLE
143 | 2001-02 |  |  | 12070612001 |  | BEAM PIPE SUPPLY | [45 empty] | C100-CM-INSP-SUBP |  | R1 |  | 4 | [5 empty] | LEAK DETECTED AT DAMAGED BELLOW. |  |  |  | JOSEPHM |  |  |  | Return To Vendor |  | Part needs to be returned to vendor for rework. Passing leak check is a requirem
144 | 57 |  |  | 12070201000 |  | HELIUM VESSEL ASSEMBLY | [45 empty] | C100-CM-INSP-HELV |  | R2 |  | 10 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING |  |  |  | CHENG |  |  |  | Use As Is |  | Helium vessel parts from this vendor typically have out-of-roundness issues. We
145 | 59 |  |  | 12070201000 |  | HELIUM VESSEL ASSEMBLY | [45 empty] | C100-CM-INSP-HELV |  | R2 |  | 11 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING |  |  |  | CHENG |  |  |  | Use As Is |  | Helium vessel parts from this vendor typically have out-of-roundness issues. We
146 | 2001-05 |  |  | 12070612001 |  | BEAM PIPE SUPPLY | [45 empty] | C100-CM-INSP-SUBP |  | R1 |  | 9 | [5 empty] | NICK ON KNIFE EDGE;2.75 DIA. CON FLAT. LIGHT DENT ON LARGE BELLOW. |  |  |  | JOSEPHM |  |  |  | Modify |  | Knife edge repaired. Small dent acceptable.
147 | 5 |  |  | 12070612002 |  | BEAM PIPE RETURN | [45 empty] | C100-CM-INSP-RTBP |  | R1 |  | 6 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING |  |  |  | JOSEPHM |  |  |  | Use As Is
148 | 6 |  |  | 12070612002 |  | BEAM PIPE RETURN | [45 empty] | C100-CM-INSP-RTBP |  | R1 |  | 7 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. |  |  |  | JOSEPHM |  |  |  | Use As Is
149 | 2001-06 |  |  | 12070612001 |  | BEAM PIPE SUPPLY | [45 empty] | C100-CM-INSP-SUBP |  | R1 |  | 10 | [5 empty] | DENT ON LARGE BELLOW |  |  |  | JOSEPHM |  |  |  | Use As Is |  | Bellows passed leak check.
150 | 2 |  |  | CRM-120-7043-0000 |  | Outer Magnetic Shield | [45 empty] | C100-CM-INSP-OMAG |  | R1 |  | 3 | [5 empty] | Missing item 26 3" wide tape |  |  |  | CHENG |  |  |  | Use As Is |  | We have tapes needed for assembly.
151 | CRM1207002-0001-019 |  |  | 12070020001 |  | TOPHAT SINGLE | [45 empty] | C100-CM-INSP-TPHTS |  | R1 |  | 10 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING |  |  |  | KWILSON |  |  |  | Use As Is
152 | CRM1207002-0001-020 |  |  | 12070020001 |  | TOPHAT SINGLE | [45 empty] | C100-CM-INSP-TPHTS |  | R1 |  | 11 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING |  |  |  | KWILSON |  |  |  | Use As Is
153 | CRM1207002-0001-022 |  |  | 12070020001 |  | TOPHAT SINGLE | [45 empty] | C100-CM-INSP-TPHTS |  | R1 |  | 13 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING |  |  |  | KWILSON |  |  |  | Use As Is
154 | CRM1207002-0001-023 |  |  | 12070020001 |  | TOPHAT SINGLE | [45 empty] | C100-CM-INSP-TPHTS |  | R1 |  | 14 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING |  |  |  | KWILSON |  |  |  | Use As Is
155 | CRM1207002-0001-024 |  |  | 12070020001 |  | TOPHAT SINGLE | [45 empty] | C100-CM-INSP-TPHTS |  | R1 |  | 15 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING |  |  |  | KWILSON |  |  |  | Use As Is
156 | CRM1207002-0001-025 |  |  | 12070020001 |  | TOPHAT SINGLE | [45 empty] | C100-CM-INSP-TPHTS |  | R1 |  | 16 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING |  |  |  | KWILSON |  |  |  | Use As Is
157 | CRM1207002-0001-030 |  |  | 12070020001 |  | TOPHAT SINGLE | [45 empty] | C100-CM-INSP-TPHTS |  | R1 |  | 17 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING |  |  |  | KWILSON |  |  |  | Use As Is
158 | MMT-022 |  |  | FAC400-5656-1016 |  | WAVEGUIDE SWEEP | [45 empty] | C100-CM-INSP-SWEEP |  | R1 |  | 13 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING |  |  |  | KWILSON |  |  |  | Use As Is
159 | MMT-046 |  |  | FAC400-5656-1016 |  | WAVEGUIDE SWEEP | [45 empty] | C100-CM-INSP-SWEEP |  | R1 |  | 12 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING |  |  |  | KWILSON |  |  |  | Use As Is
160 | MMT-025 |  |  | FAC400-5656-1016 |  | WAVEGUIDE SWEEP | [45 empty] | C100-CM-INSP-SWEEP |  | R1 |  | 14 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING |  |  |  | KWILSON |  |  |  | Use As Is
161 | MMT-031 |  |  | FAC400-5656-1016 |  | WAVEGUIDE SWEEP | [45 empty] | C100-CM-INSP-SWEEP |  | R1 |  | 15 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING |  |  |  | KWILSON |  |  |  | Use As Is
162 | MMT-026 |  |  | FAC400-5656-1016 |  | WAVEGUIDE SWEEP | [45 empty] | C100-CM-INSP-SWEEP |  | R1 |  | 16 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING |  |  |  | KWILSON |  |  |  | Use As Is
163 | MMT-040 |  |  | FAC400-5656-1016 | [47 empty] | C100-CM-INSP-SWEEP |  | R1 |  | 17 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING |  |  |  | KWILSON |  |  |  | Use As Is
164 | MMT-027 |  |  | FAC400-5656-1016 |  | WAVEGUIDE SWEEP | [45 empty] | C100-CM-INSP-SWEEP |  | R1 |  | 18 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING |  |  |  | KWILSON |  |  |  | Use As Is
165 | MMT-047 |  |  | FAC400-5656-1016 |  | WAVEGUIDE SWEEP | [45 empty] | C100-CM-INSP-SWEEP |  | R1 |  | 19 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING |  |  |  | KWILSON |  |  |  | Use As Is
166 | MMT-035 |  |  | FAC400-5656-1016 |  | WAVEGUIDE SWEEP | [45 empty] | C100-CM-INSP-SWEEP |  | R1 |  | 20 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING |  |  |  | KWILSON |  |  |  | Use As Is
167 | MMT-049 |  |  | FAC400-5656-1016 |  | WAVEGUIDE SWEEP | [45 empty] | C100-CM-INSP-SWEEP |  | R1 |  | 21 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING |  |  |  | KWILSON |  |  |  | Use As Is
168 | 6 |  |  | CRM-120-7090-1000 |  | Vacuum Vessel | [45 empty] | C100-CM-INSP-VV |  | R3 |  | 8 | [5 empty] | Stains on one tuner mount flange. |  |  |  | CHENG |  |  |  | Return To Vendor |  | Communicated with the vendor and they want the stained flange assembly back for
169 | 9 |  |  | CRM-120-7042-0000 |  | Inner Mag Shield | [45 empty] | C100-CM-INSP-IMAG |  | R3 |  | 14 | [5 empty] | 0010-72-1 Half Shell Tab Side severely damaged |  |  |  | CHENG |  |  |  | Return To Vendor |  | Return the damaged part for repair and re-annealing.
170 | R13 |  |  | 12070750041 |  | WAVEGUIDE SPOOL RIGHT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 14 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING |  |  |  | KWILSON |  |  |  | Use As Is
171 | R18 |  |  | 12070750041 |  | WAVEGUIDE SPOOL RIGHT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 16 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING |  |  |  | KWILSON |  |  |  | Use As Is
172 | L09 |  |  | 12070750040 |  | WAVEGUIDE SPOOL LEFT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 18 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING |  |  |  | KWILSON |  |  |  | Use As Is
173 | L10 |  |  | 12070750040 |  | WAVEGUIDE SPOOL LEFT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 19 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING |  |  |  | KWILSON |  |  |  | Use As Is
174 | R20 |  |  | CRM-120-7075-0041 |  | WAVEGUIDE SPOOL RIGHT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 22 | [5 empty] | Diameter of flange thru holes out of tolerance. Please see pdf attached. |  |  |  | KWILSON |  |  |  | Use As Is
175 | 2001-07 |  |  | 12070612001 |  | BEAM PIPE SUPPLY | [45 empty] | C100-CM-INSP-SUBP |  | R1 |  | 11 | [5 empty] | NICKS ON KNIFE EDGE OF 5.975 DIA.CONFLAT. |  |  |  | JOSEPHM |  |  |  | Modify |  | Repair knife edge.
176 | 2001-08 |  |  | 12070612001 |  | BEAM PIPE SUPPLY | [45 empty] | C100-CM-INSP-SUBP |  | R1 |  | 12 | [5 empty] | NICKS ON KNIFE EDGE OF 5.975 DIA.CONFLAT. |  |  |  | JOSEPHM |  |  |  | Modify |  | Repair knife edge. | [31 empty] | knife edge repaired 6-13-11 by mccrea
177 | 2001-09 |  |  | 12070612001 |  | BEAM PIPE SUPPLY | [45 empty] | C100-CM-INSP-SUBP |  | R1 |  | 13 | [5 empty] | NICKS ON KNIFE EDGE OF 5.975 DIA.CONFLAT. SMALL DENT ON LARGE BELLOW |  |  |  | JOSEPHM |  |  |  | Modify |  | Repair knife edge ASAP to verify repair is feasible. | [31 empty] | knife edge repaired on 6-13 -11 by mccrea
178 | 2001-11 |  |  | 12070612001 |  | BEAM PIPE SUPPLY | [45 empty] | C100-CM-INSP-SUBP |  | R1 |  | 15 | [5 empty] | NICKS ON KNIFE EDGE OF 5.975 DIA.CONFLAT. |  |  |  | JOSEPHM |  |  |  | Modify |  | Repair knife edge ASAP to verify repair is feasible. | [31 empty] | knife edge repaired on 6-13-11 by mccrea
179 | 2002-07 |  |  | 12070612002 |  | BEAM PIPE RETURN | [45 empty] | C100-CM-INSP-RTBP |  | R1 |  | 8 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING |  |  |  | JOSEPHM |  |  |  | Use As Is
180 | 2002-08 |  |  | 12070612002 |  | BEAM PIPE RETURN | [45 empty] | C100-CM-INSP-RTBP |  | R1 |  | 9 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING |  |  |  | JOSEPHM |  |  |  | Use As Is
181 | 2002-09 |  |  | 12070612002 |  | BEAM PIPE RETURN | [45 empty] | C100-CM-INSP-RTBP |  | R1 |  | 10 | [5 empty] | SMALL DENT ON BELLOW. NICK ON KNIFE EDGE OF 2.74 DIAMETER CONFLAT. SEE ATTACHM |  |  |  | JOSEPHM |  |  |  | Modify |  | Repair knife edge. | [31 empty] | knife edge repaired on 6-13-11 by mccrea
182 | 2002-10 |  |  | 12070612002 |  | BEAM PIPE RETURN | [45 empty] | C100-CM-INSP-RTBP |  | R1 |  | 11 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING |  |  |  | JOSEPHM |  |  |  | Use As Is
183 | 2002-11 |  |  | 12070612002 |  | BEAM PIPE RETURN | [45 empty] | C100-CM-INSP-RTBP |  | R1 |  | 12 | [5 empty] | KNIFE EDGE NICKS ON 2.74 DIAMETER CONFLAT SEE ATTACHMENT FOR DIMENSIONAL NON-CO |  |  |  | JOSEPHM |  |  |  | Modify |  | Repair knife edge. | [31 empty] | knife edge repaired 6-13-11 by mccrea
184 | L12 |  |  | CRM1207075-0042 |  | WG Spool | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 20 | [5 empty] | Bad spot in interior of wg. see pics. M. Wiseman recommendedReturn To Vendor. |  |  |  | KWILSON |  |  |  | Return To Vendor |  | Please package carefully for return. Waiting for return instructions from vendo
185 | 68 |  |  | 927874800 |  | HELIUM VESSEL ASSEMBLY | [45 empty] | C100-CM-INSP-HELV |  | R2 |  | 12 | [5 empty] | BELLOW;4840-136 SMALL DENT ON CONVOLUTIONS INSIDE SURFACE. BELLOW;4840-135 SMAL |  |  |  | CHENG |  |  |  | Use As Is |  | All bellows will be checked for imperfections and some, like the two bellows for
186 | 78 |  |  | 927874800 |  | HELIUM VESSEL ASSEMBLY | [45 empty] | C100-CM-INSP-HELV |  | R2 |  | 13 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING |  |  |  | CHENG |  |  |  | Use As Is |  | There is a bellows having significant out-of-roundness inner diameter on one end
187 | 83 |  |  | 927874800 |  | HELIUM VESSEL ASSEMBLY | [45 empty] | C100-CM-INSP-HELV |  | R2 |  | 14 | [5 empty] | SMALL DENT ON BELLOW CONVOLUTIONS. SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMIN |  |  |  | CHENG |  |  |  | Use As Is |  | Besides the bellows dent issue, the rest looks fine. We will check all bellows t
189 | CRM12070020001-041 |  |  | 12070020001 |  | TOPHAT SINGLE | [45 empty] | C100-CM-INSP-TPHTS |  | R1 |  | 18 | [5 empty] | LIGHT SCRATCHES ON FLANGE FACE WITHIN SEAL AREA SEE PHOTO. 10.13 ID NOT MEASURE |  |  |  | KWILSON |  |  |  | Use As Is
190 | CRM1207002-0000-007 |  |  | 12070020000 |  | TOPHAT DOUBLE | [45 empty] | C100-CM-INSP-TPHTD |  | R1 |  | 6 | [5 empty] | 13.310 FLANGE ID NOT MEASURED. FLANGE PRESSED ON TOO FAR AND NOT STRAIGHT. UN-AB |  |  |  | KWILSON |  |  |  | Use As Is
191 | 2 |  |  | CRM-120-7043-0030 |  | Secondary Outer Mag | [45 empty] | C100-CM-INSP-OMAGS |  | R1 |  | 3 | [5 empty] | Please see attached PDF |  |  |  | CHENG |  |  |  | Modify |  | The design for -0028 and -0029 have been changed. Internal rework is necessary o
192 | CRM12070020001-049 |  |  | 12070020001 |  | TOPHAT SINGLE | [45 empty] | C100-CM-INSP-TPHTS |  | R1 |  | 19 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING |  |  |  | KWILSON |  |  |  | Use As Is
193 | 52 |  |  | CRM1207020-1000 |  | Helium Vessel Assembly | [45 empty] | C100-CM-INSP-HELV |  | R2 |  | 15 | [5 empty] | Bellows inspection only in this traveler, s/n 4840-104 Dents in convolutions |  |  |  | CHENG |  |  |  | Return To Vendor |  | The inspected bellows 4840-104 is judged to be unusable.
194 | 65 |  |  | CRM1207020-1000 |  | Helium Vessel Assembly | [45 empty] | C100-CM-INSP-HELV |  | R2 |  | 16 | [5 empty] | Bellows inspection only in this traveler, s/n 4840-130 Dents in convolutions |  |  |  | CHENG |  |  |  | Return To Vendor |  | The bellows 4840-130 is determined to be unusable.
195 | 67 |  |  | CRM1207020-1000 |  | Helium Vessel Assembly | [45 empty] | C100-CM-INSP-HELV |  | R2 |  | 17 | [5 empty] | Bellows inspection only in this traveler, s/n 4840-133 Dents in convolutions |  |  |  | CHENG |  |  |  | Return To Vendor |  | Bellows 4840-133 is determined to be unusable.
196 | 67 |  |  | CRM1207020-1000 |  | Helium Vessel Assembly | [45 empty] | C100-CM-INSP-HELV |  | R2 |  | 18 | [5 empty] | Bellows inspection only in this traveler, s/n 4840-134 Dents in convolutions |  |  |  | CHENG |  |  |  | Return To Vendor |  | Bellows 4840-134 is determined to be unusable.
197 | 72 |  |  | CRM1207020-1000 |  | Helium Vessel Assembly | [45 empty] | C100-CM-INSP-HELV |  | R2 |  | 19 | [5 empty] | Bellows inspection only in this traveler, s/n 4840-143 Dents in convolutions |  |  |  | CHENG |  |  |  | Return To Vendor |  | Bellows 4840-143 is determined to be unusable.
198 | 72 |  |  | CRM1207020-1000 |  | Helium Vessel Assembly | [45 empty] | C100-CM-INSP-HELV |  | R2 |  | 20 | [5 empty] | Bellows inspection only in this traveler, s/n 4840-144 Dents in convolutions |  |  |  | CHENG |  |  |  | Return To Vendor |  | Bellows 4840-144 is determined to be unusable.
199 | 79 |  |  | CRM1207020-1000 |  | Helium Vessel Assembly | [45 empty] | C100-CM-INSP-HELV |  | R2 |  | 21 | [5 empty] | Bellows inspection only in this traveler, s/n 4840-157 Dents in convolutions |  |  |  | CHENG |  |  |  | Return To Vendor |  | Bellows 4840-157 is determined to be unusable.
200 | 87 |  |  | CRM1207020-1000 |  | Helium Vessel Assembly | [45 empty] | C100-CM-INSP-HELV |  | R2 |  | 22 | [5 empty] | Bellows inspection only in this traveler, s/n 4840-173 Dents in convolutions |  |  |  | CHENG |  |  |  | Return To Vendor |  | Bellows 4840-173 is determined to be unusable.
201 | 90 |  |  | CRM1207020-1000 |  | Helium Vessel Assembly | [45 empty] | C100-CM-INSP-HELV |  | R2 |  | 23 | [5 empty] | Bellows inspection only in this traveler, s/n 4840-179 Scratches in convolution |  |  |  | CHENG |  |  |  | Return To Vendor |  | Bellows 4840-179 is determined to be unusable.
207 | 3 |  |  | CRM-120-7043-0000 |  | Primary Outer Mag Shield | [45 empty] | C100-CM-INSP-OMAG |  | R2 |  | 5 | [5 empty] | Missing Item #26 3M tape |  |  |  | CHENG |  |  |  | Use As Is
210 | 4840-072 |  |  | 12070200007 |  | BELLOW | [55 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING |  |  |  | CHENG
211 | MMT-053 |  |  | FAC400-5656-1016 |  | WAVEGUIDE SWEEP | [45 empty] | C100-CM-INSP-SWEEP |  | R1 |  | 22 | [5 empty] | SMALL GOUGE NEAR PORT INTERIOR SURFACE.SCRATCH ON INTERIOR SURFACE. SEE ATTACHM |  |  |  | KWILSON |  |  |  | Use As Is
212 | MMT-063 |  |  | FAC400-5656-1016 |  | WAVEGUIDE SWEEP | [45 empty] | C100-CM-INSP-SWEEP |  | R1 |  | 23 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. |  |  |  | KWILSON |  |  |  | Use As Is
213 | MMT-067 |  |  | FAC400-5656-1016 |  | WAVEGUIDE SWEEP | [45 empty] | C100-CM-INSP-SWEEP |  | R1 |  | 24 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. |  |  |  | KWILSON |  |  |  | Use As Is
214 | MMT-071 |  |  | FAC400-5656-1016 |  | WAVEGUIDE SWEEP | [45 empty] | C100-CM-INSP-SWEEP |  | R1 |  | 25 | [5 empty] | VERY LIGHT SCRATCH ON INTERIOR SURFACE. SEE ATTACHMENT FOR DIMENSIONAL NON-CONF |  |  |  | KWILSON |  |  |  | Use As Is
215 | MMT-080 |  |  | FAC400-5656-1016 |  | WAVEGUIDE SWEEP | [45 empty] | C100-CM-INSP-SWEEP |  | R1 |  | 26 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. |  |  |  | KWILSON |  |  |  | Use As Is
216 | MMT-018A |  |  | FAC400-5656-1016 |  | WAVEGUIDE SWEEP | [45 empty] | C100-CM-INSP-SWEEP |  | R1 |  | 27 | [5 empty] | NOTE THIS PART WAS RETURNED TO THE VENDOR FOR REWORK. SLIGHT INDENTION AND ROUG |  |  |  | KWILSON |  |  |  | Use As Is
217 | L15 |  |  | 12070750040 |  | WAVEGUIDE SPOOL LEFT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 25 | [5 empty] | SPOTS ON COPPER AND NON-COPPER SURFACE OF FLANGE FACE. LIGHT DISCOLORATION ON I |  |  |  | KWILSON |  |  |  | Use As Is
218 | L20 |  |  | 12070750040 |  | WAVEGUIDE SPOOL LEFT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 26 | [5 empty] | BLACK RESIDUE ON EXTERIOR SURFACE. SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMIN |  |  |  | KWILSON |  |  |  | Modify |  | Dimensions ok. Please see if black stuff can be cleaned off. Spots would not come off. Use as is.
219 | L27 |  |  | 12070750040 |  | WAVEGUIDE SPOOL LEFT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 27 | [5 empty] | LIGHT DISCOLORATION OF COPPER SURFACES. SEE ATTACHMENT FOR DIMENSIONAL NON-CONF |  |  |  | KWILSON |  |  |  | Use As Is
220 | R01 |  |  | 12070750041 |  | WAVEGUIDE SPOOL RIGHT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 28 | [5 empty] | SLIGHT DISCOLORATION INTERNAL COPPER SURFACE. STAIN SPOTS ON EXTERIOR SURFACE O |  |  |  | KWILSON |  |  |  | Modify |  | Dimensions ok. Please clean to remove residue. Spots would not come off. Use as is.
221 | R02 |  |  | 12070750041 |  | WAVEGUIDE SPOOL RIGHT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 29 | [5 empty] | CONTAMINATION ON 2.75 CONFLAT IN REGION OF KNIFE EDGE. DOES NOT APPEAR TO BE ON |  |  |  | KWILSON |  |  |  | Use As Is |  | Dimensions ok. Please clean to remove residue.
222 | R11 |  |  | 12070750041 |  | WAVEGUIDE SPOOL RIGHT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 30 | [5 empty] | DISCOLORATION ON INTERIOR COPPER SURFACE. |  |  |  | KWILSON |  |  |  | Use As Is
223 | R15 |  |  | 12070750041 |  | WAVEGUIDE SPOOL RIGHT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 31 | [5 empty] | SMALL NICK ON KNIFE EDGE. |  |  |  | KWILSON |  |  |  | Modify |  | Please repair nick in knife edge.
224 | R17 |  |  | 12070750041 |  | WAVEGUIDE SPOOL RIGHT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 32 | [5 empty] | WELD SLAG IN SMALL PORT TUBE. BLACK RESIDUE ON COPPER SURFACE AT FLANGE FACE. U |  |  |  | KWILSON |  |  |  | Use As Is
225 | 4840-072 |  |  | 12070200007 |  | BELLOW | [55 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. |  |  |  | CHENG |  |  |  | Return To Vendor |  | The bellows has parallelism of 0.099", which is intolerable.
226 | 4840-065 |  |  | 12070200007 |  | BELLOW | [55 empty] | REWORKED BELLOW: SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. |  |  |  | CHENG
227 | 4840-065 |  |  | 12070200007 |  | BELLOW | [55 empty] | REWORKED BELLOW: SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. |  |  |  | CHENG |  |  |  | Use As Is |  | Parallelism is 0.070"
228 | 4840-061 |  |  | 12070200007 |  | BELLOW | [55 empty] | REWORKED BELLOW: SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING |  |  |  | CHENG |  |  |  | Return To Vendor |  | Parallelism is 0.152", which is intolerable.
229 | 183 |  |  | 12070200025 |  | HELIUM VESSEL HEAD ASSEMBLY | [45 empty] | C100-CM-INSP-HVHD |  | R1 |  | 36 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING |  |  |  | CHENG |  |  |  | Use As Is |  | The out-of-tolerance values are minor.
230 | 184 |  |  | 12070200025 |  | HELIUM VESSEL HEAD ASSEMBLY | [45 empty] | C100-CM-INSP-HVHD |  | R1 |  | 37 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING |  |  |  | CHENG |  |  |  | Use As Is |  | The out-of-tolerance values are minor.
231 | 185 |  |  | 12070200025 |  | HELIUM VESSEL HEAD ASSEMBLY | [45 empty] | C100-CM-INSP-HVHD |  | R1 |  | 38 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. |  |  |  | CHENG |  |  |  | Use As Is |  | The out-of-tolerance values are minor.
232 | 186 |  |  | 12070200025 |  | HELIUM VESSEL HEAD ASSEMBLY | [45 empty] | C100-CM-INSP-HVHD |  | R1 |  | 39 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. |  |  |  | CHENG |  |  |  | Use As Is |  | The out-of-tolerance values are minor.
233 | 187 |  |  | 12070200025 |  | HELIUM VESSEL HEAD ASSEMBLY | [45 empty] | C100-CM-INSP-HVHD |  | R1 |  | 40 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. |  |  |  | CHENG |  |  |  | Use As Is |  | The out-of-tolerance values are minor.
234 | 188 |  |  | 12070200025 |  | HELIUM VESSEL HEAD ASSEMBLY | [45 empty] | C100-CM-INSP-HVHD |  | R1 |  | 41 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. |  |  |  | CHENG |  |  |  | Use As Is |  | The out-of-roundness value of 0.044" caused attention. Tried the fit with a bell
235 | 189 |  |  | 12070200025 |  | HELIUM VESSEL HEAD ASSEMBLY | [45 empty] | C100-CM-INSP-HVHD |  | R1 |  | 42 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. |  |  |  | CHENG |  |  |  | Use As Is |  | The out-of-tolerance values are minor.
236 | 190 |  |  | 12070200025 |  | HELIUM VESSEL HEAD ASSEMBLY | [45 empty] | C100-CM-INSP-HVHD |  | R1 |  | 43 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. |  |  |  | CHENG |  |  |  | Use As Is |  | The out-of-tolerance values are minor.
237 | 191 |  |  | 12070200025 |  | HELIUM VESSEL HEAD ASSEMBLY | [45 empty] | C100-CM-INSP-HVHD |  | R1 |  | 44 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. |  |  |  | CHENG |  |  |  | Use As Is |  | The out-of-tolerance values are minor.
238 | 192 |  |  | 12070200025 |  | HELIUM VESSEL HEAD ASSEMBLY | [45 empty] | C100-CM-INSP-HVHD |  | R1 |  | 45 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. |  |  |  | CHENG |  |  |  | Use As Is |  | The out-of-tolerance values are minor.
239 | 193 |  |  | 12070200025 |  | HELIUM VESSEL HEAD ASSEMBLY | [45 empty] | C100-CM-INSP-HVHD |  | R1 |  | 46 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. |  |  |  | CHENG |  |  |  | Use As Is |  | The out-of-tolerance values are minor.
240 | 194 |  |  | 12070200025 |  | HELIUM VESSEL HEAD ASSEMBLY | [45 empty] | C100-CM-INSP-HVHD |  | R1 |  | 47 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. |  |  |  | CHENG |  |  |  | Use As Is |  | The out-of-tolerance values are minor.
241 | 195 |  |  | 12070200025 |  | HELIUM VESSEL HEAD ASSEMBLY | [45 empty] | C100-CM-INSP-HVHD |  | R1 |  | 48 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. |  |  |  | CHENG |  |  |  | Use As Is |  | The out-of-tolerance values are minor.
242 | 196 |  |  | 12070200025 |  | HELIUM VESSEL HEAD ASSEMBLY | [45 empty] | C100-CM-INSP-HVHD |  | R1 |  | 49 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. |  |  |  | CHENG |  |  |  | Use As Is |  | The out-of-tolerance values are minor.
243 | 197 |  |  | 12070200025 |  | HELIUM VESSEL HEAD ASSEMBLY | [45 empty] | C100-CM-INSP-HVHD |  | R1 |  | 50 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. |  |  |  | CHENG |  |  |  | Use As Is |  | The out-of-tolerance values are minor.
244 | 198 |  |  | 12070200025 |  | HELIUM VESSEL HEAD ASSEMBLY | [45 empty] | C100-CM-INSP-HVHD |  | R1 |  | 51 | [5 empty] | SEE ATTACHMENT FOR DIMENSIONAL NON-CONFORMING. |  |  |  | CHENG |  |  |  | Use As Is |  | The out-of-tolerance values are minor.
245 | 98 |  |  | 927874800 |  | HELIUM VESSEL ASSEMBLY | [45 empty] | C100-CM-INSP-HELV |  | R2 |  | 65 | [5 empty] | BELLOW 4840-197 SMALL DENT IN CONVOLUTION. SEE ATTACHMENT FOR DIMENSIONAL NON-C |  |  |  | CHENG |  |  |  | Modify |  | Both bellows have large out-of-roundness numbers, as large as 0.057". This has b
246 | 2787-094 |  |  | CRM12070201000 |  | HELIUM VESSEL ASSEMBLY | [45 empty] | C100-CM-INSP-HELV |  | R2 |  | 69 | [5 empty] | BELLOW INSPECTION ONLY IN THIS TRAVELER,SN4840-188 DENTS IN CONVOLUTIONS. |  |  |  | CHENG |  |  |  | Return To Vendor |  | There are 3 dents that are deep and created obvious bumps on the reverse side of
247 | 3 |  |  | CRM-120-7043-0030 |  | Secondary Outer Mag Shield | [45 empty] | C100-CM-INSP-OMAGS |  | R1 |  | 4 | [5 empty] | Tophat single#2 inside diameter out of tolerance. Nominal 10.53" Measured 10.62 |  |  |  | CHENG |  |  |  | Use As Is |  | The slightly bigger tophat single secondary shield eats up clearance between th
248 | 2001-02 |  |  | 12070612001 |  | BEAM PIPE SUPPLY | [45 empty] | C100-CM-INSP-SUBP |  | R1 |  | 16 | [5 empty] | Multiple dents in large bellow. Nick in large con flat knife edge. See attachme |  |  |  | JOSEPHM | [5 empty] | Knife edge repaired. Small dents acceptable. | [31 empty] | repaired knife edge c wilcox Verified Knife edge repair. 1/17/12 J.O.
249 | 2002-02 |  |  | 12070612002 |  | BEAM PIPE RETURN | [45 empty] | C100-CM-INSP-RTBP |  | R1 |  | 13 | [5 empty] | Nick on large con flat knife edge. See attachment for dimensional non-conforming |  |  |  | JOSEPHM |  |  |  | Modify |  | Repair knife edge.
250 | 4 |  |  | CRM-120-7043-0030 |  | Secondary Outer Mag Shield | [45 empty] | C100-CM-INSP-OMAGS |  | R1 |  | 5 | [5 empty] | Tophat single #1 and #2 inside diameter out of tolerance. Both Tophats ID measur |  |  |  | CHENG |  |  |  | Use As Is |  | The primary shield parts that go inside the secondary tophat single have a nomin
251 | L17 |  |  | 12070750040 |  | WAVEGUIDE SPOOL LEFT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 34 | [5 empty] | See attachment for dimensional non-conforming. Blister and large gouge inside o |  |  |  | KWILSON |  |  |  | Return To Vendor
252 | L28 |  |  | 12070750040 |  | WAVEGUIDE SPOOL LEFT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 35 | [5 empty] | See attachment for dimensional non-conforming. |  |  |  | KWILSON |  |  |  | Use As Is
253 | L16 |  |  | 12070750040 |  | WAVEGUIDE SPOOL LEFT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 36 | [5 empty] | See attachment for dimensional non-conforming. |  |  |  | KWILSON |  |  |  | Use As Is
254 | L18 |  |  | 12070750040 |  | WAVEGUIDE SPOOL LEFT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 37 | [5 empty] | Nick on knife edge. See attachment for dimensional non-conforming. |  |  |  | KWILSON |  |  |  | Use As Is |  | Please repair knife edge. Dimensions ok.
255 | L19 |  |  | 12070750040 |  | WAVEGUIDE SPOOL LEFT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 38 | [5 empty] | Knife edge on angled con flat shows unusual surface texture. |  |  |  | KWILSON |  |  |  | Modify |  | Please verify small conflat is leak tight during leak test.
256 | R26 |  |  | 12070750041 |  | WAVEGUIDE SPOOL RIGHT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 39 | [5 empty] | See attachment for dimensional non-conforming. |  |  |  | KWILSON |  |  |  | Use As Is
257 | R27 |  |  | 12070750041 |  | WAVEGUIDE SPOOL RIGHT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 40 | [5 empty] | Weld splatter on inside diameter of angled port. |  |  |  | KWILSON |  |  |  | Use As Is
258 | R23 |  |  | 12070750041 |  | WAVEGUIDE SPOOL RIGHT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 42 | [5 empty] | See attachment for dimensional non-conforming. |  |  |  | KWILSON |  |  |  | Use As Is
259 | R16 |  |  | 12070750041 |  | WAVEGUIDE SPOOL RIGHT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 43 | [5 empty] | See attachment for dimensional non-conforming. |  |  |  | KWILSON |  |  |  | Use As Is
260 | R40 |  |  | 12070750041 |  | WAVEGUIDE SPOOL RIGHT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 44 | [5 empty] | See attachment for dimensional non-conforming. |  |  |  | KWILSON |  |  |  | Use As Is
261 | R32 |  |  | 12070750041 |  | WAVEGUIDE SPOOL RIGHT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 46 | [5 empty] | See attachment for dimensional non-conforming. |  |  |  | KWILSON |  |  |  | Use As Is
262 | R33 |  |  | 12070750041 |  | WAVEGUIDE SPOOL RIGHT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 47 | [5 empty] | Spot of copper plating on conflat surface. Does not appear to be on apex of knif |  |  |  | KWILSON |  |  |  | Use As Is
263 | R37 |  |  | 12070750041 |  | WAVEGUIDE SPOOL RIGHT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 51 | [5 empty] | See attachment for dimensional non-conforming. |  |  |  | KWILSON |  |  |  | Use As Is
264 | R41 |  |  | CRM-120-7075-0041 |  | WAVEGUIDE SPOOL RIGHT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 55 | [5 empty] | Bolts are tight to go in with test plate for counter sink positions; 2 positions |  |  |  | KWILSON |  |  |  | Modify |  | Please retap holes that need it. Thanks.
265 | R21 |  |  | CRM-120-7075-0041 |  | WAVEGUIDE SPOOL RIGHT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 57 | [5 empty] | Small nick in flange knife edge. |  |  |  | KWILSON |  |  |  | Modify |  | Please repair nick in knife edge.
266 | L24 |  |  | 12070750040 |  | WAVEGUIDE SPOOL LEFT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 58 | [5 empty] | Weld protrusion and roughness on small port inside diameter. Small nick on knif |  |  |  | KWILSON |  |  |  | Use As Is |  | Pls repair knife edge. Weld ok.
267 | L25 |  |  | 12070750040 |  | WAVEGUIDE SPOOL LEFT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 59 | [5 empty] | Weld protrusion in small port inside diameter. |  |  |  | KWILSON |  |  |  | Use As Is
268 | R45 |  |  | CRM-120-7075-0041 |  | WAVEGUIDE SPOOL RIGHT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 61 | [5 empty] | Tapped holes for countersink positions require excess force to thread bolts to c |  |  |  | KWILSON |  |  |  | Modify |  | Please retap holes that need it. Thanks.
269 | R42 |  |  | CRM-120-7075-0041 |  | WAVEGUIDE SPOOL RIGHT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 62 | [5 empty] | Tapped holes for countersink positions require excess force to thread bolts to c |  |  |  | KWILSON |  |  |  | Modify |  | Please retap holes that need it. Thanks.
270 | R36 |  |  | CRM-120-7075-0041 |  | WAVEGUIDE SPOOL RIGHT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 63 | [5 empty] | Black spot;blister? on copper surface. Small nick on or near the apex of the kn |  |  |  | KWILSON |  |  |  | Modify |  | Pls touch up knife edge.
271 | C100-04 |  |  |  |  | HELIUM VESSEL BELLOWS | [45 empty] | C100-CM-ASSY-COLD |  | R3 |  | 10 | [5 empty] | HELIUM VESSEL BELLOWS LEAKS, HELIUM VESSEL KIT 035, CAVITY SN 027. STRING WILL BE REWORKED WITH NEW CAV-HV IN ITS PLACE |  |  |  | FISCHER |  |  |  | Modify |  | STRING WILL BE SENT BACK TO CLEANROOM FOR REPLACEMENT OF LEAKY HELIUM VESSEL BEL
272 | L35 |  |  | 12070750040 |  | WAVEGUIDE SPOOL LEFT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 64 | [5 empty] | See attachment for dimensional non-conforming. |  |  |  | KWILSON |  |  |  | Use As Is
273 | L31 |  |  | 12070750040 |  | WAVEGUIDE SPOOL LEFT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 65 | [5 empty] | Missing area of plating in small port inside diameter. |  |  |  | KWILSON |  |  |  | Use As Is
274 | BC-02138-09 |  |  |  |  | Return end can | [45 empty] | C100-CM-INSP-RENC |  | R1 |  | 11 | [5 empty] | The Primary return bayonet is extremely tight going in, and will not seat at the |  |  |  | FHUMPHRY |  |  |  | Return To Vendor |  | The vendor completed repair of all end can bayonet seats on site. Conformance was verified via test fitting dummy bayonet into all seats by lead JLab technician. | [31 empty] | vendor on site to hone bayonets. Problems fixed.
275 | R35 |  |  | CRM-120-7075-0041 |  | WAVEGUIDE SPOOL RIGHT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 49 | [5 empty] | Black residue at internal weld of angled conflat tube. Will not come off with ac |  |  |  | KWILSON |  |  |  | Use As Is
276 | L29 |  |  | 12070750040 |  | WAVEGUIDE SPOOL LEFT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 53 | [5 empty] | Small surface spot seen on copper plating. |  |  |  | KWILSON |  |  |  | Use As Is
277 | R37 |  |  | 12070750041 |  | WAVEGUIDE SPOOL RIGHT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 51 | [5 empty] | Black residue at internal weld of angled conflat tube. Will not come off with ac |  |  |  | KWILSON |  |  |  | Use As Is
278 | 4 |  |  | CRM12070430000 |  | Outer Mag Shield Primary | [45 empty] | C100-CM-INSP-OMAG |  | R2 |  | 6 | [5 empty] | Missing item 26 3" tape |  |  |  | CHENG |  |  |  | Use As Is
279 | R22 |  |  | 12070750040 |  | WAVEGUIDE SPOOL RIGHT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 67 | [5 empty] | Spot of discoloration inside small port inside diameter. See attachment for dim |  |  |  | KWILSON |  |  |  | Use As Is
280 | L43 |  |  | 12070750040 |  | WAVEGUIDE SPOOL LEFT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 69 | [5 empty] | Small spot of weld on inside diameter of small port tube. See attachment for di |  |  |  | KWILSON |  |  |  | Use As Is
281 | 2787-094 |  |  | 12070201000 |  | HELIUM VESSEL ASSEMBLY | [45 empty] | C100-CM-INSP-HELV |  | R2 |  | 71 | [5 empty] | This traveler is for Bellow serial number 4840-188 only. Rejected per earlier in |  |  |  | CHENG |  |  |  | Use As Is |  | The parallelism between planes of cuffs for bellows 4840-188 is now acceptable.
282 | 2787-030 |  |  | 12070201000 |  | HELIUM VESSEL ASSEMBLY | [45 empty] | C100-CM-INSP-HELV |  | R2 |  | 72 | [5 empty] | This traveler is for Bellow serial number 4840-061 only. Rejected per earlier in |  |  |  | CHENG |  |  |  | Use As Is |  | The parallelism between planes of cuffs for bellows 4840-061 is now acceptable.
283 | 2787-037 |  |  | 12070201000 |  | HELIUM VESSEL ASSEMBLY | [45 empty] | C100-CM-INSP-HELV |  | R2 |  | 73 | [5 empty] | This traveler is for Bellow serial number 4840-072 only. Rejected per earlier in |  |  |  | CHENG |  |  |  | Use As Is |  | The parallelism between planes of cuffs for bellows 4840-072 is now acceptable.
284 | R08 |  |  | 12070750041 |  | WAVEGUIDE SPOOL RIGHT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 70 | [5 empty] | See attachment for dimensional non conforming. |  |  |  | KWILSON |  |  |  | Use As Is
285 | R10 |  |  | 12070750041 |  | WAVEGUIDE SPOOL RIGHT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 71 | [5 empty] | See attachment for dimensional non conforming. |  |  |  | KWILSON |  |  |  | Use As Is
286 | L45 |  |  | 12070750040 |  | WAVEGUIDE SPOOL LEFT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 72 | [5 empty] | Some threaded holes difficult to thread sample bolt to full depth with inspectio |  |  |  | KWILSON |  |  |  | Modify |  | Please run tap through 2 holes that are tight.
287 | L01 |  |  | 12070750040 |  | WAVEGUIDE SPOOL LEFT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 73 | [5 empty] | See attachment for dimensional non conforming. |  |  |  | KWILSON |  |  |  | Use As Is
288 | 4840-063 |  |  |  |  | He Vessel Bellows | [55 empty] | Bellows has damage to convolution, also has a -6 range leak. |  |  |  | CHENG |  |  |  | Return To Vendor |  | This bellows cannot be used.
289 | R15 |  |  | 12070750041 |  | Waveguide Spool Right | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 31 | [5 empty] | The purpose of this NCR is to capture the rework process for NCR TRAV_SEQ_NUM: 2 |  |  |  | GEORGED |  |  |  | Use As Is |  | Repair small nick on knife edge.
295 | R21 | GVWGSN |  | CRM1207075-0041 |  | WAVEGUIDE SPOOL - RIGHT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 57 |  | GEORGED |  | 2011-10-25 |  | The purpose of this NCR is to capture the rework data for NCR TRAV_SEQ_NUM: 265. |  |  |  | GEORGED |  | 2011-10-25 |  | Modify |  | Repair nick on the knife edge. |  |  |  | WILCOX |  | 2011-10-25 |  | Knife edge repaired. |  |  |  |  | GEORGED |  | 2011-10-25 |  | Knife edge okay. |  |  | GEORGED |  | 2011-10-25 |  | Use As Is |  | Knife edge is acceptable, part okay to use.
296 | R08 | GVWGSN |  | CRM1207075-0041 |  | WAVEGUIDE SPOOL - RIGHT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 70 |  | GEORGED |  | 2011-10-31 |  | Rough edge/divot on copper plating. Copper deteriorated on face of flange. |  | 8A8CCE91-AD2C-93C9-A9E6FFD4BABFECB9 |  | KWILSON |  | 2012-01-20 |  | Use As Is
297 | 67 | HESSN |  | CRM1207020-1000 |  | HELIUM VESSEL ASSEMBLY (RH) | [45 empty] | C100-CM-INSP-HELV |  | R2 |  | 74 |  | OGLE |  | 2011-10-31 |  | This traveler is for Bellow serial number 4840-134 only. Rejected per earlier inspection and dispositioned as return to vendor.Attached file is for dimensional nonconforming after "rework". |  | 9CE17F70-A8F4-5686-385BDE9877237847 |  | CHENG |  | 2011-11-01 |  | Use As Is
298 | 79 | HESSN |  | CRM1207020-1000 |  | HELIUM VESSEL ASSEMBLY (RH) | [45 empty] | C100-CM-INSP-HELV |  | R2 |  | 75 |  | OGLE |  | 2011-10-31 |  | This traveler is for Bellow serial number 4840-157 only. Rejected per earlier inspection and dispositioned as return to vendor.Attached file is for dimensional non conforming after "rework". Visual inspection is accepted with no dents or visual defects. |  | 9D8363ED-F6BF-1F89-42CB6E92851B7C0D |  | CHENG |  | 2011-11-02 |  | Use As Is
299 | 90 | HESSN |  | CRM1207020-1000 |  | HELIUM VESSEL ASSEMBLY (RH) | [45 empty] | C100-CM-INSP-HELV |  | R2 |  | 76 |  | OGLE |  | 2011-10-31 |  | This traveler is for Bellow serial number 4840-185 only. (replacement for 4840-179)Rejected per earlier inspection and dispositioned as return to vendor.Attached file is for dimensional non conforming after "rework". Visual inspection is accepted with no dents or visual defects. |  | 9DAFD659-EC2B-2B20-44C3C67F2D4DAABE |  | CHENG |  | 2011-11-02 |  | Use As Is
300 | 72 | HESSN |  | CRM1207020-1000 |  | HELIUM VESSEL ASSEMBLY (RH) | [45 empty] | C100-CM-INSP-HELV |  | R2 |  | 77 |  | OGLE |  | 2011-10-31 |  | This traveler is for Bellow serial number 4840-143 only. Rejected per earlier inspection and dispositioned as return to vendor.Attached file is for dimensional non conforming after "rework". Visual inspection is accepted with no dents or visual defects. |  | 9DD21754-F893-7730-82D4A05BA9D85D14 |  | CHENG |  | 2011-11-02 |  | Use As Is
301 | 87 | HESSN |  | CRM1207020-1000 |  | HELIUM VESSEL ASSEMBLY (RH) | [45 empty] | C100-CM-INSP-HELV |  | R2 |  | 78 |  | OGLE |  | 2011-10-31 |  | This traveler is for Bellow serial number 4840-173 only. Rejected per earlier inspection and dispositioned as return to vendor.Attached file is for dimensional non conforming after "rework". Visual inspection is accepted with no dents or visual defects. |  | 9DF33B3A-9D0E-A766-530C47D096E6BE55 |  | CHENG |  | 2011-11-02 |  | Use As Is
302 | 52 | HESSN |  | CRM1207020-1000 |  | HELIUM VESSEL ASSEMBLY (RH) | [45 empty] | C100-CM-INSP-HELV |  | R2 |  | 79 |  | OGLE |  | 2011-10-31 |  | This traveler is for Bellow serial number 4840-104 only. Rejected per earlier inspection and dispositioned as return to vendor.Attached file is for dimensional non conforming after "rework". .38 +/-.01 checks .364; Visual inspection is accepted with no dents or visual defects. |  | 9E13FCB3-084F-F68F-5AAC3FA8D2328075 |  | CHENG |  | 2011-11-02 |  | Use As Is
303 | 65 | HESSN |  | CRM1207020-1000 |  | HELIUM VESSEL ASSEMBLY (RH) | [45 empty] | C100-CM-INSP-HELV |  | R2 |  | 80 |  | OGLE |  | 2011-10-31 |  | This traveler is for Bellow serial number 4840-130 only. Rejected per earlier inspection and dispositioned as return to vendor.Attached file is for dimensional non conforming after "rework". Visual inspection is accepted with no dents or visual defects. |  | 9E2B4AE7-A3CA-00C4-8FBD6786EC6AFA74 |  | CHENG |  | 2011-11-02 |  | Use As Is
304 | 67 | HESSN |  | CRM1207020-1000 |  | HELIUM VESSEL ASSEMBLY (RH) | [45 empty] | C100-CM-INSP-HELV |  | R2 |  | 81 |  | OGLE |  | 2011-10-31 |  | This traveler is for Bellow serial number 4840-133 only. Rejected per earlier inspection and dispositioned as return to vendor.Attached file is for dimensional non conforming after "rework". Visual inspection is accepted with no dents or visual defects. |  | 9E41E624-A4B9-2054-6EACFD2975907188 |  | CHENG |  | 2011-11-02 |  | Use As Is
305 | 72 | HESSN |  | CRM1207020-1000 |  | HELIUM VESSEL ASSEMBLY (RH) | [45 empty] | C100-CM-INSP-HELV |  | R2 |  | 82 |  | OGLE |  | 2011-10-31 |  | This traveler is for Bellow serial number 4840-144 only. Rejected per earlier inspection and dispositioned as return to vendor.Attached file is for dimensional non conforming after "rework". Visual inspection is accepted with no dents or visual defects. |  | 9E585735-9B06-B377-8DF38E8BA6FD78DE |  | CHENG |  | 2011-11-02 |  | Use As Is
306 | L42 | GVWGSN |  | CRM1207075-0040 |  | WAVEGUIDE SPOOL - LEFT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 74 |  | OGLE |  | 2011-11-02 |  | Damaged knife edge. |  | 08AE9664-9B89-A8CE-03E0C03887852236 |  | KWILSON |  | 2011-11-03 |  | Modify |  | Please repair knife edge |  |  |  | WILCOX |  | 2011-11-30 |  | Repaired damaged knife edge before leak test. |  |  |  |  | OGLE |  | 2011-12-08 |  | Verified knife edge rework. JO 12/8/11 |  |  | KWILSON |  | 2012-01-12 |  | Use As Is
307 | L39 | GVWGSN |  | CRM1207075-0040 |  | WAVEGUIDE SPOOL - LEFT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 75 |  | OGLE |  | 2011-11-02 |  | See attachment for dimensional non conforming. |  | 08BA4EAB-B8E3-BB88-C656C4213F8851D2 |  | KWILSON |  | 2011-11-02 |  | Use As Is
308 | 5 | OMAGSN |  | CRM1207043-0000 |  | OUTER MAGNETIC SHIELD ASSEMBLY | [45 empty] | C100-CM-INSP-OMAG |  | R2 |  | 7 |  | GEORGED |  | 2011-11-10 |  | Missing item no. 26 (3" tape) |  |  |  | CHENG |  | 2011-11-10 |  | Use As Is
309 | CRM1207002-0001_004 | TPHTS_SerialID |  | 12070020001 |  | TOPHAT SINGLE | [45 empty] | C100-CM-INSP-TPHTS |  | R1 |  | 20 |  | OGLE |  | 2011-11-11 |  | See attachment for dimensional non conforming. Note: This unit is a reworked part returned to the supplier at an earlier date. Sample Waveguide will bolt up with the above condition. |  | D572664A-F5A9-D045-20072FB3FAA1C93D |  | KWILSON |  | 2011-11-16 |  | Use As Is
310 | CRM1207002-0001_016 | TPHTS_SerialID |  | 12070020001 |  | TOPHAT SINGLE | [45 empty] | C100-CM-INSP-TPHTS |  | R1 |  | 21 |  | OGLE |  | 2011-11-11 |  | Small nick on mounting surface in or very near seal area. See attachment for dimensional non conforming. Sample Waveguide will bolt up with the above condition. Note: This is a "reworked" unit returned to the supplier at an earlier date. |  | D634F473-0320-8BDB-091B966E8D4336BA |  | KWILSON |  | 2011-11-16 |  | Use As Is
311 | CRM1207002-0001_015 | TPHTS_SerialID |  | 12070020001 |  | TOPHAT SINGLE | [45 empty] | C100-CM-INSP-TPHTS |  | R1 |  | 22 |  | OGLE |  | 2011-11-11 |  | Note: This a "reworked unit returned to the supplier at an earlier date. Spot in o-ring groove. Tool marks on face in seal area. Surface finish exceeds 32 RA. Checks:30,31,33,38 See attachment for dimensional non conforming. Sample Waveguide will bolt up with the above condition. |  | D753AF92-DE0D-5F50-6D7AF451A3056441 |  | KWILSON |  | 2011-11-16 |  | Use As Is
312 | CRM1207002-0000_004 | TPHTD_SerialID |  | 12070020000 |  | TOPHAT DOUBLE | [45 empty] | C100-CM-INSP-TPHTD |  | R1 |  | 8 |  | OGLE |  | 2011-11-11 |  | Note: This is a "rework" unit returned to the supplier at an earlier date. Surface finish on seal plate exceeds 32 finish requirement. RA checks: 42,31,24,22 See attachment for dimensional non conforming. Sample Waveguide will bolt up with the above condition. |  | D82F3919-A1DE-5C1D-27D88B99F0A62FF0 |  | KWILSON |  | 2011-11-16 |  | Use As Is
313 | R43 | GVWGSN |  | CRM1207075-0040 |  | WAVEGUIDE SPOOL - LEFT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 76 |  | OGLE |  | 2011-11-29 |  | Tapped hole tight with sample bolt. Bolt will not go to full depth. |  |  |  | KWILSON |  | 2011-12-09 |  | Modify |  | Please retap as needed. |  |  |  | WILCOX |  | 2011-11-30 |  | Cleaned up threads in indicated hole. used sample bolt as test piece. Will go full depth now |  |  |  |  | OGLE |  | 2011-12-09 |  | Verified rework is complete. JO |  |  | KWILSON |  | 2012-01-12 |  | Use As Is
314 | L30 | GVWGSN |  | CRM1207075-0040 |  | WAVEGUIDE SPOOL - LEFT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 78 |  | OGLE |  | 2011-11-29 |  | See attachment for dimensional non conforming. Knife edge damage. Final inspection see step 11. |  | 79F6E78E-0818-8BFF-190F4B301590DEEA |  | KWILSON |  | 2011-12-09 |  | Modify |  | Dimensions ok. Please repair knife edge. |  |  |  | WILCOX |  | 2011-12-15 |  | repaired knife edge on large flange as per NCR |  |  |  |  | GEORGED |  | 2011-12-15 |  | Visual done, part okay. |  | A006CC96-0C8F-B535-1CA0430D3CED0525 | KWILSON |  | 2012-01-12 |  | Use As Is
315 | L26 | GVWGSN |  | CRM1207075-0040 |  | WAVEGUIDE SPOOL - LEFT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 80 |  | OGLE |  | 2011-11-30 |  | Small stains on interior copper surface. |  | A9365512-F55F-0CEB-A8263EE58DCED399 |  | KWILSON |  | 2011-12-01 |  | Use As Is
316 | L21 | GVWGSN |  | CRM1207075-0040 |  | WAVEGUIDE SPOOL - LEFT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 81 |  | OGLE |  | 2011-11-30 |  | Black spot on interior copper surface. Will not come off with alcohol. See attachment for dimensional non conforming. |  | AAB7E8B9-C334-FCDD-84C48428A0268CA7 |  | KWILSON |  | 2011-12-01 |  | Use As Is
317 | L40 | GVWGSN |  | CRM1207075-0040 |  | WAVEGUIDE SPOOL - LEFT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 84 |  | OGLE |  | 2011-12-01 |  | Machined mark on back surface into apex of knife edge. See attachment for dimensional non conforming. |  | D9CCA977-B92C-D618-EB93D349A5638744 |  | KWILSON |  | 2011-12-01 |  | Modify |  | Dimensions ok. Knife edge ok if part passes leak check. |  |  |  | WILCOX |  | 2011-12-09 |  | unit passed leak check |  |  |  |  | OGLE |  | 2011-12-16 |  | Knife edge not reworked; accepted if passes leak check,per NCR disposition. JO |  |  | KWILSON |  | 2011-12-20 |  | Use As Is
318 | L48 | GVWGSN |  | CRM1207075-0040 |  | WAVEGUIDE SPOOL - LEFT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 85 |  | OGLE |  | 2011-12-01 |  | See attachment for dimensional non conforming. |  | DB504C19-D6FA-EFF8-AF6063CF3ECA3DE6 |  | KWILSON |  | 2011-12-01 |  | Use As Is
319 | L48 | GVWGSN |  | CRM1207075-0040 |  | WAVEGUIDE SPOOL - LEFT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 85 |  | OGLE |  | 2011-12-01 |  | See attachment for dimensional non conforming. |  | DCD25089-E24B-3435-11CBFF448F53ABC5 |  | KWILSON |  | 2011-12-01 |  | Use As Is
320 | L34 | GVWGSN |  | CRM1207075-0040 |  | WAVEGUIDE SPOOL - LEFT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 86 |  | OGLE |  | 2011-12-01 |  | Black spot on interior copper surface. |  | DDD7F5C6-0C1B-946E-722DE25A9E502848 |  | KWILSON |  | 2011-12-12 |  | Use As Is
321 | L41 | GVWGSN |  | CRM1207075-0040 |  | WAVEGUIDE SPOOL - LEFT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 87 |  | OGLE |  | 2011-12-01 |  | 2 Threaded holes difficult to thread sample bolt to proper depth. See attachment for dimensional non conforming. |  | E0E984B1-C2B9-BEA3-C7DA6CA395884FD1 |  | KWILSON |  | 2011-12-09 |  | Modify |  | Please retap as needed. |  |  |  | WILCOX |  | 2011-12-15 |  | Ran a tap through marked holes and tested with sample bolt. All Ok |  |  |  |  | OGLE |  | 2012-01-02 |  | Verified re-tap of tight threads. OK |  |  | KWILSON |  | 2012-01-03 |  | Use As Is
322 | L38 | GVWGSN |  | CRM1207075-0040 |  | WAVEGUIDE SPOOL - LEFT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 88 |  | OGLE |  | 2011-12-01 |  | See attachment for dimensional non conforming. |  | E1E60463-9531-63CE-55FEB8774F5E5B9F |  | KWILSON |  | 2011-12-02 |  | Use As Is
323 | R46 | GVWGSN |  | CRM1207075-0040 |  | WAVEGUIDE SPOOL - LEFT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 90 |  | OGLE |  | 2011-12-05 |  | Two positions difficult to thread bolt to proper depth. |  |  |  | KWILSON |  | 2011-12-06 |  | Modify |  | Please retap as needed. |  |  |  | WILCOX |  | 2011-12-09 |  | repaired damaged threads on marked holes. tested with sample bolt. all ok now |  |  |  |  | OGLE |  | 2011-12-16 |  | Verified rework of tight threads. JO |  |  | KWILSON |  | 2011-12-20 |  | Use As Is
324 | R47 | GVWGSN |  | CRM1207075-0040 |  | WAVEGUIDE SPOOL - LEFT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 91 |  | OGLE |  | 2011-12-05 |  | Two threaded positions tight with sample bolt. Difficult to thread to proper depth. See attachment for dimensional non conforming. |  | A9B493F8-A673-6A33-E5E118C6CAEDAA41 |  | KWILSON |  | 2011-12-06 |  | Modify |  | Dimensions ok. Please retap as needed. |  |  |  | WILCOX |  | 2011-12-09 |  | repaired damaged threads on 2 marked tapped holes as per NCR. tested with sample bolt. All ok |  |  |  |  | OGLE |  | 2011-12-16 |  | Verified rework of tapped holes. JO |  |  | KWILSON |  | 2011-12-20 |  | Use As Is
325 | R48 | GVWGSN |  | CRM1207075-0040 |  | WAVEGUIDE SPOOL - LEFT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 92 |  | OGLE |  | 2011-12-05 |  | See attachment for dimensional non conforming. |  | ABAF8A27-E153-5655-6D36772EFE488086 |  | KWILSON |  | 2011-12-06 |  | Use As Is
326 | L33 | GVWGSN |  | CRM1207075-0040 |  | WAVEGUIDE SPOOL - LEFT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 94 |  | OGLE |  | 2011-12-06 |  | See attachment for dimensional non conforming. |  | D976D7C8-945B-82E5-A9C02F4ADA89EE88 |  | KWILSON |  | 2011-12-07 |  | Use As Is
327 | L32 | GVWGSN |  | CRM1207075-0040 |  | WAVEGUIDE SPOOL - LEFT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 96 |  | OGLE |  | 2011-12-06 |  | See attachment for dimensional non conforming. |  | E04D11BA-E754-96A7-123300F993AE65ED |  | KWILSON |  | 2011-12-07 |  | Use As Is
328 | L44 | GVWGSN |  | CRM1207075-0040 |  | WAVEGUIDE SPOOL - LEFT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 97 |  | OGLE |  | 2011-12-06 |  | Small conflat knife edge nicked. Two threaded positions tight with sample bolt; difficult to thread to proper depth. See attachment for dimensional non conforming. |  | E27F738F-ABC0-338C-658E763297B42F00 |  | KWILSON |  | 2011-12-07 |  | Modify |  | Dimensions ok. Please repair bolt hole and retap knife edge. |  |  |  | WILCOX |  | 2011-12-09 |  | repaired damaged threads on marked tapped holes. polished knife edge on small conflat before leak test. unit passed |  |  |  |  | OGLE |  | 2011-12-09 |  | Verified rework. |  |  | KWILSON |  | 2011-12-20 |  | Use As Is
329 | R48 | GVWGSN |  | CRM1207075-0040 |  | WAVEGUIDE SPOOL - LEFT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 92 |  | OGLE |  | 2011-12-12 |  | Minor stain spots on interior copper surface. |  | 14662E88-B3BC-7465-4460CA6EA44401C2 |  | KWILSON |  | 2012-01-12 |  | Use As Is
330 | 31 | HESSN |  | CRM1207020-1000 |  | HELIUM VESSEL ASSEMBLY (RH) | [45 empty] | C100-CM-INSP-HELV |  | R2 |  | 83 |  | OGLE |  | 2011-12-15 |  | This traveler is for Bellow serial no. 4840-198 only. It a replacement for Bellow serial number 4840-063. See the attachment for dimensional non conforming. |  | A0F53366-EA91-CC13-0977B73836FF5EE6 |  | CHENG |  | 2011-12-15 |  | Use As Is
331 | 8 | VVSN |  | CRM1207090-1000 |  | VACUUM VESSEL | [45 empty] | C100-CM-INSP-VV |  | R3 |  | 11 |  | GEORGED |  | 2012-01-12 |  | One NW320 Half nipple flange has damage, raised edge present on the face of the flange. There is a small ding inside the seal path on one tophat flange. |  | 3A500495-DB47-6111-8EA1388534D75F11 |  | CHENG |  | 2012-01-12 |  | Use As Is |  | The damage on NW320 flange may not affect the use of this port. If it does, then a fix is needed. As to the ding on tophat port, if the ding is found to be right on the seal path during assembly, then this port needs to be ground to specified surface finish.
332 | 7 | VVSN |  | CRM1207090-1000 |  | VACUUM VESSEL | [45 empty] | C100-CM-INSP-VV |  | R3 |  | 10 |  | GEORGED |  | 2012-01-13 |  | Please see attached pictures and PDF. |  | 72AAD802-C159-AA98-5E0DD8D6CA520C19 |  | CHENG |  | 2012-01-13 |  | Use As Is |  | Per communication with John Fischer, this port can be repaired during cryomodule assembly if necessary.
333 | 6 | OMAGSN |  | CRM1207043-0000 |  | OUTER MAGNETIC SHIELD ASSEMBLY | [45 empty] | C100-CM-INSP-OMAG |  | R2 |  | 8 |  | GEORGED |  | 2012-01-18 |  | Missing Item 26 (3" wide tape). |  |  |  | CHENG |  | 2012-01-18 |  | Use As Is
334 | 10 | VVSN |  | CRM1207090-1000 |  | VACUUM VESSEL | [45 empty] | C100-CM-INSP-VV |  | R3 |  | 13 |  | GEORGED |  | 2012-02-06 |  | Vacuum Vessel is missing 2 Half Nipple flanges (Item #11). Pits in seal path on Tuner Mount flanges. Fingerprints found on one loose Tuner Mount flange. |  | 4E82AD01-FC67-9177-8804BAEC77C7CF31 |  | CHENG |  | 2012-02-10 |  | Return To Vendor |  | Please return this vacuum vessel to vendor for repair.
335 | L33 | GVWGSN |  | CRM1207075-0040 |  | WAVEGUIDE SPOOL - LEFT | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 94 |  | OGLE |  | 2012-03-09 |  | Small nick on knife edge. No photo...camera packed for move. Repairable based on previous units. JO. |  |  |  | KWILSON |  | 2012-03-09 |  | Modify | [5 empty] | WILCOX |  | 2012-03-12 |  | Polished knife edge as per NCR disposition |  |  |  |  | OGLE |  | 2012-03-13 |  | Repair verified and accepted. JO 3/13/12 |  |  | KWILSON |  | 2012-03-13 |  | Use As Is
336 | 67 | ELFTSN |  | 13085-01-CF |  | 8 PIN DOUBLE ENDED FEED THRU MINI CONFLAT | [45 empty] | C100-CM-INSP-ELFT |  | R1 |  | 68 |  | FISCHER |  | 2012-03-27 |  | after cold shock feedthru leaked. Will be sent back to vendor for replacement. |  |  |  | FISCHER |  | 2012-03-28 |  | Return To Vendor
337 | 64 | ELFTSN |  | 13085-01-CF |  | 8 PIN DOUBLE ENDED FEED THRU MINI CONFLAT | [45 empty] | C100-CM-INSP-ELFT |  | R1 |  | 65 |  | FISCHER |  | 2012-03-27 |  | feedthru had leak after VTA cold shock. Will be sent back to vendor for replacement. |  |  |  | FISCHER |  | 2012-03-28 |  | Return To Vendor
338 | 62 | ELFTSN |  | 13085-01-CF |  | 8 PIN DOUBLE ENDED FEED THRU MINI CONFLAT | [45 empty] | C100-CM-INSP-ELFT |  | R1 |  | 63 |  | FISCHER |  | 2012-03-27 |  | Feedthru developed leak after VTA coldshock. Unit will be shipped back to vendor for replacement. |  |  |  | FISCHER |  | 2012-03-28 |  | Return To Vendor
339 | L44 |  |  | CRM1207075-0040 |  | Secondary Waveguide | [45 empty] | C100-CM-INSP-GVWG |  | R1 |  | 97 |  | FISCHER |  | 2012-05-17 |  | Secondary waveguide spool has scratch running across rectangular flange knife edge,noticed after pre assembly steps were taken. returned to inventory |  |  |  | KWILSON |  | 2012-05-17 |  | Modify |  | Please repair new scratch on knife edge. | [5 empty] | 2012-05-17 | [8 empty] | 2012-05-17 | [16 empty] | Part returned to inventory area, should be repaired by qualified technician.Upon successful repair, put back in the que for assembly.

## C100R-NCR
Status Report for C100R-NCR on 10-Jun-20
TRAV_SEQ_NUM | SN | SNTYPE | REWORKDATE1 | DRAWINGNUMBER | REWORKTECH1 | PARTDESCRIPTION | REWORKCOMMENT1 | SHIPPINGDAMAGE | REWORKFILE1 | DIMENSIONAL | REWORKDATE2 | REWORKTECH2 | LEAKTESTFAILURE | REWORKCOMMENT2 | ROUGHNESS | REWORKFILE2 | FLATNESS | PARALLELISM | REWORKDATE3 | REWORKTECH3 | SCRATCHES | REWORKCOMMENT3 | DINGS | REWORKFILE3 | DELAMINATION | PLATING | REWORKDATE4 | STAINS | REWORKTECH4 | REWORKCOMMENT4 | OXIDATION | RESIDUES | REWORKFILE4 | INSPECTIONFAILURE | REWORKDATE5 | RFTESTFAILURE | REWORKTECH5 | REWORKCOMMENT5 | VACUUM | RFCMTF | REWORKFILE5 | INSTRUMENTATION | VERIFYDATE1 | MECHANICAL | VERIFYTECH1 | CMTFRFPERFORMANCE | VERIFYCOMMENT1 | VERIFYFILE1 | VTARFPERFORMANCE | OTHER | VERIFYDATE2 | TRAVELERID | VERIFYTECH2 | TRAVELERREVISION | VERIFYCOMMENT2 | TRAVSEQNUMBER | VERIFYFILE2 | DESCRIBESTAFF | VERIFYDATE3 | DESCRIBEDATE | VERIFYTECH3 | DESCRIBECOMMENT | VERIFYCOMMENT3 | DESCRIBEDOCS | VERIFYFILE3 | RESPONDINGENGINEER | VERIFYDATE4 | RESPONDINGDATE | VERIFYTECH4 | DISPOSITION | VERIFYCOMMENT4 | DISPOSITIONCOMMENT | VERIFYFILE4 | DISPOSITIONDOCS | VERIFYDATE5 | REWORKSTAFF | VERIFYTECH5 | REWORKDATE | VERIFYCOMMENT5 | REWORKCOMMENT | VERIFYFILE5 | RESPONDINGDATE1 | REWORKDOCS | RESPONDINGENGR1 | VERIFYSTAFF | RESPONDINGCOMMENT1 | VERIFYDATE | RESPONDINGFILE1 | VERIFYCOMMENT | RESPONDINGDATE2 | VERIFYDOCS | FINALENGINEER | RESPONDINGENGR2 | FINALDATE | RESPONDINGCOMMENT2 | FINALDISPOSITION | RESPONDINGFILE2 | FINALCOMMENT | RESPONDINGDATE3 | FINALDOCS | RESPONDINGENGR3 | CONNECTIVITYCOMMENT | RESPONDINGCOMMENT3 | NOTIFICATIONCOMMENT | RESPONDINGFILE3 | RESPONDINGDATE4 | RESPONDINGENGR4 | RESPONDINGCOMMENT4 | RESPONDINGFILE4 | RESPONDINGDATE5 | RESPONDINGENGR5 | RESPONDINGCOMMENT5 | RESPONDINGFILE5
1 | C100-RI-039 | CAVSN |  |  |  | C100R O39 HV Assy |  |  |  | 1 | [41 empty] | C100R-CAV-ASSY-HELV |  | R1 |  | 2 |  | JOGLE |  | 2018-12-07 |  | Dimensional non-conforming see attached CMM Report. |  | 6EF39D39-FABB-9A81-BC439835195B0AF7 |  | FISCHER |  | 2019-10-31 |  | Use As Is |  | Unit was used in the rebuild of C100-06R and worked. Position 5
2 | C100-RI-055 | CAVSN | 2019-01-16 | C100R -RI-055 | GEORGED | C100R O55 HV Assy | Bending performed on Fpc flange. True position okay after adjustment. |  | 687E2D6A-F819-E56A-D6853FBDB02D535B | 1 | 2019-01-16 | OVERTONR |  | remeasured good |  | 68B68C09-A974-8F55-CB0560F3E796361D | [26 empty] | 2019-01-16 |  | GEORGED |  | Please see attached cmm report. |  |  |  |  | C100R-CAV-ASSY-HELV |  | R1 |  | 3 |  | JOGLE |  | 2018-12-07 |  | See attachment for dimensional non-conforming. |  | 73767762-B7B0-F8EA-62AC9236A0C8E12D |  | ARI |  | 2018-12-14 |  | Hold/Modify |  | return to tuning, we will try and tune in the HV | [9 empty] | 2019-01-16 |  | ARI |  | good |  |  |  | 2019-01-16 |  | ARI | ARI | 2019-01-16 | good | Use As Is |  | both field flattens and dimentional within spec
3 | C100-RI-053 | CAVSN |  | C100R -053 |  | C100 053 HV ASSY |  |  |  | 1 | [41 empty] | C100R-CAV-ASSY-HELV |  | R1 |  | 4 |  | JOGLE |  | 2018-12-11 |  | Dimensional non-conforming.See attachment for dimensional inspection results. |  | 45570524-CFBB-57DE-682F9157BE13C93E |  | ARI |  | 2018-12-14 |  | Use As Is |  | unchanged from 2010 when originally recieved
4 | C100-RI-051 | CAVSN | [7 empty] | 1 | [41 empty] | C100R-CAV-ASSY-HELV |  | R1 |  | 5 |  | GEORGED |  | 2018-12-12 |  | Dimension out of tolerance. Please see attached. |  | 69D79511-0CAA-E306-75BB4D1D1F027FB2 |  | KDAVIS |  | 2018-12-12 |  | Use As Is
5 | C100-RI-049 | CAVSN | 2019-01-14 | C100R-RI-049 | GEORGED | C100R O49 HV Assy | Bending performed on Fpc flange to improve true position of beam tube and perpendicularity. |  |  | 1 | 2019-01-16 | OVERTONR |  | remeasured good |  | 70C05FA4-C6CE-FF53-0FF72ED1A0FD4F7C | [26 empty] | 2019-01-14 |  | GEORGED |  | RI-049 remeasured on CMM, dimension no longer out of tolerance. True position okay. Please see attached report. | 0B8E32CE-EADF-33A5-C7C66ECE69B4C6ED |  |  |  | C100R-CAV-ASSY-HELV |  | R1 |  | 6 |  | JOGLE |  | 2018-12-14 |  | Minor dimensional non-conforming. See Attachment for details. |  | DCA8EEF2-06F7-6EFB-DA857D4BB272953B |  | ARI |  | 2019-01-14 |  | Hold/Modify |  | bend back into tolerance and then do RF check | [9 empty] | 2019-01-16 |  | ARI |  | good |  |  |  | 2019-01-16 |  | ARI | ARI | 2019-01-16 | good | Use As Is
6 | C100-RI-043 | CAVSN | 2019-01-18 | C100R -RI-043 | GEORGED | C100R O43 HV Assy | Bending performed on Fpc flange to improve true position and perpendicularity of beam tube. |  |  | 1 | [32 empty] | 2019-01-18 |  | GEORGED |  | Passed inspection. CMM report attached. | D7D64EFC-FA91-56E5-510B5F537EE9E156 |  |  |  | C100R-CAV-ASSY-HELV |  | R1 |  | 7 |  | JOGLE |  | 2018-12-17 |  | Dimensional non conforming. See attachment for detailed dimensional report. |  | 77FD1D2F-9884-06A4-8ED309EFFFA66F97 |  | KDAVIS |  | 2019-01-02 |  | Proceed to next WorkCenter |  | field flatness as measured allows for deformation during string assembly. | [19 empty] | FISCHER |  | 2019-10-31 |  | Use As Is |  | Cavity was used in the C100-06R Cryomodule and worked with no noted issues.
7 | 51 | cav | 2018-12-18 |  | ASHLEYA |  | HPR | 1 | [34 empty] | 2018-12-18 |  | BOOKWALT |  | fine |  |  |  |  | C100R-CAV-INSP |  | R1 |  | 2 |  | BOOKWALT |  | 2018-12-18 | [5 empty] | BOOKWALT |  | 2018-12-18 |  | Hold/Re-measure |  | please remeasure | [9 empty] | 2018-12-18 |  | BOOKWALT | [7 empty] | BOOKWALT |  | 2018-12-18 |  | Use As Is
8 | C100-RI-047 | CAVSN |  | C100R -RI-047 |  | C100R O47 HV Assy |  |  |  | 1 | [41 empty] | C100R-CAV-ASSY-HELV |  | R1 |  | 8 |  | JOGLE |  | 2018-12-18 |  | Dimensional non conforming.See attachment for dimensional inspection report. |  | A832543F-C151-7CBF-394FE125C51590B9 |  | KDAVIS |  | 2018-12-19 |  | Use As Is
9 | C100-RI-039 | CAVSN | [49 empty] | C100R-CAV-VTRF |  | R1 |  | 3 |  | LZHAO |  | 2018-12-20 |  | Lock frequency was 1496.7079 MHz, slightly out of spec which is 1496.400-1496.700 MHz |  |  |  | ARI |  | 2019-01-16 |  | Use As Is |  | cavity can still be tuned in the string
10 | C100-RI-056 | CAV | 2019-01-06 |  | DREYFUSS |  | return to chemistry in cage | [28 empty] | 1 | [15 empty] | C100R-CAV-VTRF |  | R1 |  | 4 |  | POWEN |  | 2018-12-21 |  | FE Onset at 11 MV/m |  |  |  | ARI |  | 2019-01-16 |  | Hold/Re-measure |  | return to chemistry to re-processed from degreasing forward | [9 empty] | 2019-01-16 |  | ARI | [7 empty] | ARI |  | 2019-01-16 |  | Use As Is
11 | C100-RI-051 | C100R | [43 empty] | 1 | [5 empty] | C100R-CAV-VTRF |  | R1 |  | 6 |  | POWERS |  | 2019-01-09 |  | Field probe Q is slightly low at 7.55e11 range 0.8 to 1.8e12. |  |  |  | ARI |  | 2019-01-16 |  | Use As Is |  | just out of spec, can be used without modification
12 | C100-RI-047 | CAVSN | 2019-01-16 |  | OVERTONR |  | re-measure complete |  | 68D15919-0D55-0E7E-97B12E11B8FA5E2A | [24 empty] | 1 | [15 empty] | 1 |  | C100R-CAV-RFIN |  | R1 |  | 8 |  | OVERTONR |  | 2019-01-15 |  | Cavity Field flatness at 87% FPC Body Qext 3.09E7 |  | 350A5601-B757-FE4E-1FF8B896FF001375 |  | ARI |  | 2019-01-16 |  | Hold/Re-measure |  | remeasure verifying background is flat | [9 empty] | 2019-01-16 |  | ARI |  | good | [5 empty] | ARI |  | 2019-01-16 |  | Use As Is
13 | C100-RI-047 | CAVSN | [49 empty] | C100R-CAV-VTRF |  | R1 |  | 11 |  | POWEN |  | 2019-01-28 |  | FE onset 8 MV/m |  |  |  | KDAVIS |  | 2019-10-29 |  | Hold/Modify |  | Rinse/Retest VTRF | [19 empty] | KDAVIS |  | 2019-10-29 |  | Reject |  | requires rinse/retest
14 | L16 | GVWGSN | 2019-02-07 |  | LPAGE |  | Polished with stone. | [15 empty] | 1 | [19 empty] | 2019-02-07 |  | GEORGED |  | Visual okay. | D6BC83B9-F68B-CD2B-513299FD968B912A |  |  |  | C100R-CM-INSP-GVWG |  | R2 |  | 2 |  | GEORGED |  | 2019-01-31 |  | Knife edge damaged on 2.75" conflat flange. Please see attached picture of damage. |  | 702436F5-9CA4-3518-DB2F503EAC23164D |  | FISCHER |  | 2019-02-04 |  | Hold/Modify |  | We will polish the flange and leak check to qualify. Unit will then be returned to the CMM Inspection area for re-evaluation. | [19 empty] | FISCHER |  | 2019-10-31 |  | Use As Is |  | Waveguide was installed and leak checked good on the C100-06R Cryomodule |  | 62C14DB0-9AE0-0B99-65C4AA84C4DEC5F7
15 | C100R | CAV | [33 empty] | 1 | [15 empty] | C100R-CAV-VTRF |  | R1 |  | 15 |  | POWEN |  | 2019-02-08 |  | FE onset at 8 MV/m after activation at 17 MV/m |  |  |  | KDAVIS |  | 2019-10-29 |  | Reject |  | Requires rinse/retest

## F100-CAV-NCR
Status Report for F100-CAV-NCR on 10-Jun-20
TRAV_SEQ_NUM | SN | SNTYPE | DRAWINGNUMBER | PARTDESCRIPTION | CONNECTIVITYCOMMENT | TRAVELERID | TRAVELERREVISION | TRAVSEQNUMBER | DESCRIBECOMMENT | DESCRIBESTAFF | DESCRIBEDATE | DESCRIBEDOCS | NOTIFICATIONCOMMENT | RESPONDINGENGINEER | RESPONDINGDATE | DISPOSITIONCOMMENT | DISPOSITION | DISPOSITIONDOCS | REWORKCOMMENT | REWORKSTAFF | REWORKDATE | REWORKDOCS | VERIFYCOMMENT | VERIFYSTAFF | VERIFYDATE | VERIFYDOCS | FINALENGINEER | FINALDATE | FINALCOMMENT | FINALDISPOSITION | FINALDOCS
1 | 48643-0128 | WBASN | CRM1207075-0024 | BRAZEMENT ADAPTER ASSY |  | F100-CAV-INSP-ADPT | R1 | 2 | See report. | CARPENTE | 2012-08-29 | 7B858FC5-DFAD-5033-8A8E0805000E6D44 |  | HKPARK | 2012-09-04 | The level of deviation is not affecting the performance. | Use As Is | B1E3837F-C1D6-7066-CE52545DC96CE789
2 | test | test | test | test | testing file upload | F100-CAV-NCR | R7 | 2 | testing | BOOKWALT | 2012-09-04 | B2057188-D669-3774-82A4978D5F5F8CB7 |  | BOOKWALT | 2013-09-10 | test | Modify |  | test | BOOKWALT | 2013-09-10 |  | test | BOOKWALT | 2013-09-10
3 | 48643-0129 |  | unkn | unkn |  | F100-CAV-INSP-ADPT | R1 | 3 | Dimensional, please see report. | CARPENTE | 2012-09-04 | B829FC17-CBF0-440F-68999A3D85E25247 |  | HKPARK | 2012-09-04 | Fit into the flange pocket and if there is no interference use as is. | Use As Is
4 | 48643-0130 |  | unkn | unkn |  | F100-CAV-INSP-ADPT | R1 | 4 | Braze issue on eyelet, and dimensional, please see report. | CARPENTE | 2012-09-04 | B826702D-D0B3-47F8-73643F5D3B42EFAC |  | HKPARK | 2012-09-11 | Checked the intereference of the raised braze spot. No interference. Use as is. | Use As Is
5 | 48643-0131 |  | unkn | unkn |  | F100-CAV-INSP-ADPT | R1 | 5 | Dimensional, please see report. | CARPENTE | 2012-09-04 | B823072E-B9C6-CFBA-EDC69AB669E60E89 |  | HKPARK | 2012-09-04 |  | Use As Is
6 | 48643-0132 |  | unkn | unkn |  | F100-CAV-INSP-ADPT | R1 | 6 | Dimensional, please see report. | CARPENTE | 2012-09-04 | B81D9780-0AB0-9B6E-D5BC11E609C23567 |  | HKPARK | 2012-09-04 |  | Use As Is
7 | 48643-0133 |  | unkn | unkn |  | F100-CAV-INSP-ADPT | R1 | 7 | Dimensional, please see report. | CARPENTE | 2012-09-04 | B818ADD7-B9D5-63FC-09EDF1E86F1AFCC5 |  | HKPARK | 2012-09-04 |  | Use As Is
8 | 48643-0134 |  | unkn | unkn |  | F100-CAV-INSP-ADPT | R1 | 8 | Dimensional, please see report. | CARPENTE | 2012-09-04 | B85E29A7-CAED-1613-12153EFFDDB33EB6 |  | HKPARK | 2012-09-04 |  | Use As Is
9 | 48643-0136 |  | unkn | unkn |  | F100-CAV-INSP-ADPT | R1 | 10 | Dimensional, please see report. | CARPENTE | 2012-09-04 | B8A52ED6-D290-FB2D-35AA7187182FF15C |  | HKPARK | 2012-09-04 |  | Use As Is
10 | 48643-0137 | WBASN | unkn | unkn |  | F100-CAV-INSP-ADPT | R1 | 11 | Black spot on ceramic, and dimensional, please see report. | CARPENTE | 2012-09-04 | BA2E53A1-980D-3834-D7B49F0701580029 |  | HKPARK | 2012-09-11 | Very slight spot. Please proceed to the next step. Rf test will determine whether it's acceptable or not. | Proceed
11 | 48643-0139 |  | unkn | unkn |  | F100-CAV-INSP-ADPT | R1 | 13 | Dimensional, please see report. | CARPENTE | 2012-09-04 | BA573D6A-9227-0361-C6997917CF077533 |  | HKPARK | 2012-09-04 |  | Use As Is
12 | 48643-0140 |  | unkn | unkn |  | F100-CAV-INSP-ADPT | R1 | 14 | Dimensional, please see report. | CARPENTE | 2012-09-04 | BA6C6ADE-BD3D-E3A5-4A9526C47FFEE4DE |  | HKPARK | 2012-09-04 |  | Use As Is
13 | 48643-0141 |  | unkn | unkn |  | F100-CAV-INSP-ADPT | R1 | 15 | Dimensional, please see report. | CARPENTE | 2012-09-04 | BAF616B8-C715-FC5A-55B3A7BE6D2E930E |  | HKPARK | 2012-09-04 | Fit into the flange pocket if it fits use as is. | Use As Is
14 | 48643-0142 |  | unkn | unkn |  | F100-CAV-INSP-ADPT | R1 | 16 | Dimensional, please see report. | CARPENTE | 2012-09-04 | BB0AC066-E18F-13F9-579AB0E8C368DA2C |  | HKPARK | 2012-09-04 |  | Use As Is
15 | 48643-0143 |  | unkn | unkn |  | F100-CAV-INSP-ADPT | R1 | 17 | Dimensional, please see report. | CARPENTE | 2012-09-04 | BB2774E5-B6A5-E9EF-8AC142B6A73E8428 |  | HKPARK | 2012-09-04 |  | Use As Is
16 | 48643-0144 | WBASN | unkn | unkn |  | F100-CAV-INSP-ADPT | R1 | 18 | Dimensional, please see report. | CARPENTE | 2012-09-04 | BB467D1E-F44D-7D26-8EB988F4C0E8745E
17 | 48643-0145 | WBASN | unkn | unkn |  | F100-CAV-INSP-ADPT | R1 | 19 | Black spot on ceramic, and dimensional, please see report. | CARPENTE | 2012-09-04 | BB64661A-A21A-08F5-A293870FE4083FAF |  | HKPARK | 2012-09-11 | Thickness is OK. Rf test is need to determine the spot is acceptable or not. Proceed to the next step. | Proceed
18 | 48643-0146 |  | unkn | unkn |  | F100-CAV-INSP-ADPT | R1 | 20 | Dimensional, please see report. | CARPENTE | 2012-09-05 | DD942CEE-FD3E-D3DB-5E5C3C87974EFBF2 |  | HKPARK | 2012-09-11 |  | Use As Is
19 | 48643-0147 |  | unkn | unkn |  | F100-CAV-INSP-ADPT | R1 | 21 | Dimensional, please see report. | CARPENTE | 2012-09-05 | DDAEE918-9643-94F7-FF46516C1D34C043 |  | HKPARK | 2012-09-10 |  | Use As Is
20 | 48643-0148 |  | unkn | unkn |  | F100-CAV-INSP-ADPT | R1 | 22 | Dimensional, please see report. | CARPENTE | 2012-09-05 | DDD0EAE6-DFD2-4C93-911A521500D7300D |  | HKPARK | 2012-09-10 |  | Use As Is
21 | 48643-0149 |  | unkn | unkn |  | F100-CAV-INSP-ADPT | R1 | 23 | Dimensional, please see report. | CARPENTE | 2012-09-05 | DDE61D76-BB87-8A77-0C1575FA8B025705 |  | HKPARK | 2012-09-10 |  | Use As Is
22 | 48643-0150 |  | unkn | unkn |  | F100-CAV-INSP-ADPT | R1 | 24 | Dimensional, please see report. | CARPENTE | 2012-09-05 | DE0CC398-D7AB-678F-16380358A6604E61 |  | HKPARK | 2012-09-10 |  | Use As Is
23 | 48643-0151 |  | unkn | unkn |  | F100-CAV-INSP-ADPT | R1 | 25 | Nick on eyelet flange, and dimensional, please see report. | CARPENTE | 2012-09-05 | DE2638A7-B1E3-8627-4CB1449BA2A90E4B |  | HKPARK | 2012-09-11 | Thickness is OK. The nick is benign. It does not create the leak path. use as is. | Use As Is
24 | F100HOM16 |  | unkn | unkn |  | F100-CAV-INSP-HMFT | R1 | 17 | Antenna is visually bent | CARPENTE | 2012-12-05 | 3D991A9E-C7D8-3BC6-61DBBF1B68DE5DEE |  | CARPENTE | 2012-12-05 |  | Reject
25 | FLG0037-251 |  | unkn | unkn |  | F100-CAV-INSP-FLNG | R1 | 2 | Roughness in one spot of knife edge. Should be repairable. | CARPENTE | 2012-12-07 |  |  | HKPARK | 2012-12-11 | Polish the rough spot and re-inspect. | Modify |  | REWORKED KNIFE EDGE AS PER NCR DISPOSITION | WILCOX | 2013-01-04 |  | Visual done, knife edge okay. | GEORGED | 2013-01-04 |  | HKPARK | 2013-01-07 |  | Use As Is
26 | TEST | TEST | TEST | TEST | TEST | F100-CAV-NCR | R1 | 1 | TEST | BOOKWALT | 2013-09-10 |  | TEST | BOOKWALT | 2013-09-10 | TEST ONLY | Use As Is | [9 empty] | BOOKWALT | 2013-09-10 | TEST | Use As Is
27 | FLG0037-259 |  | unkn | unkn |  | F100-CAV-INSP-FLNG | R1 | 10 | (1) bolt hole location out of tolerance 0.007" | CARPENTE | 2012-12-07 |  |  | HKPARK | 2012-12-11 | Please check with the mating flange if it can be assembled. | Use As Is
29 | FLG0037-252 | DSRFSN | MDC 203046-2000 | DOUBLE SIDED WINDOW FLANGE |  | F100-CAV-INSP-FLNG | R1 | 3 | Rough area on knife edge. | CARPENTE | 2012-12-10 |  |  | HKPARK | 2012-12-11 | Polish the rough area and re-inspect. | Modify |  | REWORKED KNIFE EDGE AS PER NCR DISPOSITION. | WILCOX | 2013-01-04 |  | Visual done, knife edge okay. | GEORGED | 2013-01-07 |  | HKPARK | 2013-01-07 |  | Use As Is
30 | FLG0037-260 |  | unkn | unkn |  | F100-CAV-INSP-FLNG | R1 | 11 | Nick on knife edge and corner of knife edge pocket. | CARPENTE | 2012-12-10 |  |  | HKPARK | 2012-12-11 | Repair. If it does not work return to vendor. | Modify |  | REWORKED KNIFE EDGE AS PER NCR DISPOSITION. GOT SECOND OPINION FROM MIKE MCCREA AFTER WORK WAS COMPLETED. | WILCOX | 2013-01-04 |  | Visual done, knife edge okay. | GEORGED | 2013-01-07 |  | HKPARK | 2013-01-07 |  | Use As Is
31 | FLG0037-266 |  | unkn | unkn |  | F100-CAV-INSP-FLNG | R1 | 17 | Knife edge has a rough spot. | CARPENTE | 2012-12-10 |  |  | HKPARK | 2012-12-11 | If it can be done polish it and re-inspect. | Modify |  | REWORKED KNIFE EDGE AS PER NCR DISPOSITION. | WILCOX | 2013-01-04 |  | Visual done, knife edge okay. | GEORGED | 2013-01-07 |  | HKPARK | 2013-01-07 |  | Use As Is
32 | FLG0037-269 |  | unkn | unkn |  | F100-CAV-INSP-FLNG | R1 | 20 | Knife edge has a rough spot. | CARPENTE | 2012-12-10 |  |  | HKPARK | 2012-12-11 | it it's repairable polish the rough spot and re-inspect. | Modify |  | REWORKED KNIFE EDGE AS PER NCR DISPOSITION | WILCOX | 2013-01-04 |  | Visual done, knife edge okay. | GEORGED | 2013-01-07 |  | HKPARK | 2013-01-07 |  | Use As Is
33 | F034887_23 |  | unkn | unkn |  | F100-CAV-INSP-WGD | R1 | 2 | Nicks on knife edge | CARPENTE | 2013-06-25 |  |  | KWILSON | 2013-06-25 | Please repair knife edge. | Use As Is |  | Knife edge reworked. | WILCOX | 2013-07-11 |  | Visual good. | GEORGED | 2013-07-11
34 | F034887_07 |  | unkn | unkn |  | F100-CAV-INSP-WGD | R1 | 3 | Nick on knife edge | CARPENTE | 2013-06-25 |  |  | KWILSON | 2013-06-25 | Please repair knife edge. | Use As Is |  | Knife edge reworked. | WILCOX | 2013-07-11 |  | Visual good. Some copper was removed during rework process. | GEORGED | 2013-07-11 | 16B12550-BD7D-727A-6AC00C085913E875
35 | F028589_04 |  | unkn | unkn |  | F100-CAV-INSP-WGD | R1 | 4 | Nick on knife edge. | CARPENTE | 2013-06-25 |  |  | KWILSON | 2013-06-25 | Please repair knife edge. | Use As Is |  | Knife edge reworked. | WILCOX | 2013-07-11 |  | Visual good. | GEORGED | 2013-07-11
37 | JL015 | CAVSN | 115400-1004 | CRYOCAVITY WELDMENT ASSEMBLY |  | F100-CAV-LAP | R2 | 5 | Both HOM flanges need seal surface repair due to pits and a rough finish. | CARPENTE | 2013-07-02 | 74534757-C845-E4D8-DE772DC42BC0DA62 |  | CASTAGNO | 2013-07-03 |  | Re-process |  | Both HOM flange sealing surfaces need to be reworked. The JLAB machine shop has fixturing from C100 cavities that can be used to correct the pitted flanges. The cavity should be taken to the machine shop for repair. SC | CASTAGNO | 2013-07-03 |  | HOM and field probe flanges have been inspected on the CMM, deviations have been found and need to be addressed. Visual acceptable. RF measurements were performed on this cavity and no issues were found. Probe tips will be reset on all feed-throughs. Cavity can be used "as is". Seal surface dimensional variations have been addressed as per attached document. | CARPENTE | 2013-07-23 | 7470F005-C37C-27AA-C4F831DD32458529 | CARPENTE | 2013-08-09 |  | Use As Is
38 | JL017 | CAVSN | unkn | unkn |  | F100-CAV-VTA-HPRF | R1 | 2 | Lock frequency for F100_JL-017 1496.3679 MHz - outside the specification range | STIRBET | 2013-07-11 |  | It is about 230kHz lower than target frequency due to the EP. This requires tuner to pre-stretch the cavity at warm after string assembly by
40 | G103046-04 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | F100-CAV-INSP-WGD | R1 | 5 | Scratches are in seal path of the small flange. | GEORGED | 2013-07-11 | 15FFE93E-DFDB-AE44-A813F1C6360F35F0 |  | KWILSON | 2013-07-16 | Please lightly polish to remove scratches. | Use As Is |  | polished small flange | CASTAGNO | 2013-08-28 |  | Visual,finish and flatness good. | CARPENTE | 2013-08-28
41 | G103046-01 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | F100-CAV-INSP-WGD | R1 | 9 | Scratches present in the seal path area of small flange. | GEORGED | 2013-07-11 | 1F0BB6B5-0D94-71DF-5E33812B0B3146B4 | [6 empty] | 1. Used Dremel tool with 40 micron sand paper. 2. Used Scotch Brite to polish flange. | DANIELS | 2013-08-28 |  | Visual, flatness, and finish ok, | CARPENTE | 2013-08-28
42 | G103046-03 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | F100-CAV-INSP-WGD | R1 | 10 | Scratches and stains present on the small flange. Copper plating deteriorated on large flange. | GEORGED | 2013-07-12 | 4A79A398-F4AA-55DA-D5FA208B61853B62 | [6 empty] | Dremel and hand lap | CASTAGNO | 2013-09-18 |  | Visual good. Flatness = .001" Finish = 15 uin. | CARPENTE | 2013-09-18
43 | G096314-03 |  | unkn | unkn |  | F100-CAV-INSP-WGD | R1 | 12 | Large flange has one bad threaded hole, the bolt doesn't go completely down to proper depth. | GEORGED | 2013-07-16 |  |  | KWILSON | 2013-07-16 | Please retap bad hole. Thanks. | Use As Is |  | Re-tapped hole as per NCR Disposition and tested with sample bolt. OK now | WILCOX | 2013-07-18 |  | Bolt hole ok. | GEORGED | 2013-08-09
44 | JL012 |  | unkn | unkn |  | F100-CAV-LAP | R2 | 8 | Field probe flange has machining marks and rough spots in some areas. | GEORGED | 2013-07-16 | 0C1177F3-BE61-DB96-4701D5499A8E357D |  | CARPENTE | 2013-08-07 |  | Re-process |  | Both HOM flange sealing surfaces need to be reworked. The JLAB machine shop has fixturing from C100 cavities that can be used to correct the pitted flanges. The cavity should be taken to the machine shop for repair. SC | CASTAGNO | 2013-08-07 |  | HOM and field probe flanges have been inspected after rework on the CMM, deviations have been found and need to be addressed. Visual acceptable. RF and seal surface dimensional variations have been addressed as per attached documents. | CARPENTE | 2013-07-23 | 7480B265-0123-5125-D48C1BC7E11F9844 | CARPENTE | 2013-08-08 |  | Use As Is
45 | G096314-12 |  | unkn | unkn |  | F100-CAV-INSP-WGD | R1 | 13 | One bad threaded hole in the large flange. | GEORGED | 2013-07-16 |  |  | KWILSON | 2013-07-16 | Please retap. Thanks. | Use As Is |  | Re-tapped hole as per NCR Disposition and tested with sample bolt. OK now | WILCOX | 2013-07-18 |  |  | WILCOX | 2013-07-18
46 | G103046-07 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | F100-CAV-INSP-WGD | R1 | 14 | Scratches present on small flange. | GEORGED | 2013-07-16 |  |  | KWILSON | 2013-07-16 | Please gently polish small flange to remove scratches. Thanks. | Use As Is |  | Dremel tool and Scotchs Brite. | DANIELS | 2013-09-19 |  | Visual ok. Flatness = .002" Finish = 5 uin. | CARPENTE | 2013-09-19 |  | KWILSON | 2013-09-27 |  | Use As Is
47 | G103046-02 |  | unkn | unkn |  | F100-CAV-INSP-WGD | R1 | 15 | Scratches on small flange. Large flange has one bad tapped hole. | GEORGED | 2013-07-17 |  |  | KWILSON | 2013-07-17 | Please gently polish to remove scratches. Please retap bad hole as needed. Thanks. Edit: Please do not use due to oversized bolt hole. Please keep in case we get desperate. Thanks. | Reject |  | Re-tapped hole as per NCR Disposition and tested with sample bolt. Hole OK now. Small flange has not been reworked yet! | WILCOX | 2013-07-18 |  | Small flange is good. Surface finish measured 4 microinches. Flatness measured 0.001". The retapped hole is bad (oversized), nogo gauge fits. | GEORGED | 2013-08-14 |  | DANIELS | 2013-08-08 | Use dremel tool and scotch-brite to remove scratches | Proceed
48 | G103046-06 | WGDSN | unkn | unkn |  | F100-CAV-INSP-WGD | R1 | 18 | Scratches are in the seal area on the small flange. | GEORGED | 2013-07-17 |  |  | KWILSON | 2013-07-17 | Please gently polish to remove scratches. | Use As Is |  | Dremel and hand lap | CASTAGNO | 2013-09-18 |  | Visual good. Flatness = .001" Finish = 6 uin. | CARPENTE | 2013-09-18
49 | JL016 | CAVSN | unkn | unkn |  | F100-CAV-VTA-HPRF | R1 | 4 | Cavity lock freq = 1496.2702 MHz. Specification is 1496.400 to 1496.700 | KDAVIS | 2013-07-18 |  |  | HAIPENG | 2013-09-11 | The lower frequency by 230-250kHz than the target range can be re-tuned back by the tuner preset during the string assembly, according to John Fisher, that means the cavity would be longer by ~1mm than the original when it is under the tension. | Use As Is
50 | JL014 | CAVSN | unkn | unkn |  | F100-CAV-VTA-HPRF | R1 | 5 | Cavity lock frequency was 1496.307; specification was 1496.4 to 1496.7 MHz | KDAVIS | 2013-07-25 |  | Tuner assembly preset can tune the target frequency back | HAIPENG | 2013-10-28 | The lower frequency by 230-250kHz than the target range can be re-tuned back by the tuner preset during the string assembly, according to John Fisher, that means the cavity would be longer by ~1mm than the original when it is under the tension. | Use As Is
51 | JL014 | CAVSN | unkn | unkn |  | F100-CAV-VTA-HPRF | R1 | 5 | Lorenzt coefficient was -1.46 Hz/((MV/m)^2), specification -2 to -3. Scaling 29.4 max gradient to sqrt(-1.46/2.5) means corrected max gradient was 22.5. Actual Lorentz coefficient for these cavities will not be known we have tested more cavities. | KDAVIS | 2013-07-25 |  |  | HAIPENG | 2013-10-28 |  | Use As Is
52 | JL010 |  | unkn | unkn |  | F100-CAV-LAP | R2 | 10 | Significant pits are present on FPC Beam line flange | CARPENTE | 2013-07-30 | 11C6D961-F22E-9CF8-E4CC2FE4C8220A53 |  | CASTAGNO | 2013-08-26 | FPC beam line flange is in the same condition as it was when it was assembled as a string. No leaks were detected after the initial string assembly. Please us the cavity as is. | Use As Is
53 | JL016 | CAVSN | unkn | unkn |  | F100-CAV-VTA-HPRF | R1 | 6 | Cavity Q0 at 20 MV/m was 6.98e+9, below the expected value of 8e9. Low field Q0 was 7.53e+9. | KDAVIS | 2013-08-07 |  | The cavity should use Nb Tophat to retest it | HAIPENG | 2013-10-28 |  | Modify
54 | JL016 | CAVSN | unkn | unkn |  | F100-CAV-VTA-HPRF | R1 | 6 | Cavity lock frequency was 1496.256 MHz; specifcation is 1496.4 to 1496.7 MHz | KDAVIS | 2013-08-07 |  |  | HAIPENG | 2013-09-11 | The lower frequency by 230-250kHz than the target range can be re-tuned back by the tuner preset during the string assembly, according to John Fisher, that means the cavity would be longer by ~1mm than the original when it is under the tension. | Use As Is
55 | ASSEM0035-202 | WINSN | CRM1207075-0035 | ADAPTER WINDOW ASSEMBLY |  | F100-CAV-INSP-ASSY-WIN | R1 | 22 | Nick on knife edge, repairable | CARPENTE | 2013-08-28 |  |  | HKPARK | 2013-09-09 |  | Modify |  | REPAIRED KNIFE EDGE AS PER NCR CONSULTED WITH MIKE MCCREA AFTER REPAIR ON THIS ONE ALL OK | WILCOX | 2013-09-11 |  | Ok | CARPENTE | 2013-09-19 |  | HKPARK | 2013-09-24 |  | Use As Is
56 | JL011 | CAVSN | unkn | unkn |  | F100-CAV-VTA-HPRF | R1 | 7 |  | STIRBET | 2013-08-30 |  | JL011 has lick frequency 1496.2715 MHz out of the specificantion range. | HAIPENG | 2013-09-11 | The lower frequency by 230-250kHz than the target range can be re-tuned back by the tuner preset during the string assembly, according to John Fisher, that means the cavity would be longer by ~1mm than the original when it is under the tension. | Use As Is
57 | JL011 | CAVSN | unkn | unkn |  | F100-CAV-VTA-HPRF | R1 | 7 |  | STIRBET | 2013-08-30 |  | JL011 Emax is 17 MV/m
58 | JL011 | CAVSN | unkn | unkn |  | F100-CAV-VTA-HPRF | R1 | 7 |  | STIRBET | 2013-08-30 |  | JL011 Lorentx coefficent is -2.6704 - out of specification range. | HAIPENG | 2013-10-28 |  | Use As Is
59 | ASSEM0035-209 | WINSN | CRM1207075-0035 | ADAPTER WINDOW ASSEMBLY |  | F100-CAV-INSP-ASSY-WIN | R1 | 5 | (2) nicks on knife edge | CARPENTE | 2013-09-09 |  |  | HKPARK | 2013-09-09 | Please modify if the damage is repairable. | Modify |  | REPAIRED KNIFE EDGE AS PER NCR | WILCOX | 2013-09-11 |  | Ok | CARPENTE | 2013-09-19 |  | HKPARK | 2013-09-24 |  | Use As Is
60 | ASSEM0035-207 | WINSN | CRM1207075-0035 | ADAPTER WINDOW ASSEMBLY |  | F100-CAV-INSP-ASSY-WIN | R1 | 24 | Nicks on knife edge | CARPENTE | 2013-09-09 |  |  | HKPARK | 2013-09-09 | Please modify if the damage is repairable. | Modify |  | REPAIRED KNIFE EDGE AS PER NCR | WILCOX | 2013-09-11 |  | Ok | CARPENTE | 2013-09-19 |  | HKPARK | 2013-09-24 |  | Use As Is
61 | G010213-02 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | F100-CAV-INSP-WGD | R1 | 17 | Failed tape and blister test. see pics in traveler | CARPENTE | 2013-09-09 |  |  | KWILSON | 2013-09-25 | Please do not throw out. | Reject
62 | G096314-01 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | F100-CAV-INSP-WGD | R1 | 16 | (3) black spots in wg after blister test | CARPENTE | 2013-09-09 |  |  | KWILSON | 2013-09-25 | Let's see if Leonard can run a RF test to look for heating at the black spots. (Don't need to do this if we don't need the WG.) | Modify
63 | G010213-01 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | F100-CAV-INSP-WGD | R1 | 11 | Failed tape and blister test | CARPENTE | 2013-09-10 |  |  | KWILSON | 2013-09-25 | Please do not throw out. | Reject
64 | F356224-06 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | F100-CAV-INSP-WGD | R1 | 8 | Failed tape test. | CARPENTE | 2013-09-10 |  |  | KWILSON | 2013-10-07 | Please scotchbrite flange face as needed to remove loose copper. | Modify |  | Loose copper plating removed on face of small flange. | CARPENTE | 2013-10-14 |  | Done very well. | CARPENTE | 2013-10-14 |  | KWILSON | 2013-10-14 |  | Use As Is
65 | G096314-12 | WGDSN | CRM1207070-0000 | 12GEV WAVEGUIDE ASSEMBLY |  | F100-CAV-INSP-WGD | R1 | 13 | Issues after blister test | CARPENTE | 2013-09-10
66 | JL013 |  | unkn | unkn |  | F100-CAV-VTA-HPRF | R1 | 9 | Lorentz coefficient -2.47 out of specification (-6,-4) range | STIRBET | 2013-10-04 |  | LFD spec is wrong here. With the OC cavity shape and HV, this number could be higher but with the tuner installed could be going down. New digital LLRF can overcome this LFD for restart from a trip. | HAIPENG | 2013-10-28 |  | Use As Is
67 | G103046-07 |  | unkn | unkn |  | F100-CAV-INSP-WGD | R1 | 14 | Blisters in wg plating, and stains on stainless. please see pictures in traveler | CARPENTE | 2013-10-04
68 | G103046-04 |  | unkn | unkn |  | F100-CAV-INSP-WGD | R1 | 5 | Scratches in seal area and dark spot on wg. Please see pictures in traveler. | CARPENTE | 2013-10-04 |  |  | KWILSON | 2013-10-07 | Please polish to remove scratches on small flange. Discolored spot on plating ok. | Modify |  | Used dremel tool to remove scratches. ID3 Used scotch brite to finish surface. SC | DANIELS | 2013-10-10 |  | visual good, finish= 4, flatness= 0.002" | CARPENTE | 2013-10-14 |  | KWILSON | 2013-10-14 |  | Use As Is
69 | G103046-03 |  | unkn | unkn |  | F100-CAV-INSP-WGD | R1 | 10 | lots of small blisters at opening of wg. please see traveler for pictures. | CARPENTE | 2013-10-07 |  |  | KWILSON | 2013-10-07 |  | Reject
70 | G103046-06 |  | unkn | unkn |  | F100-CAV-INSP-WGD | R1 | 18 | Blisters found after baking. Please see picture in traveler. | CARPENTE | 2013-10-07
71 | G103046-01 |  | unkn | unkn |  | F100-CAV-INSP-WGD | R1 | 9 | Blisters and small black spots after blister test. Please see pictures in traveler | CARPENTE | 2013-10-07 |  |  | KWILSON | 2013-10-09 | remove blister, clean and RF test. | Modify |  | blister removed 10/9/13, pictures 8008 - 8011 after blister removal. wg cleaned on 10/10/13. Pictures 001 - 005 after cleaning, new blisters have appeared. | CARPENTE | 2013-10-09 | 3105EA2C-ABD5-1D4E-D6AA375E1F9470DE |  | CARPENTE | 2013-10-09
72 | JL017 |  | unkn | unkn |  | F100-CAV-VTA-HPRF | R1 | 10 |  | STIRBET | 2013-10-16 |  | Lock frequency 1496.27661 out of the specification range 1496.4 - 1496.6 MH. | DRURY | 2013-10-28 |  | Use As Is
73 | JL011 |  | unkn | unkn |  | F100-CAV-VTA-HPRF | R1 | 11 | Lock frequency 1496.32455 MHz out of specification range (1496.4 - 1496.7 MHz). | STIRBET | 2013-10-18 |  | Use as it is, the tuner preset at tuner assembly will take care this small difference. | HAIPENG | 2013-10-28 |  | Use As Is
74 | JL012 |  | unkn | unkn |  | F100-CAV-VTA-HPRF | R1 | 13 | Lock frequency 1496.33588MHz out of specification range 1496.400 - 1496.600 MHz. | STIRBET | 2013-10-30 |  |  | KDAVIS | 2013-12-09 | Cavity tuner can compensate for this discrepancy | Use As Is
75 | JL011 |  | unkn | unkn |  | F100-CAV-VTA-HPRF | R1 | 14 | Cavity lock frequency (measured=1496.324MHz) is out of specification (Low:1496.400MHz, High: 1496.700 MHz) | KDAVIS | 2013-11-04 |  |  | HAIPENG | 2013-11-04 |  | Use As Is
76 | JL011 |  | unkn | unkn |  | F100-CAV-VTA-HPRF | R1 | 14 | Cavity Q0 rolls off steeply at 5.5 MV/m, FE onset at 5 MV/m | KDAVIS | 2013-11-04 |  |  | HAIPENG | 2013-11-04 | Interior optical inspection of this cavity is necessary before next reprocess. | Re-process |  |  | FOLLKIEJ | 2013-11-07 |  | Cavity completed high power RF qualification test in VTA | KDAVIS | 2013-12-09 |  | KDAVIS | 2013-12-09 |  | Proceed
77 | 1032-01 |  | unkn | F100 REN Window |  |  |  |  | nicks on knife edge | CARPENTE | 2013-11-19 |  |  | CARPENTE | 2013-11-19 | repair knife edge | Modify |  | Repaired Knife Edge as per NCR disposition | WILCOX | 2013-11-19 |  |  | CARPENTE | 2013-11-19 |  | CARPENTE | 2013-11-19 |  | Use As Is
78 | 1032-02 |  | unkn | unkn | [5 empty] | CARPENTE | 2013-11-19 |  |  | CARPENTE | 2013-11-19 | please repair knife edge | Modify |  | repaired knife edge as per NCR | WILCOX | 2013-11-19 |  |  | CARPENTE | 2013-11-19 |  | CARPENTE | 2013-11-19 |  | Use As Is
79 | 1032-02 |  | unkn | unkn |  |  |  |  | nicks on knife edgo | CARPENTE | 2013-11-19
80 | 1032-02 |  | unkn | unkn |  |  |  |  | nicks on knife edge | CARPENTE | 2013-11-19
81 | JL011 |  | unkn | unkn |  | F100-CAV-VTA-HPRF | R1 | 17 | cavity lock frequency 1496.232MHz out of specification range of 1496.4 to 1496.7 MHz | KDAVIS | 2013-11-21 |  |  | KDAVIS | 2013-12-09 | Cavity tuner can accommodate this discrepancy | Use As Is
82 | JL011 |  | unkn | unkn |  | F100-CAV-VTA-HPRF | R1 | 18 | cavity lock frequency 1496.224 was out of the specification range | KDAVIS | 2013-11-26 |  |  | KDAVIS | 2013-11-26 | cavity tuner can accommodate this discrepancy | Use As Is
83 | JL011 | CAVSN | 115400-1004 | CRYOCAVITY WELDMENT ASSEMBLY |  | F100-CAV-VTA-HPRF | R1 | 16 | Cavity lock frequency 1496.22379 is outside of the specification range of 1496.4 to 1496.7 MHz | KDAVIS | 2013-12-09 |  |  | KDAVIS | 2013-12-09 | Cavity Tuner can compensate for this discrepancy | Use As Is
84 | JL017 |  | unkn | unkn |  | F100-CAV-VTA-HPRF | R1 | 10 | Q0 7.62e9 at 15 MV/m is below specification value of 8e9 | KDAVIS | 2013-12-09 |  |  | KDAVIS | 2013-12-09 | Cavity performance is within error bar of Q0/Eacc specifications. Use as is | Use As Is

## F100-NCR
Status Report for F100-NCR on 10-Jun-20
TRAV_SEQ_NUM | SERIALID | SNTYPE | DRAWINGNUMBER | PARTDESCRIPTION | CONNECTIVITYCOMMENT | TRAVELERID | TRAVELERREVISION | TRAVSEQNUMBER | DESCRIBECOMMENT | DESCRIBESTAFF | DESCRIBEDATE | DESCRIBEDOCS | NOTIFICATIONCOMMENT | RESPONDINGENGINEER | RESPONDINGDATE | DISPOSITIONCOMMENT | DISPOSITION | DISPOSITIONDOCS | REWORKCOMMENT | REWORKSTAFF | REWORKDATE | REWORKDOCS | VERIFYCOMMENT | VERIFYSTAFF | VERIFYDATE | VERIFYDOCS | FINALENGINEER | FINALDATE | FINALCOMMENT | FINALDISPOSITION | FINALDOCS
1 |  | TEST | TEST | TEST | TEST | F100-CM-DISA | R1 | 1 | TEST | BOOKWALT | 2012-03-02 |  | TEST
3 |  |  | unkn | unkn |  | F100-CM-INSP-OMAG | R1 | 2 | Some of the parts are contaminated and many of the shields have extra holes that aren't shown on the drawings. 6 of 7 tophats arrived for drawing #115650-1086D, 1 missing. | GEORGED | 2013-01-10 | 73FA5848-0D55-4715-5DA5E40F177B3098 |  | CHENG | 2013-01-15 | For the contaminated parts, bag and tag with wording like "F100 contaminated outer magnetic shields. DO NOT USE!" | Reject
4 |  |  | unkn | unkn |  | F100-CM-INSP-GVWG | R1 | 2 | Please see traveler. thank you. | CARPENTE | 2013-07-02 |  |  | KWILSON | 2013-07-02 | Small gap in copper on flange face ok. If possible, please use fine sandpaper or similar to gently remove bmp/bubble in copper inside tube. Small bare spot acceptable there. Please remove copper on knife edge as required for sealing. Out of tolerance bolt hole ok. Please retap tapped holes as required. | Use As Is |  | RAN TAP THROUGH BAD HOLES AND TESTED WITH SAMPLE SILVER PLATED BOLT ALL OK. CAREFULLY REMOVED EXCESS COPPER PLATING ON KNIFE EDGE. | WILCOX | 2013-08-06 |  | Tapped holes and knife edge okay. | GEORGED | 2013-08-15
5 |  |  | unkn | unkn |  | F100-CM-INSP-GVWG | R1 | 3 | Plating and dimensional issues. Please see traveler. thank you. | CARPENTE | 2013-07-02 |  |  | KWILSON | 2013-07-02 | Dimensions ok. Please remove plating from knife edge as required to seal. Please clean extra-thoroughly before baking and re-examine black spots in plating after baking. Thanks. | Use As Is |  | REPAIRED KNIFE EDGE AS PER DISPOSTION | WILCOX | 2013-08-06 |  |  | CARPENTE | 2013-09-10
6 |  |  | unkn | unkn |  | F100-CM-INSP-GVWG | R1 | 4 | Plating and dimensional issues. Please see traveler. Thank you. | CARPENTE | 2013-07-02 |  |  | KWILSON | 2013-07-03 | Plating ok. Dimensions ok. Please clean up knife edge as required. Thanks. | Use As Is |  | CLEANED UP KNIFE EDGE | WILCOX | 2013-08-06 |  | Knife edge okay | GEORGED | 2013-08-15
7 |  |  | unkn | unkn |  | F100-CM-INSP-GVWG | R1 | 5 | Plating, dimensional, and bolt holes issues. Please see traveler. Thanks. | CARPENTE | 2013-07-02 |  |  | KWILSON | 2013-07-03 | Plating ok. Dimensions ok. Please clean up knife edge as required. Please retap holes as required. Thanks. | Use As Is |  | RAN TAP THROUGH BAD HOLES AND TESTED WITH SAMPLE SILVER PLATED BOLT ALL OK. CLEANED UP COPPER PLATING FROM KNIFE EDGE. | WILCOX | 2013-08-06 |  | Knife edge ok. Tapped holes ok. | GEORGED | 2013-08-15
8 |  |  | unkn | unkn |  | F100-CM-INSP-GVWG | R1 | 6 | Dis-coloration on mini-CF weld. | CARPENTE | 2013-07-08 |  |  | KWILSON | 2013-07-09 |  | Use As Is
9 |  | GVWGSN | CRM1207075-0040 | WAVEGUIDE SPOOL - LEFT |  | F100-CM-INSP-GVWG | R1 | 7 | Missing copper inside of miniCF tube | CARPENTE | 2013-07-11 |  |  | KWILSON | 2013-07-11 |  | Use As Is
10 |  |  | unkn | unkn |  | F100-CM-INSP-GVWG | R1 | 8 | gouges or blisters inside of waveguide. Blisters in copper plating. Conflat locations out of tolerance. see report below. | CARPENTE | 2013-07-11 |  |  | KWILSON | 2013-07-24 | Plating does not look good due to gouges in the plating. Prefer not to use. Please hold on to part in case we run out of other options and have to try to clean it up. | Reject
11 |  |  | unkn | unkn |  | F100-CM-INSP-GVWG | R1 | 9 | Black stains in waveguide and mini conflat tubes. copper on knife edge. Nick on knife edge, tapped flange. | CARPENTE | 2013-07-11 |  |  | KWILSON | 2013-07-11 | Please try to gently remove copper from knife edge and repair nick to knife edge if possible. Discoloration of plating ok. | Use As Is |  | REPAIRED KNIFE EDGE AS PER DISPOSITION | WILCOX | 2013-08-06 |  | Knife edge ok | GEORGED | 2013-08-15
12 |  |  | unkn | unkn |  | F100-CM-INSP-GVWG | R1 | 10 | 1 bad tapped hole | CARPENTE | 2013-07-11 |  |  | KWILSON | 2013-07-11 | Please retap as required. | Use As Is |  | Re-tapped hole as per NCR Disposition and tested with sample bolt. OK now | WILCOX | 2013-07-18 |  | Tapped hole ok | GEORGED | 2013-08-15
13 |  |  | unkn | unkn |  | F100-CM-INSP-GVWG | R1 | 11 | Conflat locations out of tolerance. (1) bad tapped hole. | CARPENTE | 2013-07-11 |  |  | KWILSON | 2013-07-24 | Please retap bad hole. | Use As Is |  | RETAPPED MARKED HOLE AND TESTED WITH SAMPLE BOLT. OK | WILCOX | 2013-08-07 |  | Tapped hole ok | GEORGED | 2013-08-15
14 |  |  | unkn | unkn |  | F100-CM-INSP-GVWG | R1 | 12 | 5 tapped holes are bad. Cu plating is on the knife edge of a 1.33 mini conflat. Plating inconsistent on face of flange. Flaking inside tube. | GEORGED | 2013-07-23 | 822F4AC7-F514-3448-D3E7493E0F0F15C0 |  | KWILSON | 2013-07-23 | Please retap holes as needed and carefully clean up knife edge. | Use As Is |  | REPAIRED THREADS OF SAID TAPPED HOLES AND TESTED WITH SAMPLE SILVER PLATED BOLT ALL OK. CAREFULLY REMOVED COPPER PLATING ON SAID KNIFE EDGE. | WILCOX | 2013-08-06 |  |  | WILCOX | 2013-08-06
15 |  |  | unkn | unkn |  | F100-CM-INSP-GVWG | R1 | 13 | Fingerprints and discoloration on copper plating. | GEORGED | 2013-07-25 | D8CAA6A9-E5FD-76B4-EB908A5C6626BF3F |  | KWILSON | 2013-07-25 |  | Use As Is
16 |  |  | unkn | unkn |  | F100-CM-INSP-GVWG | R1 | 14 | Two tapped holes are bad. Copper flaking inside tube. | GEORGED | 2013-07-26 | 0C235CD2-EEB8-0A69-CFFECCEB3B4B10C1 |  | KWILSON | 2013-07-26 | Please retap holes as needed. | Use As Is |  | RETAPPED MARKED HOLES AND TESTED WITH SAMPLE BOLT. THREADS OK NOW | WILCOX | 2013-08-07 |  | Tapped holes ok | GEORGED | 2013-08-15
17 |  |  | unkn | unkn |  | F100-CM-INSP-GVWG | R1 | 15 | 4 bad tapped holes. Copper present on knife edge (1.33 CF). | GEORGED | 2013-07-26 |  |  | KWILSON | 2013-08-05 | Please retap holes and clean up knife edge. | Use As Is |  | RAN TAP THROUGH MARKED HOLES AND TESTED WITH A SAMPLE BOLT. ALL OK. CLEANED UP KNIFE EDGE AS BEST AS POSSIBLE. | WILCOX | 2013-08-07 |  | Tapped holes ok. Knife edge ok. | GEORGED | 2013-08-14
18 |  |  | unkn | unkn |  | F100-CM-INSP-GVWG | R1 | 16 | Large gouge/dent in copper plating. Blisters and rough spots present. | GEORGED | 2013-08-05 | 1FE60E01-B3AC-B04B-1A1E031105F6756D
19 |  |  | unkn | unkn |  | F100-CM-INSP-GVWG | R1 | 17 | Two bad tapped holes. | GEORGED | 2013-08-06 |  |  | KWILSON | 2013-08-06 | Please retap bad holes. | Use As Is |  | RETAPPED MARKED HOLES AND TESTED WITH SAMPLE BOLT. ALL OK NOW | WILCOX | 2013-08-07 |  | Tapped holes ok | GEORGED | 2013-08-15
20 |  | GVWGSN | unkn | unkn |  | F100-CM-INSP-GVWG | R1 | 14 | Issues after blister test. | CARPENTE | 2013-09-10 |  |  | KWILSON | 2013-10-07 |  | Use As Is
24 |  | GVWGSN | unkn | unkn |  | F100-CM-INSP-GVWG | R1 | 7 | Black spot in wg. Please see picture in inspection traveler. | CARPENTE | 2013-09-16 |  |  | KWILSON | 2013-09-25 |  | Use As Is
25 |  |  | unkn | unkn |  | F100-CM-INSP-GVWG | R1 | 11 | Streaks inside of wg after cleaning and blister test. Black spots still on face of flange. see traveler for pics | CARPENTE | 2013-09-25 |  |  | KWILSON | 2013-09-25 |  | Use As Is
26 |  |  | unkn | unkn |  | F100-CM-INSP-GVWG | R1 | 15 | Plating looks thin at end of wg. please see pics in traveler. | CARPENTE | 2013-09-25 |  |  | KWILSON | 2013-09-25 |  | Use As Is
27 |  |  | unkn | unkn |  | F100-CM-INSP-GVWG | R1 | 2 | Dark streaks in plating | CARPENTE | 2013-09-25 |  |  | KWILSON | 2013-09-25 | 1. Please repair tapped bolt holes. 2. Please scotch brite off small blister in small tube. 3. We can live with stains on copper. Please clean well. Thanks. | Modify
28 |  | GVWGSN | unkn | unkn |  | F100-CM-INSP-GVWG | R1 | 13 | Dark streaks in wg, black spots on face of wg | CARPENTE | 2013-09-25 |  |  | KWILSON | 2013-10-07 |  | Use As Is
29 |  |  | unkn | unkn |  | F100-CM-INSP-GVWG | R1 | 10 | Dis-coloration and black spots in plating | CARPENTE | 2013-10-07
30 |  |  | unkn | unkn |  | F100-CM-INSP-GVWG | R1 | 12 | Dis-coloration or small blisters in wg after blister test. Please see traveler for pictures | CARPENTE | 2013-10-07
31 |  |  | unkn | unkn |  | F100-CM-INSP-GVWG | R1 | 3 | Dis-coloration and black spots in wg. Please see traveler for pictures. | CARPENTE | 2013-10-07 |  |  | KWILSON | 2013-10-07 |  | Use As Is
32 |  |  | unkn | unkn |  | F100-CM-INSP-THTD | R1 | 1 | OD of tube circularity and loose flange ID out of tolerance. see report. | CARPENTE | 2013-11-08 | 322A6646-E340-AE19-071DA8B9F6031BBC |  | KWILSON | 2013-11-14 |  | Return To Vendor
33 |  |  | unkn | unkn |  | F100-CM-INSP-THTS | R1 | 4 | OD of tube circularity and loose flange ID out of tolerance. see report. | CARPENTE | 2013-11-08 | 32526256-FF9F-8B43-11F8D29E8828A35E |  | KWILSON | 2013-11-14 |  | Return To Vendor
34 |  |  | unkn | unkn |  | F100-CM-INSP-THTS | R1 | 3 | OD of tube circularity and loose flange ID out of tolerance. see report. | CARPENTE | 2013-11-08 | 325D51C6-ECCF-5E38-94F6A021ED210B92 |  | KWILSON | 2013-11-14 |  | Return To Vendor
35 |  |  | unkn | unkn |  | F100-CM-INSP-THTS | R1 | 2 | OD of tube circularity and loose flange ID out of tolerance. see report. | CARPENTE | 2013-11-08 | 326BB5D2-AC89-DDEE-70FCA0CB0DF43979 |  | KWILSON | 2013-11-14 |  | Return To Vendor
36 |  |  | unkn | unkn |  | F100-CM-INSP-THTS | R1 | 1 | OD of tube circularity and loose flange ID out of tolerance. see report. | CARPENTE | 2013-11-08 | 32794170-036A-D655-10A4EBA713C16C1B |  | KWILSON | 2013-11-14 |  | Return To Vendor
37 |  |  | unkn | unkn |  | F100-CM-INSP-THTS | R1 | 6 | OD of tube circularity and loose flange ID out of tolerance. see report. | CARPENTE | 2013-11-08 | 3281B478-D76B-C8BD-93CBA54FAAD3C267 |  | KWILSON | 2013-11-14 |  | Return To Vendor
38 |  |  | unkn | unkn |  | F100-CM-INSP-THTS | R1 | 5 | OD of tube circularity,loose flange ID, and several bolt holes out of tolerance. see report. | CARPENTE | 2013-11-08 | 3291FD2B-01F6-7745-E5419C80FC8F044B |  | KWILSON | 2013-11-14 |  | Return To Vendor

## R100-CAV-NCR
Status Report for R100-CAV-NCR on 10-Jun-20
TRAV_SEQ_NUM | IMMEDIATE_ATTENTION_REQUIRED | SERIALID | VENDORID | DRAWINGNUMBER | PARTDESCRIPTION | CONNECTIVITYCOMMENT | TRAVELERID | TRAVELERREVISION | TRAVSEQNUMBER | DESCRIBECOMMENT | SRF_CVP_USER | NCR_ISSUE_DATE | DESCRIBEDOCUMENTS | NOTIFICATIONCOMMENT | ENGINEERTOEMAIL | RESPONDINGENGINEER | NCR_RESPONSE_DATE | DISPOSITIONCOMMENT | NCR_DISPOSITION | DOCUMENTATIONCOMMENT | NCR_DOCUMENTS
2 |  | R100-002 |  | CRM-120-7013-0001 | C100 Cryomodule Cavity |  |  |  |  | FPC flange perpendicularity and End Flange True Position (FPC End) are out of tolerance | GEORGED | 2010-07-08 |  |  | drury | DRURY |  |  | Use As Is
3 |  | R100-002 |  | CRM-120-7013-0001 | C100 Cryomodule Cavity |  | R100-CAV-INSP2 | R1 | 1 | FPC Flange Perpendicularity and End Flange True Position (FPC End) are out of tolerance | GEORGED | 2010-07-08 |  |  | drury | DRURY |  |  | Use As Is
4 |  | R100-004 |  | CRM12070130001 | R100 cavity |  | R100-CAV-INSP2 | R1 | 4 | FPC Flange Height and End Flange True Position (FPC End) are out of tolerance | GEORGED | 2010-07-22 |  |  | drury | DRURY |  |  | Use As Is
5 |  | R100-005 |  | CRM12070130001 | R100 cavity |  | R100-CAV-INSP2 | R1 | 5 | Overall length and End Flange True Position (HOM End) are out of tolerance | GEORGED | 2010-07-23 |  |  | drury | DRURY |  |  | Use As Is
6 |  | R100-06 |  | CRM12070130001 | R100 Cavity |  | R100-CAV-INSP2 | R1 | 6 | FPC Flange Height is out of tolerance | GEORGED | 2010-07-27 |  |  | drury | DRURY |  |  | Use As Is
7 |  | R100-07 |  | CRM12070130001 | R100 Cavity |  | R100-CAV-INSP2 | R1 | 7 | FPC Flange Height, End Flange True Position (FPC End), and end flange true position (HOM End)are out of tolerance | GEORGED | 2010-08-04 |  |  | drury | DRURY |  |  | Use As Is
8 | [6 empty] | R100-CAV-VTRF | R1 | 1 | lock frequency out of specification high |  | 2010-08-20 |  |  | kdavis | DRURY |  |  | Use As Is
11 |  | 47726-0010 |  | CRM-120-7014-0020 | HOM PROBE FEED THRU |  | R100-CAV-INSP-HMFT | R2 | 16 | Center connector pin slightly off-center. | OGLE | 2010-09-02 |  |  |  | HKPARK | 2010-09-23 | If the center pin fits the gauge, center the pin carefully. | Use As Is
12 |  | 47726-0011 |  | CRM-120-7014-0020 | HOM PROBE FEED THRU |  | R100-CAV-INSP-HMFT | R2 | 17 | Center connector pin slightly off-center. Antenna length 1.459/1.469 out of tolerance. | OGLE | 2010-09-02 |  |  |  | HKPARK | 2010-11-03 | 1.464+/-.010 is acceptable. | Use As Is
13 |  | 47726-0012 |  | CRM-120-7014-0020 | HOM PROBE FEED THRU |  | R100-CAV-INSP-HMFT | R2 | 18 | Center connector pin very slightly off-center. | OGLE | 2010-09-02 |  |  |  | HKPARK | 2010-09-23 | If the center pin fits the gauge, center the pin carefully. | Use As Is
14 |  | 47726-0013 |  | CRM-120-7014-0020 | HOM PROBE FEED THRU |  | R100-CAV-INSP-HMFT | R2 | 19 | Center connector pin off-center. Antenna length 1.459/1.469 out of tolerance. | OGLE | 2010-09-02 |  |  |  | HKPARK |  | If the pin fits in the gauge, center the pin carefully. If not, do not use. The disposition of the out of tolerance dimension will be decided at the time of assembly. The critical dimension is .04" from the probe tip to the inner conductor which is indeterminate at this point. | Use As Is
15 |  | 47726-0014 |  | CRM-120-7014-0020 | HOM PROBE FEED THRU |  | R100-CAV-INSP-HMFT | R2 | 20 | Antenna length 1.459/1.469 out of tolerance. | OGLE | 2010-09-02 |  |  |  | HKPARK | 2010-09-23 | If the pin fits in the gauge, center the pin carefully. If not, do not use. The disposition of the out of tolerance dimension will be decided at the time of assembly. The critical dimension is .04" from the probe tip to the inner conductor which is indeterminate at this point. | Use As Is
16 | [6 empty] | R100-CAV-VTRF | R1 | 13 |  |  | 2010-09-07 |  | Lock frequency for R100_1 with He vessel out of specifications |  | DRURY | 2010-09-08 | We have to use as is and hope that the cavity can be tuned correctly. | Use As Is
17 |  | R100_1 |  |  | R100 Cavity |  | R100-CAV-VTRF | R1 | 14 | Cavity lock frequency out of specification, low |  | 2010-09-13 |  |  | areilly | DRURY | 2012-01-05 |  | Use As Is
18 |  | R100_1 |  |  | R100 Cavity |  | R100-CAV-VTRF | R1 | 14 | Radiation burst resulted in the cavity FE at 11MV/m. Recommend reprocessing cavity. |  | 2010-09-13 |  |  | areilly | DRURY | 2010-09-16 | Cavity must undergo high pressure rinse | Modify
19 |  | 1 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | R100-CAV-INSP-ADPT | R1 | 3 | Braze material on outside edge of part.Gap between ceramic and copper.Braze material in gap between copper and stainless adapter. | OGLE | 2010-09-14 |  |  |  | HKPARK | 2010-10-05 | Based on the additional 5 cycle RF test, the amount of braze alloy bridging on this part is acceptable. | Use As Is
20 |  | 4 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | R100-CAV-INSP-ADPT | R1 | 1 | Braze material found in gap between copper and stainless.Braze residue on outside of stainless adapter. | OGLE | 2010-09-14 |  |  |  | HKPARK | 2010-10-05 | Based on the additional 5 cycle RF test, the amount of braze alloy bridging on this part is acceptable. | Use As Is
21 |  | 7 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | R100-CAV-INSP-ADPT | R1 | 4 | Ceramic surface shows slight black particles...not clean. Braze material on outside of stainless adapter. | OGLE | 2010-09-14 |  |  |  | HKPARK | 2010-09-28 | Braze residue on side is OK as long as it fits in the flange. The black mark seems to be OK. | Use As Is
22 |  | 8 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | R100-CAV-INSP-ADPT | R1 | 8 | Braze material found on outside edge of stainless adapter. Note:serial no 03 marked out on bag and remarked 08. Used serial no. 08 for this unit. | OGLE | 2010-09-14 |  |  |  | HKPARK | 2010-09-28 | Braze material on outside edge of adapter is OK as long as it does not interfere the flange fit. | Use As Is
23 |  | 10 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | R100-CAV-INSP-ADPT | R1 | 9 | Braze material on outside of stainless adapter | OGLE | 2010-09-14 |  |  |  | HKPARK | 2010-09-28 | Braze material on outside edge of adapter is OK as long as it does not interfere the flange fit. | Use As Is
24 |  | 13 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | R100-CAV-INSP-ADPT | R1 | 10 | Braze material on outside of stainless adapter. Braze material in gap between copper and stainless adapter. | OGLE | 2010-09-14 |  |  |  | HKPARK | 2010-10-05 | Based on the additional 5 cycle RF test, the amount of braze alloy bridging on this part is acceptable. | Use As Is
25 |  | 17 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | R100-CAV-INSP-ADPT | R1 | 5 | Braze residue on outside surface of stainless adapter. Braze material in gap between copper and stainless adapter. | OGLE | 2010-09-14 |  |  |  | HKPARK | 2010-10-05 | Based on the additional 5 cycle RF test, the amount of braze alloy bridging on this part is acceptable. | Use As Is
26 |  | 18 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | R100-CAV-INSP-ADPT | R1 | 7 | Braze residue on outside surface of stainless adapter. Braze material in gap between copper and stainless adapter. | OGLE | 2010-09-14 |  |  |  | HKPARK | 2010-10-05 | Based on the additional 5 cycle RF test, the amount of braze alloy bridging on this part is acceptable. | Use As Is
27 |  | 20 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | R100-CAV-INSP-ADPT | R1 | 6 | Braze material in gap between copper and stainless adapter. | OGLE | 2010-09-14 |  |  |  | HKPARK | 2010-10-05 | Based on the additional 5 cycle RF test, the amount of braze alloy bridging on this part is acceptable. | Use As Is
28 |  | 21 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | R100-CAV-INSP-ADPT | R1 | 2 | Light smudge mark on ceramic surface. Braze residue on outside surface of stainless adapter. | OGLE | 2010-09-14 |  |  |  | HKPARK | 2010-09-28 | The smudge mark is OK since this adapter passed RF test. | Use As Is
29 |  | MPF-0003 |  | CRM12070140020 | HOM Feedthru |  | R100-CAV-INSP-HMFT | R1 | 12 | Antennae angle is 1.8 degrees. should be .5 max. | CARPENTE | 2010-09-20 |  |  |  | HKPARK | 2010-11-03 | Do not use. Due to the angle the probe tip location might not be optimum to the inner conductor flat surface. Qurantine. | Scrap
30 |  | MPF-0005 |  | CRM12070140020 | HOM Feedthru |  | R100-CAV-INSP-HMFT | R1 | 14 | Antenna angle is 1.3, should be .5 max. | CARPENTE | 2010-09-20 |  |  |  | HKPARK | 2010-11-03 | Do not use. Due to the angle the probe tip location might not be optimum to the inner conductor flat surface. Qurantine. | Scrap
31 |  | 47726-0011 |  | CRM12070140020 | HOM Feedthru |  | R100-CAV-INSP-HMFT | R2 | 17 | Antenne lenght is 1.456, should be 1.459/1.469 | CARPENTE | 2010-09-20 |  |  |  | CARPENTE | 2011-09-20 | Used on R100 cavity string | Use As Is
32 |  | FLG0037-012 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | R100-CAV-INSP-FLNG | R1 | 2 | CIRC18 OUTTOLZ.009 CIRC20 OUTTOLZ.003 CIRC22 OUTTOLZ.001 CIRC24 OUTTOLZ.005 CIRC26 OUTTOLZ.004 CIRC28 OUTTOLZ.004 CIRC30 OUT Z.005 CIRC32 OUT Z.004 CIRC34 OUT Z.005 CIRC38 OUT Z.002 CIRC40 OUT Z.001 CIRC44 OUT Z.003 CIRC46 OUT Z.007 CIRC48 OUT Z.003 CIRC52 OUT Z.005 | OGLE | 2010-09-23 |  | **TO BE USED AS A WELD TEST*** |  | HKPARK | 2010-10-05 | Use this flange and one Omley adapter (S/N 47789-003) for a weld practice. After welding proceed leak test and label as "reject". Do not use. | Scrap
33 |  | FLG0037-014 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | R100-CAV-INSP-FLNG | R1 | 4 | THICKNESS1.360 +/-.005; OUT OF TOL. CHECKS 1.368 | OGLE | 2010-09-24 |  |  |  | HKPARK | 2010-09-30 | The out tolerance dimension does not affect assembly. It will affect ceramic to ceramic distance if this flange is used for outer window. But the adjacent waveguide has bigger tolerance range on its length so it is not critical. | Use As Is
34 |  | FLG0037-016 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | R100-CAV-INSP-FLNG | R1 | 6 | FLNG THICKNESS 1.360+/-.005 OUT OF TOL. CHECKS 1.366 BORE STEP .705+/-.005 CHECKS .698 OUT OF TOL. .002 | OGLE | 2010-09-24 |  |  |  | HKPARK | 2010-09-30 | Use as long as it does not affect assembly. | Use As Is
39 |  | 5 |  | CRM-120-7075-0024 | WINDOW ADAPTER |  | R100-CAV-INSP-ADPT | R1 | 13 | VERY MINOR BRAZE RESIDUE FOUND ON OUTSIDE SURFACR NEAR THE EDGE. BRAZE MATERIAL IN THE GAP. | OGLE | 2010-09-27 |  |  |  | HKPARK | 2010-10-05 | Based on the additional 5 cycle RF test, the amount of braze alloy bridging on this part is acceptable. | Use As Is
40 |  | 6 |  | CRM-120-7075-0024 | WINDOW ADAPTER |  | R100-CAV-INSP-ADPT | R1 | 14 | BRAZE RESIDUE ON OUTSIDE EDGE BRAZE MATERIAL IN GAP | OGLE | 2010-09-27 |  |  |  | HKPARK | 2010-10-05 | Based on the additional 5 cycle RF test, the amount of braze alloy bridging on this part is acceptable. | Use As Is
41 |  | 11 |  | CRM-120-7075-0024 | WINDOW ADAPTER |  | R100-CAV-INSP-ADPT | R1 | 16 | BRAZE RESIDUE ON OUTSIDE EDGE OF PART. BRAZE MATERIAL IN THE GAP. | OGLE | 2010-09-27 |  |  |  | HKPARK | 2010-10-05 | Based on the additional 5 cycle RF test, the amount of braze alloy bridging on this part is acceptable. | Use As Is
42 |  | 16 |  | CRM-120-7075-0024 | WINDOW ADAPTER |  | R100-CAV-INSP-ADPT | R1 | 19 | BRAZE RESIDUE ON OUTSIDE EDGE OF PART | OGLE | 2010-09-27 |  |  |  | HKPARK | 2010-09-28 | Braze material on outside edge of adapter is OK as long as it does not interfere the flange fit. | Use As Is
43 |  | FLG0037-020 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | R100-CAV-INSP-FLNG | R1 | 10 | FLNG THCK. 1.360 CKS. 1.368 OUT OF TOL .003 CIRC28 Z 1.515 CKS 1.522 OUT OF TOL.002 CIRC 47 X 1.803 CKS 1.809 OUTOL .001 CIRC 48 Z 1.515 CKS 1.521 OUTOL .001 CIRC49 Z 1.515 CKS 1.507 OUTOL .003 | OGLE | 2010-10-06 |  |  |  | HKPARK | 2010-10-07 | The out tolerance range is acceptable. | Use As Is
44 |  | FLG0037-001 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | R100-CAV-INSP-FLNG | R1 | 11 | FLANGE THICKNESS 1.360 CKS. 1.371 OUT OF TOL. .006 CIRC 22 X 2.404 CKS. 2.410 OUT OF TOL .001 CIRC 32 X 3.185 CKS. 3.177 OUT OF TOL. .003 CIRC 36 3.435 CKS. 3.429 OUT OF TOL. .001 | OGLE | 2010-10-13 |  |  |  | HKPARK | 2010-10-15 | Out tol value not affecting assembly | Use As Is
45 |  | FLG0037-002 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | R100-CAV-INSP-FLNG | R1 | 12 | FLANGE THCK. 1.360 CKS. 1.371 OUT OF TOL .006 CIRC30 Z 1.515 CKS. 1.522 OUT/TOL .002 CIRC34 X 3.435 CKS. 3.429 OUT/TOL .001 CIRC36 X 3.435 CKS. 3.427 OUT/TOL. .003 CIRC52 X 3.435 CKS 3.444 OUT/TOL .004 | OGLE | 2010-10-13 |  |  |  | HKPARK | 2010-10-15 | Out tol value not affecting assembly | Use As Is
46 |  | FLG0037-003 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | R100-CAV-INSP-FLNG | R1 | 13 | BORE LENGTH 5.548 CKS. 5.540 OUT/TOL .003 BORE WIDTH 1.732 CKS. 1.724 OUT/TOL .003 FLNG THICK. 1.360 CKS. 1.371 OUT/TOL .006 CIRC18 Z .500 CKS. .493 OUT/TOL .002 CIRC24 Z 1.515 CKS. 1.509 OUT/TOL .001 CIRC26 Z 1.515 CKS. 1.505 OUT/TOL .005 CIRC46 Z 1.515 CKS. 1.524 OUT/TOL .004 CIRC48 Z 1.515 CKS. 1.522 OUT/TOL .002 CIRC52 Z .5 CKS. .506 OUT/TOL .001 | OGLE | 2010-10-13 |  |  |  | HKPARK | 2010-10-15 | Out tol value not affecting assembly | Use As Is
47 |  | FLG0037-004 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | R100-CAV-INSP-FLNG | R1 | 14 | FLANGE THK. 1.360 CKS. 1.372 OUT/TOL .007 CIRC24 X 1.202 CKS. 1.208 OUT/TOL .001 CIRC28 X 1.202 CKS. 1.196 OUT/TOL .001 CIRC28 Z 1.515 CKS. 1.521 OUT/TOL .001 CIRC48 X 2.404 CKS. 2.412 OUT/TOL .003 CIRC50 X 3.185 CKS. 3.191 OUT/TOL .001 CIRC50 Z 1.375 CKS. 1.367 OUT/TOL .003 | OGLE | 2010-10-13 |  |  |  | HKPARK | 2010-10-15 | Out tol value not affecting assembly | Use As Is
48 |  | FLG0037-005 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | R100-CAV-INSP-FLNG | R1 | 15 | FLANGE THCK. 1.360 CKS. 1.372 OUT/TOL .007 CIRC22 X 2.204 CKS. 2.413 OUT/TOL .004 CIRC24 X 1.202 CKS. 1.210 OUT/TOL .003 CIRC28 X 1.202 CKS. 1.196 OUT/TOL .001 CIRC28 Z 1.515 CKS. 1.521 OUT/TOL .001 CIRC36 X 3.435 CKS. 3.428 OUT/TOL .002 CIRC36 Z .5 CKS. .492 OUT/TOL .003 | OGLE | 2010-10-13 |  |  |  | HKPARK | 2010-10-15 | Out tol value not affecting assembly | Use As Is
49 |  | FLG0037-006 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | R100-CAV-INSP-FLNG | R1 | 16 | BORE LENGTH 5.548 CKS. 5.540 OUT/TOL .003 BORE WIDTH 1.732 CKS. 1.719 OUT/TOL .008 FLANGE THCK. 1.360 CKS. 1.371 OUT/TOL .006 CIRC24 X 1.202 CKS. 1.212 OUT/TOL .005 CIRC32 X 3.185 CKS. 3.179 OUT/TOL .001 CIRC50 Z 1.375 CKS. 1.381 OUT/TOL .001 | OGLE | 2010-10-13 |  |  |  | HKPARK | 2010-10-15 | Out tol value not affecting assembly | Use As Is
50 |  | FLG0037-007 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | R100-CAV-INSP-FLNG | R1 | 17 | FLANGE THCK. 1.360 CKS. 1.372 OUT/TOL .007 CIRC18 Z .5 CKS. .493 OUT/TOL .002 CIRC22 Z 1.515 CKS. 1.504 OUT/TOL .006 CIRC25 Z 1.515 CKS. 1.521 OUT/TOL .001 CIRC27 Z 1.515 CKS. 1.521 OUT/TOL .001 CIRC29 Z 1.515 CKS. 1.521 OUT/TOL .001 CIRC36 X 3.435 CKS. 3.429 OUT/TOL .001 CIRC44 Z 1.515 CKS. 1.521 OUT/TOL .001 CIRC46 Z 1.515 CKS. 1.527 OUT/TOL .007 CIRC48 Z 1.515 CKS. 1.522 OUT/TOL .002 | OGLE | 2010-10-13 |  |  |  | HKPARK | 2010-11-24 | The thickness is OK but the hole location seems to be far off. The bolt hole pattern should be checked with the mating flange. Do not use until until it's checked. | Scrap
51 |  | FLG0037-008 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | R100-CAV-INSP-FLNG | R1 | 18 | FLANGE THCK. 1.360 CKS. 1.372 OUT/TOL .007 CIRC18 Z .500 CKS. .493 OUT/TOL .002 CIRC32 X 3.185 CKS. 3.177 OUT/TOL .003 | OGLE | 2010-10-13 |  |  |  | HKPARK | 2010-10-15 | Out tol value not affecting assembly | Use As Is
52 | [6 empty] | R100-CAV-INSP-FLNG | R1 | 19 | BORE LENGTH 5.548 CKS. 5.540 OUT/TOL .003 BORE WIDTH 1.732 CKS. 1.725 OUT/TOL .002 FLANGE THICKNESS 1.360 CKS. 1.371 OUT/TOL .006 CIRC18 Z .500 CKS. .494 OUT/TOL .001 CIRC24 Z 1.202 CKS. 1.196 OUT/TOL .001 CIRC28 Z 1.515 CKS. 1.509 OUT/TOL .001 CIRC28 Z 1.202 CKS. 1.209 OUT/TOL .002 CIRC30 Z 1.515 CKS. 1.509 OUT/TOL .001 CIRC44 Z 1.515 CKS. 1.522 OUT/TOL .002 CIRC50 Z 1.375 CKS. 1.381 OUT/TOL .001 | OGLE | 2010-10-13 |  |  |  | HKPARK | 2010-11-24 | Before release drop a window brazement into the pocket to see if 1.732 width is not too much undersized. If it doesn't fit do not use. | Use As Is
53 |  | FLG0037-010 |  | CRM-120-7075-0037 | DOUBLE SIDED FLANGE |  | R100-CAV-INSP-FLNG | R1 | 20 | FLANGE THICKNESS 1.360 CKS. 1.371 OUT/TOL .006 CIRC22 X 2.404 CKS. 2.411 OUT/TOL. .002 CIRC32 X 3.185 CKS. 3.178 OUT/TOL .002 CIRC34 X 3.435 CKS. 3.425 OUT/TOL .005 CIRC40 Z 1.515 CKS. 1.508 OUT/TOL .002 CIRC48 X 2.404 CKS. 2.413 OUT/TOL .004 CIRC52 X 3.435 CKS. 3.442 OUT/TOL .002 | OGLE | 2010-10-13 |  |  |  | HKPARK | 2010-10-15 | Out tol value not affecting assembly | Use As Is
54 |  | 19 |  | CRM-120-7075-0024 | Window brazement adapter assy. |  | R100-CAV-INSP-ADPT | R1 | 20 | BRAZE MATERIAL IN THE GAP BETWEEN THE COPPER EYELET AND STAINLESS ADAPTER. | OGLE | 2010-10-14 |  |  |  | DRURY | 2012-01-05 |  | Use As Is
55 |  | 5 |  |  | CAVITY |  | R100-CAV-INST-HELV | R1 | 4 | AFTER STANDING CAVITY UP INTO WELD FIXTURE WE SAW THAT THE TOP FLANGE APPEARED TO BE BENT. BEAM PIPE SPOOL AND VALVE ASSY WAS REPLACED AND ANOTHER WAS WAS PLACED ON IT AND THE SAME RESULTS WERE FOUND, WHICH LEADS US TO BELEIVE THAT THE UPPER CAVITY FLANGE IS BENT. | wilcox | 2010-10-25 |  |  |  | DRURY | 2012-01-05 |  | Use As Is
56 |  | R100_004 |  |  | R100 cavity |  | R100-CAV-VTRF | R1 | 24 | cavity quenches and FE at 23MV/m. | forehand | 2010-11-08 |  |  | kdavis | DRURY |  |  | Use As Is
57 |  | 47726-0004 |  | CRM12070140020 | hom feed thru |  | R100-CAV-INSP-HMFT | R1 | 8 | SMALL AMOUNT OF DEBRIS FOUND INTERNAL NEAR CERAMIC SURFACE CONNECTOR END. | ogle | 2010-11-10 |  |  |  | DRURY |  | Clean and use | Modify
58 |  | 47726-0008 |  | CRM120-7014-0020 | hom feed thru |  | R100-CAV-INSP-HMFT | R1 | 4 | SMALL AMOUNT OF DEBRIS INTERNAL NEAR CERAMIC SURFACE;CONNECTOR END. | ogle | 2010-11-10 |  |  |  | HKPARK | 2010-11-11 | Use if the debris can be removed. (use pressured clean air or nitrogen) | Use As Is
59 | [5 empty] | R100_01 sequence 25 | R100-CAV-VTRF | R1 | 25 | Quench processing pushed the cavity gradient over 27 MV/m, but the cavity cannot keep the gradient: after about 2 - 3 minutes of stable gradient, cavity start quenching again. Stable cavity gradient at about 24.5 MV/m | fischer | 2010-11-15 |  |  |  | DRURY |  |  | Use As Is
60 |  | R100_008 |  |  |  | cavity R100_008 with retracted pin on HOM B found during HOM survey on 9th Nov 2010. Cavity with Qo at low fields ( bellow 9 MV/m)while RF tested in Dewar 3. Suppose He Level issues. Cavity will be retested in Dewar 4. HOM survey after replacing the SiO2 cable on HOM B O.K. | R100-CAV-VTRF | R1 | 26 |  |  | 2010-11-15 |  |  |  | DRURY |  |  | Use As Is
61 |  | R100_004 |  |  |  |  | R100-CAV-VTRF | R1 | 28 | cavity quenching repeatedly at 21 MV/m, quenching occassionally at 19.5 MV/m |  | 2010-11-16 |  |  |  | DRURY |  |  | Use As Is
62 |  | 6 |  | CRM1207075-0037 | Window Double Flange |  | R100-CAV-INSP-FLNG | R1 | 16 | Blank adapter does not fit. BC 01/06/11 | carpente | 2011-01-06 |  |  |  | CARPENTE |  | Returned on 7/27/11 | Return To Vendor
63 |  | 10 |  | CRM1207075-0037 | Window Double Flange |  | R100-CAV-INSP-FLNG | R1 | 20 | Blank adapter does not fit. BC 01/06/11 | carpente | 2011-01-06 |  |  |  | CARPENTE | 2011-09-20 | Returned on 7/27/11 | Return To Vendor

## C100-D3
Status Report for C100-D3 on 10-Jun-20
TRAV_SEQ_NUM | D3CLASSIFICATION | D3OPENEDTIME | CATEGORIZETECHNICIAN | SERIALID | TRAVELERID | REVNO | SEQNO | PAGENO | STEPNO | D3DESCRIPTION | DESCRIPTIONFILES | D3ACTIONS | ACTIONFILES | D3RESOLVE | RESOLVETECHNICIAN | D3_RESOLVED_TIME | RESOLVEFILES
3 | ELNIK_FURNACE | 2010-01-13 | overtonr | WG_F061359_05 | C100-CAV-INSP-WGD | 2 | 11 |  | 19 | Blistering on copper plating
4 | ELNIK_FURNACE | 2010-01-12 | OVERTONR | WG_F061359_01 | C100-CAV-INSP-WGD | R2 | 16 |  | 19 | Black stains on small flange
5 | ELNIK_FURNACE | 2010-01-12 | overtonr | WG_F061359_11 | C100-CAV-INSP-WGD | R2 | 16 |  | 19 | Black stains on small flange
6 | RF_TEST_ABORTED | 2010-02-05 | forehand | C100_2A | C100_2A |  |  |  |  | Unable to find C1 D cavity center frequency at Kelvin in VTA dewar using network analyzer E |  | Abort VTA Kelvin RF test E
7 |  | 2010-06-07 | JDAVEN | R100_01 | C100-CAV-DEGR | R1 | 1 | 1 | 4 | Gouge discovered on one of the HOM flanges. Steve Castagnola inspected and decided cavity needed lapping on flange surface. |  | R100-01 taken back to Isiah Daniels to have HOM flange touched up with lapping. |  | Steve Castagnola inspected post-lapping flange and approved cavity progression to degreasing/UHV. | JDAVEN | 2010-06-07
8 |  | 2010-06-09 | JDAVEN | R100_01 | C100-CAV-DEGR | R1 | 2 | 1 | 3 | Peter Kneisel was using the UHV to degrease one of his seamless 3-cell cavities. R100-01 sat for 30 minutes waiting for UHV to become available. |  | Tony Reilly was notified of this breach in schedule adherence. Seamless 3-cell cavity finished UHV process. Cavity removed and rinsed. Tank emptied and refilled with fresh water for R100-01. |  | Schedule adherence needs to be emphasized on EP days. Necessary for cavity to enter ultrasonic bath immediately after being removed from EP machine. | JDAVEN | 2010-06-09
9 |  | 2010-06-09 | JDAVEN | CRM10270130001_001 | C100-CAV-EPOL | R1 | 1 | 5 | 24 | Peter Kneisel was using the UHV to degrease one of his seamless 3-cell cavities. R100-01 sat for 30 minutes waiting for UHV to become available. |  | Tony Reilly was notified of this breach in schedule adherence. Seamless 3-cell cavity finished UHV process. Cavity removed and rinsed. Tank emptied and refilled with fresh water for R100-01. |  | Schedule adherence needs to be emphasized on EP days. Necessary for cavity to enter ultrasonic bath immediately after being removed from EP machine. | JDAVEN | 2010-06-09
11 | DEVIATIONS | 2010-11-22 | FOREHAND | C100-RI-008 | C100-CAV-ASSY2 | R1 | 14 |  |  | Cavity pressure was one order of magnitude higher than normal when cavity was moved from clean room to macha's landing. |  | No actions taken |  | No resolution necessary | FOREHAND | 2010-11-22
12 | LEAK | 2010-12-09 | MACHA | F067763/06 | C100-CST-ASSY-WGD | R1 | 3 |  |  | Waveguide to tophat serpentine seal leaked. Sized leak to 2.6e-8 std. cc/sec
13 | OTHER | 2011-01-14 | aburrill | 0TEST | C100-D3 | R1 | 1
14 | LEAK | 2011-01-25 | FHUMPHRY | CRM1207001-1000_001 | C100-CM-ASSY-COLD | R3 | 7 |  |  | Upon 2k leak check, found 8e -9 leak in the return header feed thru. serial number 1385-01-34 Location was directly above helium vessel 2. The leak was identified as through the pins. |  | The feed thru was removed from the 2k circuit. |  | Installed a new feed thru at location. 1385-01-41 is now located in the return header above helium vessel 2. | FHUMPHRY | 2011-01-25
15 | OTHER | 2011-03-18 | MACHA | C100-002 | C100-CST-ASSY-ASO1 | R2 | 4 |  |  | Epoxy smell was noticed on Day four of the cavity string assembly around 5:00pm. After investigation it was believed the contractors were mixing epoxy outside and the doors were open allowing air to flow into the building.
16 |  | 2011-04-05 | MACHA | C100-002 | C100-CST-ASSY-ASO1 | R2 | 5 |  |  | Two leaks found during final leak test. Leak at serpentine seal at C100-01 Leak at serpentine seal at C100-17 | 262D413A-BBB6-5BB7-72AD3E2B23A2153D | Went to a higher torque. 43' lbs on the sides and 40' lbs on the ends. Leak size changed some. See leak test chart in the string traveler or attached to this D-3 |  | After group discussion it was decided to allow the string to move forward for cryomodule assembly.
17 | HARDWARE_UNAVAILABLE | 2011-04-05 | FISCHER | R100-01 | R100-CM-ASSY |  |  |  |  | INS VAC BLED UP TO CHECK N-TYPE RF FEEDTHRUS. UPON RECEIVING INSPECTION OF FEEDTHRUS CENTER THREADED PINS WERE FOUND TO BE LOOSE OR DISCONNECTED. |  | INS VAC BLED UP. FEEDTHRUS REMOVED, CENTER PINS CHECKED FOR TIGHTNESS. FEEDTHRUS REINSTALLED. VAC RE-ESTABLISHED. |  | LEAK CHECK PERFORMED,BACK ON TRACK. | FISCHER | 2011-04-05
18 | ELNIK_FURNACE | 2011-06-02 | OVERTONR | L09 | C100-CM-INSP-GVWG | R1 | 18 |  |  | power outage during run furnace was already cooling down |  | no action needed |  |  | OVERTONR | 2011-06-02
19 | ELNIK_FURNACE | 2011-06-02 | OVERTONR | L10 | C100-CM-INSP-GVWG | R1 | 19 | 5 |  | power outage during run furnace was already cooling down |  | no action needed |  |  | OVERTONR | 2011-06-02
20 | ELNIK_FURNACE | 2011-06-02 | OVERTONR | L12 | C100-CM-INSP-GVWG | R1 | 20 | 5 |  | power outage during run furnace was already cooling down |  | no action needed |  |  | OVERTONR | 2011-06-02
21 | ELNIK_FURNACE | 2011-06-02 | OVERTONR | L14 | C100-CM-INSP-GVWG | R1 | 21 | 5 |  | power outage during run furnace already cooling down |  | no action needed |  |  | OVERTONR | 2011-06-02
22 | DEVIATIONS | 2011-06-17 | ABURRILL | C100-003 | C100-CST-ASSY-ASO1 | R2 | 6 |  |  | Removed Cavity C100-23 and replaced with C100-31 due to HPR failure during rinsing of C100-23. See attached email to cognizant staff. |  |  | 9E9BA9FE-0593-95CE-9916836DFAAAB43F
23 | STAINS |  | MACHA | C100-003 | C100-CST-ASSY-ASO1 | R2 | 6 |  |  | Cavity RI-029 has two small stains that cross the seal path. After talking to Andrew we thought they are stains from a Sharpie used during polishing. Attempts to remove with acetone Isopropanol and DI water all failed. Andrew believed this would not be a problem so we continued to assemble. No scratches were observed. |  |  |  | The string was found to be leak tight. The stains did not cause a leak. | MACHA | 2011-06-21
24 | DEVIATIONS | 2011-06-21 | FISCHER | CRM1207001-1000_001 | C100-CM-ASSY-SFR | R1 | 2 |  |  | North end 50 k shield was installed as outlined but was low and not centered in spaceframe. we lifted the south end 50 k shield 1/2" by changing g-10 hanger hole patterns. |  | changed hole pattern in new g-10 hangers. |  | change was necessary to center the shield in the spaceframe assy. | FISCHER | 2011-06-21
25 | DEVIATIONS | 2011-06-21 | FISCHER | CRM1207001-1000_001 | C100-CM-ASSY-SFR | R1 | 2 |  |  | 50 k shield modifications |  | top panels were sectioned into 19" long peices, ease of installation. Bolt pattern for tuner side long axial rod flex straps relocated above window, straps too short. Third modification was added helicoils to the HOM strapping to ease assy. |  | modification listed above to ease assy. all completed as built. Dwg package has been reviewed and rev is being done. | FISCHER | 2011-06-21
26 | DEVIATIONS | 2011-06-21 | FISCHER | CRM1207001-1000_001 | C100-CM-ASSY-SFR | R1 | 2 |  |  | alignment of cavity string presented problems at cavity 4 HOM end of cavity. |  | alignment of cavity 4 was tried several different ways, could not achieve tolerance of +/- .020". |  | Adjoining cavities were adjusted along with cavity 4 to minimize deviation across beamline flanges. Consulted with M Wiseman. Finish torque values of cav 4, 200 in/lbs on east side (tuner), 75 in/lbs on opposing side. | FISCHER | 2011-06-21
27 | DEVIATIONS | 2011-06-21 | FISCHER | CRM1207001-1000_001 | C100-CM-ASSY-SFR | R1 | 2 |  | 11 | outer magnetic shield modification. |  | In center of cryomodule, double layer of mag shield tophat pcs did not fit over each other, G. Cheng consulted. |  | Second shield was slit to fit over inner shield . | FISCHER | 2011-06-21
28 | DETOURS | 2011-06-21 | FISCHER | CRM1207001-1000_001 | C100-CM-ASSY-VV | R1 | 2 |  |  | C100-01 Cryomodule was taken to the tunnel for Microphonic measurements after this traveler. We did not move on to the Final Assy Traveler. |  | Module sent to Accelerator tunnel, cooled down. Measurements taken. Returned to Test Lab for commissioning. |  | Module in test cave, final assembly traveler will be started at this point. | FISCHER | 2011-06-21
29 | OTHER | 2011-06-20 |  | C100-003 | C100-CST-ASSY-ASO1 | R2 | 6 |  |  | During the final day of string assembly the make-up air pre-filters were changed out on the clean room blowers without notice to assembly personnel. It is believed the filters were changed while the last two cavities were being assembled in the class 10, C100-30 and C100-28. It is possible a beam-line was being assembled in the main room on the string.
30 | DEVIATIONS | 2011-06-22 | FISCHER | C100-01 | C100-CM-ASSY | R1 | 2 |  |  | Cryomodule did not follow normal assembly path. |  | Cryomodule was sent to tunnel for micr phonics measurements prior to the Final Assembly Traveler being completed. |  | Module returned for final assembly and CMTF cooldown and testing. | FISCHER | 2011-06-22
31 | DISCREPANCIES | 2011-06-24 |  | HV Bellows see attachment | C100-CM-INSP-HELV | R2 | 12 |  |  | As a result of the inspection of sample Helium Vessel Assemblies, 100 % inspection of all in house bellows was requested. Inspection defects are of visual nature found on bellow convolutions. See attached list of serial numbers inspected and defect photos. Defect photo file: M:\cmm\Pictures\Inspection Pictures\C100\C100HeliumVessel\BELLOW INSPECTION | C1E813E0-0632-E09D-BC3FBEB5549AA1D5
32 | DEVIATIONS | 2011-07-20 | ABURRILL | C100-004 | C100-CST-ASSY-ASO1 | R2 | 7 |  |  | Cavity # C100-RI-035 was dropped from ~8" while being removed from the bakers rack. The cavity landed on the beamline flange on the HOM end of the cavity. |  | As a precautionary measure the following steps are being taken. 1. CMM will re-execute the post He vessel welding inspection and compare it to the data taken previously. 2. Another beadpull on the cavity will be carried out to ensure the file flatness is still within specification. 3. The HOM and Field probe feedthrus will be leak checked when removed from the cavity for degreasing. |  | Cavity tested better than most, ~35 MV/m. | ABURRILL
33 | LEAK | 2011-07-26 | MACHA | C100-004 | C100-CST-ASSY-ASO1 | R2 | 7 |  |  | Three leaks detected during final leak test of the cavity string. After going to a higher torque one leak sealed up. See leak test charts in the traveler for details.
34 | DEVIATIONS | 2011-07-27 | FOREHAND | CRM1207001-1000_003 | C100-CM-ASSY-COLD | R3 | 9 | 2 | 4 | Cavity string had to be re-measured due to the fact that the He circuit was evacuated without cavities locked down properly. |  | The cavity modes were measured and recorded after the He circuit was bled back up. This new list shall be referred to when calculating deltas during later steps. | 359CA1B6-C256-53D0-5AB4737F355EF25E | no resolution necessary. | FOREHAND | 2011-07-27
35 | LEAK | 2011-08-24 | MACHA | C100-002 | C100-CST-ASSY-ASO1 | R2 | 4 |  |  | D-3 was never submitted for HOM feedthrough leak. This D-3 is to capture the history. When this string was pumped down it never pumped lower than 1e-2 due to a broken HOM feedthrough at position #7 HOMA Feedthru # MPF-0028. |  | The string was disassembled and rerinsed to be assembled again. A new feedthrough was used for the leaking broken feeedthru.
36 | OTHER | 2011-09-07 | FHUMPHRY | BC-02138-09 | C100-NCR | R6 | 274 |  |  | The Primary return bayonet is extremely tight going in, and will not seat at the bottom. The male end of the bayonet has major interference with bottom female seat the last 1 1/4" of insertion. | C0ED9219-C021-B6FA-C6AE18105091F3B4 | Sent the end can back to the vender for repair. Vendor to hone interference to provide clearance. |  | Endcan arrived back from vendor. Receiving inspection was repeated, including leak checks and pressure tests. All tests good. Unit can be used for CM assembly. | FISCHER | 2012-05-09
37 | LEAK | 2011-09-09 | CASTAGNO | F067763/07 | C100-CST-ASSY-WGD | R1 | 53 |  |  | Unable to achieve vacuum tight seal between this waveguide and window. |  | Three different windows have been assembled to this waveguide. Each window was assembled twice and leaked twice. The knife edge on the large flange has a flat spot on of the corners and is the source of the leaks. "Update" Three different windows have been assembled to this waveguide. Each window was assembled twice and leaked twice. The knife edge on the large flange has a flat spot on of the corners and is the source of the leaks. "Update" This waveguide has had a history of intermittently sealing and not sealing. It has been issued to production (2) times prior to this current attempt to achieve a leak tight seal. There are (2) waveguide re-work travelers filled out in the past. The area of the knife edge where the leaks have been occurring has been reworked in an attempt to sharpen the knife edge. The height of the knife edge relative to the flange face has been lowered by approx. 0.004" by this re-work. Further flattening has occurred to other adjacent areas of the knife edge, presumably from repeated assemblies. BC 09/09/11 | 1475D840-D558-84BF-C9305ED89BE46B1C | Mark waveguide "do not use."
38 | LEAK | 2011-09-27 | MACHA | C100-005 | C100-CST-ASSY-ASO1 | R2 | 8 |  |  | Two Large leaks detected in C100-5 string. After pumping on the string over the weekend the pressure was at 8.6e-7 on Monday morning. Leak #1 was found on the beam-line to Return VAT valve metal seal gasket. After retorque of the bolts the string pressure dropped to 7.1e-7. A second leak #2 was found at the return VAT valve (472XX-XE01-ACP1/0011) bonnet seal. After tightening the bonnet seal bolts by hand the pressure quickly dropped to 8e-8. After pumping the string overnight the pressure was 5e-8. The leaks were now leak tight as well as the remainder of the string. See chart in the string traveler. Note: The bolts at the beam-line to VAT valve were not tightened properly. Techs are afraid to remove the HOM elbows and this must be done to properly tighten the three bolts at that location. |  |  |  | Use as is
39 | ELNIK_FURNACE | 2011-09-28 | OVERTONR | L35 | C100-CM-INSP-GVWG | R1 | 64 | 5 | 9 | Spontaneous temperature spikes during temperature ramp up | E3100541-0DB5-1189-5304B14EF49ECEDD | checked resistivity after run improved from 5.0 to 7.1 |  |  | OVERTONR | 2011-09-28
40 | ELNIK_FURNACE | 2011-09-28 | OVERTONR | R45 | C100-CM-INSP-GVWG | R1 | 61 | 5 | 9 | Spontaneous temperature spikes during temperature ramp up | E34B59A2-B121-F3E0-53A955218CA27D69 | checked resistivity after run improved from 5.0 to 7.1 |  |  |  | 2011-09-28
41 | ELNIK_FURNACE | 2011-09-28 | OVERTONR | R42 | C100-CM-INSP-GVWG | R1 | 62 | 5 | 9 | Spontaneous temperature spikes during temperature ramp up | E3771196-B2A2-94E3-B174F6872CA35DCE | checked resistivity after run improved from 5.0 to 7.1 |  |  | OVERTONR | 2011-09-28
42 | ELNIK_FURNACE | 2011-09-28 | OVERTONR | L31 | C100-CM-INSP-GVWG | R1 | 65 | 5 | 9 | Spontaneous temperature spikes during temperature ramp up | E3A2F785-BEDB-F661-ED189B503C7048CE | checked resistivity after run improved from 5.0 to 7.1
43 | ELECTRICAL | 2011-09-23 | KING | CRM1207001-1000_002 | C100-CM-ASSYR1 | R1 | 3 | 2 | 6 | Unable to power test resistor of IR sensor (s/n 016) in position #4A on C100-2 while performing the high power checklist, prior to CMTF HPRF testing. |  | To allow high power testing, we warmed the outside of the waveguide sweep to test proper response of the IR sensor, in place of powering up the test resistor. On 23-SEP-2011, after completion of high power testing and CM warmup, the bad sensor was replaced with sensor s/n 046 by Dave Bigelow. |  | Following its removal from the CM, inspection of the shorted IR test resistor revealed that the ceramic spacer for the 6 legs of the feedthrough/IR body was too far down (away from the sensor). This made it possible for the feedthrough's legs to be accidentally bent outward. The leg for pin D (R test +) was in contact with the vacuum feedthrough body, causing the test resistor to appear shorted to ground when powered. All vacuum IR assemblies in inventory will be inspected to ensure the ceramic spacer is properly positioned. Both IR-A and IR-B sensors should be electrically checked and have their self-test resistors powered up, prior to cooldown, if possible. | KING | 2011-09-26 | 29974A63-E8B3-CA46-B54ACAC6A1FF016E
44 | LEAK | 2011-10-11 | MACHA | C100-004_2 | C100-CST-ASSY-ASO1 | R2 | 9 |  |  | One leak found in the serpentine seal gasket at C100-23. Retorqued to higher torque, 43 Ft lbs on the sides and 40 Ft lbs on the ends. The leak size is 4.5e-7 TL/s. Retorque did not change the leak size.
45 | DISCREPANCIES | 2011-10-13 | FISCHER | C100-05 | C100-CM-ASSY-COLD | R3 | 11 |  |  | When attempting to measure HOM and connector saver profile with TDR,HOM feedthurs and connectors were observed to be wet from cavity assembly. |  | email sent to parties describing discrepancies, met on site to discuss. Connectors were removed dried and re-installed. |  | Assembly of Cryomodule has resumed. | FISCHER | 2011-10-13
46 | DISCREPANCIES | 2011-10-13 | FISCHER | CRM1207001-1000_003 | C100-CM-ASSY-VV | R1 | 4 |  |  | During TDR measurements it was noticed that the HOM connector savers were missing on all cavities. Attached pictures show with and without savers. |  | When discovered Cryomodule was already closed up, HOM feedthrus were left without savers in place. RF cable directly to HOM feedthru. Email sent, parties met on site. Leave as is. | EB71FD2B-F8D1-3DDB-E5CC4F131005C5B6 | Leave as is. Step added to assembly traveler to verify connector saver is in place prior to close up. | FISCHER | 2011-10-13
47 | SCRATCH | 2011-10-17 | GEORGED | F067763-10 | C100-CST-ASSY-WGD | R1 | 54 |  |  | In the cleanroom a scratch was discovered in the seal path of small flange. Unknown if damage occurred in the chemroom, pass-thru, or cleanroom.
48 | DEVIATIONS | 2011-11-02 | FISCHER | CRM1207001-1000_003 | C100-CM-ASSY-COLD | R3 | 9 |  |  | SEE D3-34 FOR DANNY FOREHAND's COMMENTS. CAVITY STRING WAS EVACUATED BUT NOT PROPERLY SECURED PRIOR. CAUSING FREQ IN ALL CAVITIES TO DROP. |  | USED THE COLD TUNERS TO TUNE THE CAVITIES BACK TO THE NOMINAL 1494.05-1494.1 MHZ. INSTALLED TUNER LOCKS. | F65B0002-A3FD-804B-328B4B1EED01B4BC | ALSO RECORDING THE WARM FREQS TAKEN IN CMTF, INS VAC ESTABLISHED, HELIUM CIRCUIT RETURNING GAS TO CTF.CAVITIES UNDER VAC AS WELL. | FISCHER | 2011-11-02 | F6A0D29E-B505-1519-B61C7C5024F1B00F
49 | LEAK | 2011-12-02 | MACHA | C100-007 | C100-CST-ASSY-ASO1 | R2 | 11 |  |  | A leak was detected in the HOM metal seal at C100-50 HOMa #RI-PRO-036. After repeated retorque at 45" lbs the leak sealed and is now leak tight. Two other leaks were detected in serpentine seals at C100-54 and C100-52. Leak sizes before increasing torque. C100-54 1.99e-8 TL/s and C100-52 1.34e-8 TL/s. Increase torque to 43' lbs on the sides and 40' lbs on the ends. Leak sizes after higher torque. C100-54 no change 1.99e-8 TL/s and C100-52 1.28e-8 TL/s |  | Went to higher torque.
50 | DEVIATIONS | 2011-12-14 | WILCOX | C100-03 | C100-CM-HPRF-ACTS | R1 | 4 |  |  | TUNER 047D.084; UPPER COMPRESS SPLIT SHAFT ASSEMBLY ON C100-3 CAVITY 4 A GOT BENT DURING CRYOMODULE CAVE TESTING .055" OUT OF SPEC. PART GOT BENT BY THE LIMIT SWITCH BLOCK BINDING AGAINST SECOND WARM TUNER PLATE AND DRIVEN FROM CMTF WITH THE STEPPER MOTOR UNTIL ARTICLE WAS BENT. NO OTHER ILL EFFECTS WERE NOTICED. |  | STRAIGHTEN OUT PART AND RE-LEAK CHECK. IF PART PASSES LEAK CHECK. IT IS OK FOR USE AS PER MIKE MCCREA NOTE: THESE PARTS ARRIVED FROM VENDOR .020"-.040" OUT OF SPEC. WE HAVE CHECKED 16 OF THESE AFTER REPAIR AND HAVE NOT HAD A PROBLEM YET | 6E03ED7B-FFA2-FF9E-F7ECA2EBC2FAB508 | 047D.084 PART NUMBER. PART WAS RE STRAIGHTENED THEN LEAK CHECKED AS DONE PREVIOUSLY. PART CAN BE USED AS IS FOR FUTURE CM ASSEMBLY. | FISCHER | 2012-05-09
51 | DEVIATIONS | 2011-12-14 | WILCOX | C100-03 TUNER | [5 empty] | UPPER COMPRESS SPLIT SHAFT ASSEMBLY ON C100-3 CAVITY 3 GOT BENT DURING CRYOMODULE CAVE TESTING .037" OUT OF SPEC. | 6E202053-B537-DD00-B8D7D81C41A66294 | STRAIGHTEN PART OUT, LEAK CHECK AND USE IF PART PASSES LEAK CHECK. NOTE: ALL OF THESE PIECES SO FAR HAVE ARRIVED FROM VENDOR BENT .020"-.040" HAVE NOT FOUND A LEAKY ONE YET AFTER 1ST REPAIR |  | Tuner piece was removed from C100-03 cavity 3 warm tuner. Was re-straightened and weld leak checked, still leak tight. Unit can be used as is, back within tolerance. | FISCHER | 2012-05-09
52 | LEAK | 2011-12-19 | WILCOX | CRM1207002-0001_016 | C100-CM-INSP-TPHTS | R1 | 21 | 2 | 9 | Top hat had several repeatable leaks after being welded in house. |  | Sent unit back to weld area for repair and then re leak test |  | unit passed leak check after reweld on outer ring | WILCOX | 2011-12-19 | 6EB3DAD6-0A5A-1B01-D979D6735AEA601B
53 | LEAK | 2011-12-19 | MACHA | C100-008 | C100-CST-ASSY-ASO1 | R2 | 12 |  |  | Leak detected in the Beam-line on C100-79 connection to the supply VAT valve. Leak size at 40'lbs was 7e-6 TL/s. After torquing to 45'lbs the leak improved to 4.05e-6 TL/s.
54 | DEVIATIONS | 2012-01-09 | FISCHER | C100-RI-037 | C100-CM-ASSY-COLD | R4 | 13 | 4 | 14 | When tuners were assembled 3.818 block wasn't used. This block sets the dimension between the jack blocks, holding correct angles at Ti flexures and assuring the warm tuner ball screw is at midrange when cavity is tuned to 1497. |  | Measurements were taken, both frequency and inner to outer tuner shafts, matalevich spreadsheet gave a dimension for a spacer, which was installed to correct condition. When spacer was added new hardware was used to inner shaft, threaded rod with SiBz nuts. Modulas done on Ti flexure looking at load increase, deemed to be ok. see attached. | A6FB5A47-CBAB-F0C6-7238B2FF73B74A2B | Spacers calculated and added to assembly. Tuners are repaired. Spacers not added till vv assy trav, with warm tuner section. | FISCHER | 2012-02-23
55 | DEVIATIONS | 2012-01-09 | FISCHER | C100-05 | C100-CM-ASSY-COLD | R3 | 11 | 3 | 14 | When tuners were assembled 3.818 block was not used. This block sets the distance between the jack blocks on the tuner scissors. This dimension puts the Ti flexures and warm tuners in the correct location. |  | Measurements were taken, both freq and from inner to outer shafts, numbers passed on to Matalevich, spreadsheet attached showing calcs determining spacer length which was added to the inner shaft to correct deviation. New hardware was used, threaded rod and SiBz nuts. Modulas was run on Ti flexures to verify extra angle/ force would not compromise them. | 9CE8AB45-0498-2ECA-1A75ADF34424792D | Spacers were added to inner tuner shaft during warm tuner assembly. Corrected. | FISCHER | 2012-03-05
56 | LEAK | 2012-01-23 | MACHA | C100-009 | C100-CST-ASSY-ASO1 | R2 | 13 |  |  | Leak detected in the serpentine gasket at C100-80. The leak size before higher torque is 4.4e-9 TL/s. |  | Torqued serpentine flange to a higher torque value (45'lbs on the sides & 38' lbs on the ends). Re-sized the leak after higher torque 3.62e-9 TL/s. |  | Use as is
57 | DEVIATIONS | 2012-02-03 |  | CRM1207001-1000_004 | C100-CM-ASSY-SFR | R1 | 6 | 2 | 8 | ONCE SF ASSY WAS LOADED INTO VV TUNER FIT WAS VERIFIED, STRING NEEDED TO ROLL. UNABLE TO ROLL DUE TO WHEELS ASSY HAS STRING TOO HIGH, UNABLE TO CENTER IN VV. DECISION WAS MADE TO REMOVE STRING, LOCKDOWNS WERE REMOVED IN THE WRONG ORDER. | A342DAED-AD62-E49A-E2DB9E77EDB1ED0E | Mu metal was removed, mli opened,SEC end wheel assy's were slotted and re installed, alignment was repeated. Alignment crew to re- fiducialize 2-6-2012. | A347109A-98EC-C969-E3D52E60B339AE01 | Spaceframe assy completed. Will be re installed into VV. Assy to continue. No ill effects from work performed. | FISCHER | 2012-02-03
58 | LEAK | 2012-02-13 | MACHA | C100-010 | C100-CST-ASSY-ASO1 | R2 | 14 |  |  | Leak detected at C100-62 serpentine seal. Leak size at production torque: 1.3e-6 TL/s After going to a higher torque, 43' lbs on the sides and 40' lbs on the ends, the leak size is: 8.99e-7 TL/s. |  | Went to a higher torque and re-sized leak. |  | Use as is
59 | OTHER | 2012-02-13 | FHUMPHRY | CRM1207001-1000_004 | C100-CM-ASSY-SFR | R1 | 6 |  |  | Spaceframe sitting high in vacuum vessel. The supply speaceframe wheel assembly needed to be slotted to allow the spaceframe to sit lower inside the vacuum vessel. The wheel assembly was also tack welded in place after the mounting bolts were torque to spec. |  | Slotted all supply wheel assembly bolt hole 1/4". |  | slotting the holes corrected the problem. | FHUMPHRY | 2012-02-13
60 | OTHER | 2012-02-22 | FHUMPHRY | CRM1207001-1000_004 | C100-CM-ASSY-VV | R1 | 6 |  |  | Return end can, beampipe vat valve orientation 180 degree out. |  | During the n2 bleed up of the beampipe assembly, prior to initial pump down the valve orientation issue was corrected. |  | Issue resolved. Final leak check good. Bigelow handled issue. | BIGELOW | 2012-02-21
61 | OTHER | 2012-02-22 | FHUMPHRY | CRM1207001-1000_004 | C100-CM-ASSY-SFR | R1 | 6 |  |  | Slight issue installing space frame in vacuum vessel. We thought one of the spaceframe lock downs could of compromised the space frame or cavity string .oo5 tolerance during the positioning of space frame in vacuum vessel. |  | Removed space frame assembly from vacuum vessel and had the alignment guys re-shoot the space frame assembly. |  | Problem resolved. know significant alignment changes occurred. | FHUMPHRY | 2012-02-22
62 | DEVIATIONS | 2012-02-23 | FISCHER | CRM1207001-1000_002 | C100-CM-ASSY-COLD | R3 | 8 |  |  | When CM assembly was completed warm tuner vacuum load springs (green)will not engage with the adjustable spring boss. Distance between flanges too great. Believe 3.818 block was not used, possible solutions -add spacers to inner tuner shaft or lengthen spring bosses |  | No actions taken yet, on the repair list. |  | Module was installed in the tunnel, used as is. Suggest time and resources should be allocated to properly repair tuner assemblies. | FISCHER | 2012-03-28
63 | DEVIATIONS | 2012-02-23 | FISCHER | CRM1207001-1000_002 | C100-CM-ASSY | R1 | 3 | 3 | 8 | Spacers added to inner tuner shaft. SEE COLD MASS D3 for further detail. |  | SEE COLD MASS D3 for further detail. |  | FYI | FISCHER | 2012-02-23
64 | DISCREPANCIES | 2012-02-29 | KING | CRM1207001-1000_005 | C100-CM-CRYO-ACTS | R1 | 6 | 4 | 10 | During initial 4K fill, C100-5 Supply LL readings were erratic, around 60%, and not correlating with the Return LL readings of 28%. Verified both Supply and Return LL cable and instrument are working properly with test resistor/LL simulator. Measured R values at J1-10 Supply LL feedthrough on CM and found pin I shorted to either pin H or J inside the CM. pins G-H: 5 Ohms (should be ~172 Ohms with LL at 28%) pins I-J: short (should be ~167 Ohms with LL at 28%) |  | Warm up cryomodule to 295K, bleed insulating vacuum up to atmosphere, remove and inspect J1-10 (10-pin) feedthrough for possible shorted wiring. Rework shorted components. Reinstall feedthrough and check continuity of sensor. |  | Resistance readings returned to normal during warmup. Subsequent inspection of wiring directly beneath top hat by S. Dutton showed no discrepancies. | KING | 2012-04-16
65 | DEVIATIONS | 2012-03-05 | FISCHER | CRM1207001-1000_004 | C100-CM-ASSY-VV | R1 | 6 | 2 | 9 | When installing REC beampipe it was noticed that the warm VAT valve was installed backwards during initial assembly.Before use, beampipes are leak checked, baked out on a clean turbo system, then releak checked. |  | Beampipe was bled up and valve installed correctly. Beampipe is unable to be baked once installed on Cryomodule. This is an abnormal condition. Just an FYI. All work was done clean, filtered N2, in flow hoods. |  | Problem resolved on the floor. | FISCHER | 2012-03-05
66 | DEVIATIONS | 2012-03-05 | FISCHER | CRM1207001-1000_004 | C100-CM-ASSY-VV | R1 | 6 | 3 | 14 | When tuners were assembled 3.818 block was not used. This block sets the distance between the jack blocks on the tuner scissors. This dimension puts the Ti flexures and warm tuners in the correct location. |  | Measurements were taken, both freq and from inner to outer shafts, numbers passed on to Matalevich, spreadsheet attached showing calcs determining spacer length which was added to the inner shaft to correct deviation. New hardware was used, threaded rod and SiBz nuts. Modulas was run on Ti flexures to verify extra angle/ force would not compromise them. | DFA14307-0E93-AA96-021CD18F0A1A7178 | Spacers were added to inner tuner shaft during warm tuner assembly. Corrected. | FISCHER | 2012-03-05
67 | DISCREPANCIES | 2012-03-05 | FISCHER | CRM1207001-1000_005 | C100-CM-ASSY | R1 | 5 | 2 | 6 | During secondary waveguide installation a leak in the Cu gasket from waveguide to inner RF window was found during leak check. |  | Retorqued Cu gasket to 130 in/lbs. Leak still there. Waveguide was removed and gasket replaced. During this activity the beamline pressure rose and tripped off ionpump. All work done in a clean portable flow hood arrangement. |  | Gasket replaced, joint now leak tight. See beamline D3 for actions taken. | FISCHER | 2012-03-05
68 | DISCREPANCIES | 2012-03-05 | FISCHER | CRM1207001-1000_005 | C100-CM-ASSY | R1 | 5 | 2 | 6 | During the secondary waveguide install on cavity 5, beamline RF window developed a leak. Leak tripped off ionpump. Leak was logged in the CMTF. |  | Installed portable flow hood over SECpumpdrop. Hooked up beamline quality turbo station. Joint was leak checked. Valves opened, evacuated, ionpump re established. |  | No apparent damage done. CM testing in CMTF showed no sure signs of degredation. | FISCHER | 2012-03-05
69 | LEAK | 2012-03-20 | FHUMPHRY | CRM1207001-1000_004 | C100-CM-ASSY-VV | R1 | 6 |  |  | During final pressure test on 50k shield circuit. A -8 leak was identified in the supply end can flex hose. The leak was localized to the bottom side of the flex line. The circuit was at 100 psi when sizing leak. Sub-atmospheric the leak was sized at -9 range. | 4C69F5AC-D7BD-A747-620F89FA18B1979A | Removed 50k flex line from circuit. Went out to bluecrab storage, cut off the 50k flex line from supply end can #3. Supply end can 3 has gone through all leak checks/ pressure tests as well as cold shocks. Alignment survey data attached. The alignment guys surveyed before and after the repair. | 4C8FD214-096E-DFF7-25FA0E67D8C49800 | System leak tight after installing new flex line in place. Pressurized shield circuit with helium 130 psig while monitoring insulating vacuum with inficon leak detector. Pressure test attached | FHUMPHRY | 2012-03-27 | 539FA8F5-B161-17B3-9B681361F9EE14FC
70 | OTHER | 2012-04-05 | BOOKWALT | TESTING | [5 empty] | testing | 329755BF-025C-6B36-4E533D793C7D5655
71 | DISCREPANCIES | 2012-04-10 | JJCAMP | CRM1207001-1000_006 | C100-CM-ASSY-COLD | R4 | 14 |  |  | cavity # 1 has a slight roll causing the liquid level to be too close to the hom connectors, not allowing copper straps to be installed |  | liquid level is cut off to allow copper straps to be installed,then will be rewelded then leak checked again | 596CEF08-EF03-2696-6D4198B56D02DAE8 | All welding complete, circuit re leak checked. Proceed with assembly of Cryomodule. | FISCHER | 2012-04-13 | B599A1D4-DF21-F387-E4FC8CB53EB0EC21
72 | DEVIATIONS | 2012-04-17 | DRURY | CRM1207001-1000_004 | C100-CM-HPRF-ACTS | R1 | 6 |  |  | This cryomodule will undergo helium processing. This is procedure is an attempt to address the high level of field emitted radiation and low gradient field emission onset which was most likely the result of beam line vacuum contamination. |  |  | 8D06F9C4-080C-F4D1-E1036B8FF45DE5F7
73 | OTHER | 2012-04-20 | JJCAMP | CRM1207001-1000_007 | C100-CM-ASSY-VV | R1 | 7 |  |  | during top hat 1 and 3 installation,beam line vacuum went from 9e-9 to 3e-8. |  | stopped installation,watched beam line to see if it recovered. |  | beam line started to recover and now is 9e-9 | JJCAMP | 2012-04-20
74 | DEVIATIONS | 2012-05-09 | FISCHER | CRM1207001-1000_004 | C100-CM-ASSY-VV | R1 |  | 5 | 67 | ONCE SEC WAS IN POSITION THE TOP HEADER HELIUM LINE DID NOT MATCH UP WITH THE END OF THE CAVITY STRING. |  | CUT THE END CAN FLEX LINE IN ORDER TO ROTATE ARTICLE TO FIT UP WITH STRING HEADER. WORK PERFORMED BY QUALIFIED INDIVIDUALS. |  | LINE WAS WELDED IN PLACE BY QUALIFIED WELDER, INSPECTED, LEAK CHECKED AND PRESSURE TESTED. ALL TEST RESULTS WERE GOOD. THIS JOINT IS NO DIFFERENT THEN THE ENTIRE PROCESS PIPING ASSY THAT IS INSTALLED BY SAME PERSONS. | FISCHER | 2012-05-09
75 | LEAK | 2012-05-23 | FHUMPHRY | CRM1207001-1000_004 | C100-CM-CRYO-ACTS | R1 | 7 |  |  | Downstream return heater feedthru developed at 10-6 leak during the 25k cryocycle. |  | Replaced 8 pin feedthru after the module was moved to the high bay area. All attached leak tests were preformed 5-28-12 | 02F00CBC-A41C-45ED-F4AB454F835601C6 | yes, installed new feed thu. All final pressure tests completed. | FHUMPHRY | 2012-05-29
76 | OTHER | 2012-05-23 | FHUMPHRY | CRM1207001-1000_004 | C100-CM-CRYO-ACTS | R1 | 7 |  |  | JT stem seat not sealing properly during Q's. Primary supply u-tube diode never able to warm up like it should of while valve was closed. |  |  |  | Replaced entire JT stem assembly with a new one. | FHUMPHRY | 2012-05-23

## C100-CAV-D3
Status Report for C100-CAV-D3 on 10-Jun-20
TRAV_SEQ_NUM | D3CLASSIFICATION | D3OPENEDTIME | CATEGORIZETECHNICIAN | SERIALID | TRAVELERID | REVNO | SEQNO | PAGENO | STEPNO | D3DESCRIPTION | DESCRIPTIONFILES | D3ACTIONS | ACTIONFILES | D3RESOLVE | RESOLVETECHNICIAN | D3_RESOLVED_TIME | RESOLVEFILES
2 | EP | 2010-09-21 | jdaven | C100_RI_002 | C100-CAV-EPOL | R1 | 2 | 2 | 8 | COMEDY OF ERRORS: 1. Servo Motor wouldn't engage. 2. Brand new Nitrogen flow meter installed without proper notification of scale and function to technician (went from CFM to CFH). |  | 1. Shut main power to cabinet, removed the protective bagging, disconnected power supply spraying with electric port cleaner then re-attached. Manually rotated cage. 2. Multiple restarts to discover to proper beginning flow for process step 5. Appears as long as the flow is atleast "7" (adjusted up to "10" for this run), the program continues to work properly. |  | 1. SERIOUS ATTENTION needs to be paid to Servo issue. It was NEVER properly remedied. 2. Further investigation is needed into the desired nitrogen flow rate by Rongli (transition from CFM to CFH). An overall total 1.5 LOST HOURS of efficient processing due to both issues. | jdaven
3 | EP | 2010-09-28 | jdaven | C100_RI_003 | C100-CAV-EPOL | R1 | 3 | 5 | 23 | The FPC flange was not lapped prior to the EP cycle. Upper management decided to undergo this step after the EP. |  | Stains resulting from the EP cycle weren't buffed off because the cavity is going to be lapped at a later date. |  | Process Change | jdaven | 2010-09-28
4 | HPR | 2010-09-28 | jdaven | C100_RI_003 | C100-CAV-EPOL | R1 | 4 | 5 | 25 | C100-RI-003 did not progress directly to the HPR step because upper management changed to process flow. |  | Cavity wasn't HPR immediately after the Post-EP degrease. |  | Once Lapping FPC, Tuning, and CMM data are collected cavity will be returned to Chemistry for another degrease and the HPR. | jdaven | 2010-09-28
5 | EP | 2010-10-04 | jdaven | C100_RI_005 | C100-CAV-EPOL | R1 | 6 | 3 | 12 | The servo motor ceased to work in the minutes leading up to current engage step of the EP process after working properly up to that point. |  | The process was stopped in order to remedy the issue. No diagnosis has been made in the 8 months this issue has been a constant problem with the EP process. The short term remedy has proven to be one or all of the following: 1. cutting power to the cabinet manually rotating the cage/servo motor 2. Unplugging the power cords to the servo motor and violently shaking them and/or banging them against the wall. |  | It is inefficient to continue plowing through the schedule with "crossed fingers" in hopes of the servo not acting up each and every time. A definitive diagnosis needs to be attained before any more EPs are executed. | jdaven | 2010-10-04
6 | SCRATCH | 2010-10-15 | castagno | C100_RI_004 | [5 empty] | After FPC lapping was completed, all flanges were inspected by Steve C. A series of small scratches were observed on the HOM beam line flange at approximately the 3 o'clock position. | B00CD2DA-C063-ECDD-25A44D3325E0141C | Proper tooling is not available to adequately repair the flange sealing surface. The seal surface was wet polished with maroon colored Scotch-brite in an attempt to increase the chances of a leak tight cavity after assembly. |  | Cavity was leak tight at room temperature. Proceed the cavity to the VTA. 120C bake went okay. | castagno | 2010-10-15
7 | LEAK | 2010-10-14 | kdavis | C100_RI_004 | C100-CAV-VTRF | R2 | 6 | 1 | 7 | Cavity vacuum deteriorated when pumping dewar down through lambda |  | AI to investigate possible leak sources |  | Changed lapping procedure, AlMg seal under investigation. | aburrill | 2010-12-16
8 | LEAK | 2010-10-14 | kdavis | C100_RI_004 | C100-CAV-VTRF | R2 | 6 | 2 | 15 | Unable to perform high power HOM measurements due to cavity helium leak cutting short this test. |  | Removed cavity from dewar and leak tested warm in VTA mezz. No leaks detected. Bled up cavity and performed inspection of aluminum gaskets. Determine gasket crush. (see attachment) | F9457298-BFE1-1874-37C26AA4D3882AD0 | Use of new polishing cloth could be cause of cold leak.
9 | LEAK | 2010-10-21 | castagno | C100_RI_008 | C100-CAV-ASSY2 | R1 | 10 |  |  | During the slow pump down the pressure bottomed out at 52 mbar after 5 hours of pumping. Opening the throttle valve did not lower the pressure. A large leak is suspected. |  | "Whistling" could be heard coming from the field probe feedthru area. IPA was sprayed around the HOM and FP flanges to better determine where the leak was. The FP gasket is leaking. |  | Do not use old style FP flanges in future. | aburrill | 2010-12-16
10 | EP | 2010-10-14 | stipton | C100_RI_009 | C100-CAV-EPOL | R1 | 10 | 4 | 17 | Cavity etch rate seemed to move a great deal faster than normal. |  | Andrew Burrill instructed Sarah Tipton to extend polishing until the algorithm read 35um instead of the normal projected 30um. |  | Post EP thickness measurements revealed the material removed was an average of 34um. | stipton | 2010-10-15
11 | EP | 2010-10-20 | jdaven | C100_RI_011 | C100-CAV-EPOL | R1 | 12 | 4 | 17 | Andrew Burrill instructed James Davenport to cease the polishing process when 25um were removed in reaction to the majority of the C100 EPs resulting in 40-45um being removed when 30um was shown by the algorithm. |  |  |  | Post EP thickness measurements revealed an average of only 25um was removed. | tmharris | 2010-10-21
12 | LEAK | 2010-10-27 | dreyfuss | C100_RI_006 | [5 empty] | Cold leak in dewar. |  | leak tested warm no leaks after a twenty minute helium spray in a bag Cavity was then disassembled and all gaskets were measured for crush.(See attachment) | EE387C02-F3D6-EDCD-2AFB67F2C0004F5A | Use of new polishing cloth could be cause of cold leak. Relap FPC flange using original lapping paper/procedure. degrease/HPR/ assemble cavity/ retest in VTA/watch for cold leak. | castagno | 2010-10-29
13 | OTHER | 2010-10-28 |  | C100_RI_008 | C100-CAV-ASSY2 | R1 | 11 |  |  | Black liquid emitted from scroll pump after block valve was opened. An unusual odor was detected as well. |  | After the black liquid dried, a sample was taken and given to Andy Wu for analysis. Record analog scans of the system with and without the cavity valve open and compare the readings. Disassemble the cavity and perform an interior visual inspection. | 12D04207-C48D-2CFC-D36FD5B47A580BCF | Nothing unusual indicated by surface analysis,RGA analog scan, or visual inspection. Relap flange using old lapping method and retest in VTA. | castagno | 2010-11-03
14 | DISCREPANCIES | 2010-10-28 |  | C100_RI_008 | C100-CAV-ASSY2 | R1 | 11 |  |  | Total pressure was at 1.2e-6 mbar the next day. This is one order of magnitude lower than normal. |  |  |  | Cavity had damage to inside of cavity
15 | HPR | 2010-10-22 | tmharris | C100_RI_007 | C100-CAV-HPR | R1 | 21 | 2 | 5 | HPR wand did not travel the complete length of program for the second pass. It stopped in the upper area of the cavity. | F87E502D-08DA-4E44-5AC9120E7CBEE09B | Phil Denny, AJ Jacobson, Derrick Dail, and Brian Carpenter have been commissioned to appraise the situation and remedy any and all problems. |  | HPR stepper motor and lead screw replaced | aburrill | 2010-10-22
16 | DISCREPANCIES | 2010-10-30 | STIPTON | F067763-20 | C100-CAV-INSP-WGD | R4 | 52 | 4 | 9 | the Cu plating on the interior of the waveguide is flaking off |  | Finished drying and bagged it. Currently sitting on the shelf in the Chemroom awaiting path forward |  | re-inspect in CMM | ABURRILL | 2012-01-20
17 | OTHER | 2010-11-06 |  | C100_RI_006 | [5 empty] | While preparing to remove cavity from dewar, technician noted that cavity isolation valve on the toplate was already closed. Not known if cavity was being actively pumped while high power test took place. |  | Opened valve and check rise in overall pressure. Pressure did not get above 5E-6 mbar. Appears that cavity is leak tight and test was not affected- even if valve was, in fact, closed during high power test. |  | No resolution necessary. |  | 2010-11-07
18 | OTHER | 2010-11-07 | forehand | C100_RI_007 | [5 empty] | VTA testing tophat was found to have a short at the input probe tip while measuring HOM filters. Occurred on 10/21/10 |  | Cavity was sent back to clean room for re-assembly and rinsing. |  | No resolution-----cavity had problems during HPR after dis-assembly. Cavity will be re-assembled after resolving second problem |  | 2010-11-07
19 | EP | 2010-11-02 | jdaven | C100_RI_007 | C100-CAV-EPOL | R1 | 14 | 2 | 9 | Pump rate exceedingly high off acid flow. Andrew Burrill and AJ Jacobson inspected pump 29Oct2010 but failed to discover any issues. |  |  |  | Pump scheduled to be inspected again on 04Nov2010. All fittings on pump tightened, new rate may be correct rate for pump with no leaks. | jdaven | 2010-11-05
20 | EP | 2010-11-02 | jdaven | C100_RI_007 | C100-CAV-EPOL | R1 | 14 | 3 | 12 | Weird spike in the voltage reading prior to the current being engaged on LabView (see screen shot attached to traveler). |  |  |  |  | jdaven | 2010-11-02
21 | DISCREPANCIES | 2010-11-08 | CASTAGNO | C100_RI_009 | C100-CAV-LAP | R1 | 2 |  |  | Both HOM flange seal grooves exhibit a chatter pattern from machining or similar feature. Flange B also has what appears to be a pit or material defect in the seal groove @ the 6 or 7 o'clock position when the FPC is up. |  |  |  | Cavity is being held for repair with polishing stone. | ABURRILL | 2011-07-14
22 | BCP | 2010-11-12 | overtonr | F067764/02 | C100-CAV-INSP-WGD | R4 | 57 | 5 |  | This wave guide was baked a second time on Nov. 12 |  | no action needed |  |  | overtonr | 2010-11-12
23 | DISCREPANCIES | 2010-11-12 | forehand | C100_RI_006 | [5 empty] | Technician noticed a gouge on the exterior of the beamline on the FPC end of the cavity as it was being removed from cage. There is a dent inside the beamtube that correlates with this gouge. |  | Informed the supervisor and group leader. We discussed theories as to how the damage occurred. |  | Changes will be made to the cavity caging procedure and some of the tooling will be modified to have softer edges. These two actions should resolve this from happening in the future. | forehand | 2010-11-12
24 | BCP | 2010-11-18 | jleung | clean room epoxy odor | [5 empty] | D3 as initiated by Johnny Leung. On 11/15/10, a strong odor was noted by the QC Lab around 4:30 pm. Mortenson Construction had completed a batch mix of epoxy outside of QC lab. Traveling odor may have reached into the air intake of the clean room. |  | QC Lab notified Tony Reilly and Jennifer Williams (ESH&Q) of incident. |  | Mortenson was informed not to mix epoxy inside the test lab. | jleung | 2010-11-18
25 | DEVIATIONS | 2010-12-01 | forehand | C100_RI_007 | C100-CAV-TUNE | R2 | 17 |  |  | Cavity had issues during the original attempt at VTA testing. It was returned to the EP machined and polished a second time- therefore it came back to the tuning are to be re-tuned for VTA testing. The attached file shows the field flatness after the second polishing. | A36DD307-09CA-9B2F-1E169431E009102B | The cavity was re-tuned a second time before VTA testing. The attached file is the field flatness after the second tuning. The cavity is now ready for its second attempt at VTA testing | A36F5311-BB68-1870-FCB13F32B61949DF | Cavity tuning for VTA testing complete. | forehand | 2010-12-01
26 | BCP | 2010-12-02 | castagno | C100-RI-012 | [5 empty] | Inside of FPC waveguide is stained or discolored. Presumably from EP process. |  | Continue to monitor future cavities, ensure degreasing uses adequate liquinox to clean the waveguide
27 | STAINS | 2010-12-02 | castagno | C100-RI-012 | [5 empty] | Inside of FPC waveguide is stained or discolored. Presumably from EP process. | A934F6AE-9404-35D9-CB1B7D3FF7F3DEA0
28 | WELDING | 2010-12-03 | OGLE | C100_RI_007 | C100-CAV-INSP2 | R3 | 12 |  |  | OBSERVED SURFACE DEFECT ON HOM FLANGES;OXIDATION IN SEAL SURFACE. | AE66312F-C57A-422D-B818605ABF34DE5A |  |  | polish and use as is | ABURRILL | 2012-01-20
29 | RF_TEST_ABORTED | 2010-12-04 | grenoble | C100_RI_005 | C100-CAV-VTRF | R2 | 13 | 1 | 10 | Field probe cable short near top plate connector. |  | Remove insert from dewar and replace cable. |  | replacement cables being investigated | aburrill | 2010-12-16
30 | BCP | 2010-12-14 |  | C100_RI_011 | C100-CAV-VTRF | R2 | 23 | 2 | 15 | HOM A dewar cable open just below top plate. QextA measurement is high due to open. |  | Ensure all cables are TDR'd prior to insertion into the dewar for testing.
31 | HPR | 2010-12-15 | tmharris | C100_RI_003 | C100-CAV-HPR | R1 | 52 | 1 | 3 | When I removed covers for HPR I smelled an order to me it smelled like something burnt |  | Called Steve and by the time he came in there was no longer a smell I was told to go ahead with HPR |  |  | tmharris | 2010-12-15
32 | HPR | 2010-12-16 | ABURRILL | C100_RI_012 | C100-CAV-HPR | R1 | 50 |  |  | BCP tool turned off so there is no temperature reading for the DI system |  | Install thermocouple in HPR cabinet |  |  | ABURRILL
33 | HPR | 2011-01-06 | TMHARRIS | C100_RI_008 | C100-CAV-HPR | R1 | 55 | 2 | 4 | Tony Reilly verified that HPR should continue as scheduled despite the Inlet Resistivity reading 771Mohms. |  |  |  |  | TMHARRIS | 2011-01-06
34 | BCP | 2011-01-11 |  | 0 | C100-CAV-VTRF | R2 | 18 | 2 | 15 | HOM B dewar cable and FP dewar cable open during cold test. |  | Freq vs E^2 value was low (3.3 vs. 5.0 expected). Eacc values were subsequently corrected using sqrt (3.3/5.0) |  | Cavity easily passes specification using corrected gradient. Actual Q0 values are nominally higher than calculated in the data file. | kdavis | 2011-01-11
35 |  | 2011-01-11 |  | C100_RI_006 | C100-CAV-VTRF | R2 | 18 | 3 | 21 | Q0 vs Temperature not performed |  | information only |  | no action required |  | 2011-01-11
36 | OTHER | 2011-01-14 | folsom | 0test | C100-CAV-D3 | R1 | 1 |  |  |  | 8621B075-ACC6-2448-329B45603F3BBB6E |  | 8621C629-AD46-4C9F-41AA75675D532736 |  | dia | 2011-01-14 | 8621E1CA-0D0D-6613-71399D3DF28FED8B
37 | DISCREPANCIES | 2011-01-21 | MCCREA | C100-RI-017 | C100-CAV-INST-HELV | R1 | 15 |  |  | WHEN CAVITY RI-017 WAS STOOD UP IN WELD FIXTURE IT WAS IMMEDIATELY NOTICED BY MCCREA THAT IT APPEARED TO BE BENT. IT WAS VERIFIED BY JOHN FISCHER AND JHN HOGAN. THEN IT WAS DECIDED TO BE TACKED AS IS AND SENT TO DANNY FOREHAND FOR A FIELD FLATNESS READING TO SEE IF IT WAS STILL WITHIN TOLERANCE |  |  |  | beadpull and use | ABURRILL | 2012-01-20
38 | DEVIATIONS | 2011-01-21 | JDAVEN | C100_RI_022 | C100-CAV-DEGR | R2 | 95 | 1 | 1 | The large UHV was occupied when the cavity arrived into the Chemroom from EP. |  |  |  | There was a 20 minute extended delay between final steps in EP room and beginning of UHV. | JDAVEN | 2011-01-21
39 | DISCREPANCIES | 2011-01-24 | CASTAGNO | C100-RI-001 | C100-CAV-LAP | R2 | 17 |  |  | Small void can be seen at tube/flange weld of HOM B.
40 | DISCREPANCIES | 2011-01-24 | CASTAGNO | C100-RI-013 | C100-CAV-ASSY2 | R1 | 28 |  |  | After turbo was established, the pressure fluctuated from 1.0e-3 to 5.0e-5. This lasted for about 40 minutes. It finally settled down at 3.0e-5 mbar.
41 | DISCREPANCIES | 2011-01-25 | MCCREA | C100-RI-018 | C100-CAV-INST-HELV | R1 | 16 |  |  | Bellows lifted while tacking, resulting in tuner tab on the HOM end on the big pipe side to have a 32nd gap all else is good |  | none |  |  | MCCREA | 2011-01-25
42 | DEVIATIONS | 2011-01-31 |  | C100_RI_013 | C100-CAV-VTRF | R3 | 31 |  |  | cavity vacuum was out of range because the four line valve was closed for about one hour before testing
43 | HPR | 2011-02-01 | TMHARRIS | C100-RI-017 | C100-CAV-HPR | R1 | 67 |  |  | WHEN CHECKING LEVEL OF CAVITY FOUND THAT THE TOP WAS NOT LEVEL BUT CAVITY WAS. CALLED STEVE INTO CHECK |  |  |  | All OK | CASTAGNO | 2011-02-01
44 | EP | 2011-02-02 | JDAVEN | C100_RI_027 | C100-CAV-EPOL | R1 | 27 | 2 | 8 | Servo motor would not engage the cage for rotation. |  | Cut main power to cabinet in order to manually rotate carriage. After multiple attempts this did not remedy the situation. Disconnected power supply cable to servo motor. Sprayed contacts with cleaner, banged contacts against wall, hooked back up, manually rotated cage. |  | Started program on second attempt carriage rotated, however there was a violent shake at initial start-up. This situation has never properly been resolved. It's always been a problem in the EP process. The frequency in the interruptions have been reduced, but when it does happen the result is 30-45 minutes of lost work time on the front end of a 11-13 hour process. | JDAVEN | 2011-02-02
45 | RF_TEST_ABORTED | 2011-02-02 | CASTAGNO | C100-RI-008 | C100-CAV-VTRF | R2 | 26 | 3 | 19 | FE loaded at 7 MV/m. Attempted to process through radiation barrier unsuccessfully. Total vacuum readings (warm)for this cavity were 1.1e-6 mbar after CR assembly. |  | Removed cavity from insert and placed on CR pump for additional troubleshooting.
46 | LEAK | 2011-02-02 | CASTAGNO | C100-RI-002 | C100-CAV-ASSY2 | R1 | 30 |  |  | Initial leak test showed a small rise when helium was sprayed near the top of the cavity. Cavity appears to be leak tight. |  | Moved insert (28-004) to VTA landing to leak test downpipe conflat near top plate. Conflat is leaking near top plate. Valved off cavity and bled up test stand. Replaced leaking conflat under top plate. Evacuated and leak tested. Opened cavity valve after good leak test. |  | Send to dewar. | CASTAGNO | 2011-02-02
47 | ELNIK_FURNACE | 2011-02-07 | OVERTONR | C100RI029 | C100-CAV-HEAT | R1 | 28 |  |  | Cryo pump lost power during run | 00BC1548-BFED-DE7C-C5CAFE2311A4DC0B | regenerated pump and restarted run |  | run was restarted | OVERTONR | 2011-02-07 | 00BCEF9C-0F12-B2EE-879404D63AEA3EDA
48 | ELNIK_FURNACE | 2011-02-07 | OVERTONR | C100RI030 | C100-CAV-HEAT | R1 | 29 |  |  | Cryo pump lost power during run | 00C034BD-B1AF-69DF-6C5A3CFF31CB24BF | Regenerated pump | 00C06F7D-B5E4-BBED-B88E4A05388A2892 | Restarted run | OVERTONR | 2011-02-07
49 | OTHER | 2011-02-16 | FOREHAND | 0 | C100-CAV-VTRF | R3 | 40 | 2 |  | Cable on HOMb with issues |  | Continue with RF power testing
50 | HPR | 2011-02-18 | TMHARRIS | C100-RI-017 | C100-CAV-HPR | R1 | 78 | 2 | 4 | PROBLEMS WITH ALIGNMENT THEN AT START UP OF HPR PLANT DI WATER WAS 17.50 10 MIN LATER IT WENT TO 18.32 |  |  |  | FYI
51 | HPR | 2011-02-18 | JDAVEN | C100-RI-021 | C100-CAV-HPR | R1 | 79 | 2 | 4 | DI Plant resistivity read 17.33Mohms, HPR inlet resistivity read 17.39Mohms |  |  |  | FYI | ABURRILL
52 | HPR | 2011-02-18 | JDAVEN | C100-RI-021 | C100-CAV-HPR | R1 | 79 | 2 | 4 | Anticipated a raise in the DI plant resistivity point has it had done with C100-RI-017 rinse early today (by Teena Harris), but the DI Plant actually dropped to 17.20mohms by the end of the process. |  | Contacted Andrew Burrill. |  |  | ABURRILL
53 | OTHER | 2011-02-23 | TMHARRIS | C100-RI-032 | C100-CAV-DEGR | R2 | 131 |  |  | Was removing cavity from sonic when noticed crain cable twisted lowered cavity when it came loose and fell |  | sent to cmm |  | removed C50 cavity fixture from use. | ABURRILL
54 | EP | 2011-02-24 | JDAVEN | C100_RI_008 | C100-CAV-EPOL | R1 | 32 | 2 | 8 | The thermocouple for Cell#1 equator was not functioning properly. |  | James Davenport made the decision to continue with the EP b/c all other thermocouples were operating properly. |  | Cell#1 thermocouple began to work consistently around 9:30; it was showing activity prior to that although the readings were wildly inaccurate. AJ Jacobson will be notified for further investigation/repair. | JDAVEN | 2011-02-24
55 | LEAK | 2011-02-24 | CASTAGNO | F067763/04 | [5 empty] | Leak at knife edge of large flange. Two different window assy's were installed onto this waveguide. Both of them leaked. The same two windows were installed onto two different waveguides and were leak tight. | 5D3AA31C-CC15-0DFD-5E815D45336AF670 | Removed waveguide from clean room to further evaluate the knife edge. |  |  | ABURRILL
56 | EP | 2011-02-25 | JDAVEN | C100_RI_008 | C100-CAV-EPOL | R1 | 32 | 2 | 8 | Reference previous D3 regarding EP thermocouple malfunction. |  |  |  | AJ Jacobson inspected the transmitter board to reveal loose port screws for several of the thermocouples. He tightened them. James Davenport ran a test program to check the response of the thermocouples. All properly worked. No replacement was necessary at this time. | JDAVEN | 2011-02-25 | 5E22485C-9D2A-6F98-0598573CFD37AF3E
57 | DEVIATIONS | 2011-02-25 |  | C100_RI_017 | C100-CAV-VTRF | R3 | 42 | 1 | 9 | FP Cable Return Loss measurement was Popen=24.2mW and Pfp=3.46mW (typical values Pfp=13mW). Initial cavity testing indicated problems with the FP cable, limiting the cavity test. The HOMb and the FP cables were switched under the dewar lid. Retesting of the cavity was successful using HOM-B signal for field measurements. Note: this explains 5e12 Q2 for final testing.
58 | DEVIATIONS | 2011-02-25 | JDAVEN | C100-RI-033 | C100-CAV-EPOL | R1 | 33 | 1 | 3 | C100-RI-033 is following new process flow omitting the post-Heat Treat/pre-EP degrease step. Charlie Reece, Anthony Reilly, and Andrew Burrill deemed the aforementioned step unnecessary because "the cavity will never be cleaner than after the heat treat". |  |  |  | FYI | ABURRILL | 2011-07-14
59 | STAINS | 2011-03-03 | MACHA | C100-RI-025 | C100-CAV-ASSY | R1 | 37 |  |  | Getting ready to assemble top-hat onto cavity. Noticed the FPC flange on the cavity did not look uniform in color. Thought it might still be dirty. Wiped with acetone and wiper and this removed and changed pattern of shadows confirming dirt probably left from polishing. |  | Further cleaned flange with acetone and white scotch bright followed with acetone and a wiper. The cavity FPC flange looked much cleaner and more uniform in color. We proceeded with the cavity assembly. |  | Cleaned,noted and proceeded with assembly. | MACHA | 2011-03-03
60 | LEAK | 2011-03-03 | MACHA | C100-RI-020 | [5 empty] | Cavity had cold leak in the dewar. We were unable to get a high power test due to high cavity vacuum. Disassembled the cavity in the cleanroom to inspect flanges. FPC flange had stain lines (thought to be from residue left from polishing when water dried after second HPR. Attempted to clean with acetone and wipe then white scotchbrite and acetone,unsuccessful. Used micro and white scotchbrite and most stains were removed. "Slight" stain line outside seal path remained. |  | Cleaned with micro and white scotchbrite. Rinsed well with DI water. We are going to HPR today, reassemble top-hat and BL flanges and pump tomorrow to proceed to RF-test. |  | CLTF developed | ABURRILL
61 | DEVIATIONS | 2011-03-07 |  | C100-RI-032 | C100-CAV-DEGR | R2 | 129 |  |  | Cavity was being unloaded from the ultrasonic tank when the plastic cell support collar slipped, causing the cavity to shift leading to a dent in the cell nearest the FPC, see photos. Cavity collar was not fully engaged on the cavity cell leading to this damage. | 907164E4-97EC-5392-91C04BFEC6C1A915 | cell style support collar removed from service until it is modified such that this accident cannot happen again. |  | Cavity was given an internal visual inspection which identified no sharp edges or other abnormalities. Cavity progressing to EP as scheduled.
62 | SCRATCH | 2011-03-07 |  | C100-RI-008 | [5 empty] | After several VTA cycles during which C100-8 did not meet specification the cavity was given a visual inspection of the inside of the cavity. The attached report summarizes the findings and a possible cause. | 907A6DFD-0837-D6F9-6FB4F72A8D8D3B0E | Cavity was re-electropolished and visually inspected. Inspection looks satisfactory (see attached) and cavity will progress to VTA testing. | 907AE502-AF65-C42D-081BCA6FC9D6E8EF | Cavity performance satisfactory | ABURRILL | 2011-07-14
63 | LEAK | 2011-03-05 |  | C100-RI-027 | C100-CAV-INST-HELV | R1 | 24 |  |  | 4840-070 bellow had a leak in in a seam on one of the end rings. | 959C97AF-CA43-C838-712BBE205C067DC3 | Seam was welded and leak checked. | 959D0DDF-BEDF-F2DC-1DDFC8DACA8D7E82 | Leak tight. Pictures in vessel 027 traveler. | WORLAND | 2011-03-05
64 | LEAK | 2011-03-08 | DREYFUSS | C100-RI-020 | C100-CAV-ASSY2 | R1 | 44 |  |  | had cold leak |  | remove top hat bottom flange and top flange wiped with IPA rerinsed reassembled
65 | LEAK | 2011-03-08 | DREYFUSS | C100-RI-021 | C100-CAV-ASSY2 | R1 | 45 |  |  | Test stand leak |  | removed top hot leak tested on bench found no leak rerinsed reassembled
66 | LEAK | 2011-03-09 | DREYFUSS | C100-023 | C100-CAV-ASSY2 | R1 | 46 |  |  | cold leak |  | Removed top hat and lower flange lapped top hat re-rinsed reassembled
67 | DISCREPANCIES | 2011-03-11 | DREYFUSS | MPF-039 | [5 empty] | found loose pin on HOM feed thru when putting on cavity |  | gave feed thru to Danny |  | Put on different feed thru | DREYFUSS | 2011-03-11
68 | RF_TUNING | 2011-03-11 | OVERTONR | C100-RI-029 | C100-CAV-TUNE | R3 | 41 |  |  | Cavity needed to be tuned a 4th time to meet spec.there is no place in the traveler to put this data. | BE81A110-9F11-DD85-4CABFD9DF4778FE2 | Cavity was tuned a 4th time | BE829167-B4C2-2E92-DCE29AD0DF331DB5 | After tuning a 4th time cavity was checked by CMM and no bending needed | OVERTONR | 2011-03-11
69 | EP | 2011-03-10 | JDAVEN | C100_RI_036 | C100-CAV-EPOL | R1 | 37 | 1 | 1 | YRC failed to deliver the Hydrofluoric Acid on 10-Mar-2011 as scheduled. |  |  |  | The HF will not arrive until 14-Mar-2011. C100-RI-036 EP will be moved from 14-Mar-2011 to 15-Mar-2011. | JDAVEN | 2011-03-10
70 | EP | 2011-03-15 | JDAVEN | C100_RI_036 | C100-CAV-EPOL | R1 | 37 | 1 | 4 | There is no response from WLS-TC (the wireless data transmitter of the process thermocouples). Process cannot continue until the issue is remedied. |  | WLS-TC has a 12-13 month shelf life due to the unfriendly electronic environment in EP room. Manufacturer warned the life of the part is drastically reduced by variables it operates in (acidic vapors and moisture/humidity inside the EP cabinet) and will need to be replaced yearly. | C4C99DFC-D89B-7D51-1C749448855D0AB7 | Christian Grenoble instructed James Davenport on the operation of replacing the dead WLS-TC and installing new part on computer and software recognition set-up. James wrote a new procedure to be included in the EP traveler at a later date (CP-C100-EP-HDWR). WLS-TC operational as of 11:00. Andrew Burrill requested the EP be moved to 16-Mar-2011 to avoid unnecessary overtime accrual. | JDAVEN | 2011-03-15
71 | EP | 2011-03-16 | JDAVEN | C100_RI_036 | C100-CAV-EPOL | R1 | 37 | 1 | 4 | Andrew Burrill instructed James Davenport to leave C100-RI-036 installed into the EP cabinet after the failure of the WLS-TC 15-Mar-2011. The acidic environment within the cabinet caused the exterior of the cavity to become tarnished in several areas. |  | Update 2011-04-06 (FM): No interior blemish is reported, therefore cavity will proceed as usual with remaining post-processing towards high power qualification (see summary file attached). | 2B4032E9-BE69-2744-ECBCD0EDBA0B9948 | In the future cavities will not be left inside the EP cabinet over night. The acidic environment is not a result of leaks in the plumbing for the EP system, but is more than likely a result of the direct plumbing of the water overflow port into the Trench. Fumes have been witnessed being pulled into the cabinet from this drainage hole by the EP cabinet exhaust fan. | JDAVEN | 2011-03-16
72 | EP | 2011-03-16 | JDAVEN | C100_RI_036 | C100-CAV-EPOL | R1 | 37 | 1 | 4 | Thermocouple for Cell #1 data not responding. WLS-TC was replaced on 15-Mar-2011 with all other thermocouples showing proper response. This particular sensor has been acting up as of late. |  |  |  | C1 thermocouple will need to be replaced in the near future. EP of C100-RI-036 will continue due to being pushed back 3 days on the production schedule already. | JDAVEN | 2011-03-16
73 | EP | 2011-03-16 | JDAVEN | C100_RI_036 | C100-CAV-EPOL | R1 | 37 | 2 | 9 | Flow meter not recording properly.
74 | EP | 2011-03-16 | JDAVEN | C100_RI_036 | C100-CAV-EPOL | R1 | 37 | 3 | 12 | Current oscillation is not reacting as it should during the first run of brand new electrolyte. |  | system may need to be cleaned with micro
75 | EP | 2011-03-15 | JDAVEN | C100_RI_036 | C100-CAV-EPOL | R1 | 37 | 4 | 19 | Resistivity rinse extended 37 minutes to avoid conflict in personnel. Teena Harris had to be in the Cleanroom to continue HPR of C100 String cavities, James Davenport had a 12:00 Doctor's Appt. |  |  |  | Extending rinse did not adversely effect process. | JDAVEN | 2011-03-15
76 | DETOURS | 2011-03-17 | CASTAGNO | C100-RI-027 | C100-CAV-ASSY2 | R1 | 47 |  |  | Stop current processing track of cavity to reposition tuner tabs on end dishes to avoid having to test cavity in VTA twice.
77 | EP | 2011-03-18 | JDAVEN | C100_RI_037 | C100-CAV-EPOL | R1 | 38 | 1 | 1 | C1 thermocouple continues to have malfunction, most recently it appears to have completely died. Inspection of WLS-TC revealed there is an unused lead line, C7. |  |  |  | James Davenport switched dead lead C1 with unused lead C7. A test was performed to verify proper function. C1 how has a functioning thermocouple. Wholesale replacement of thermocouples might be desireable in the near future. | JDAVEN | 2011-03-18
78 | EP | 2011-03-21 | JDAVEN | C100_RI_037 | C100-CAV-EPOL | R1 | 38 | 3 | 12 | Miniscule current oscillation noticed. Decline usually is not witnessed until the 3rd or 4th run on new electrolyte. |  | clean sump
79 | EP | 2011-03-21 | JDAVEN | C100_RI_037 | C100-CAV-EPOL | R1 | 38 | 4 | 19 | The di-water system experienced wholesale malfunctions in resistivity the entire day. |  |  |  | Not possible to forgo any of the rinse processes for the EP, so cycles continued despite low quality of water(resistivity reading 1.1Mohms). | JDAVEN | 2011-03-21
80 | DETOURS | 2011-03-21 | TMHARRIS | C100_RI_037 | C100-CAV-DEGR | R2 | 148 | 1 | 7 | Di-water system was malfunctioning entire day. Water quality was very low (1.1mohms). |  |  |  | Andrew Burrill instructed Teena Harris not to leave the cavity in UHV overnight, but to instead rinse it with the low quality DI-water. | TMHARRIS | 2011-03-21
81 | HPR | 2011-03-21 | TMHARRIS | C100-RI-032 | C100-CAV-DEGR | R2 | 149 | 1 | 3 | Cavity visibly bent. HPR process was not initiated because of the severe banana bend in cavity. Needs to be straightened by CMM. |  |  |  | Cavity straightened and ready to progress with another UHV and initial HPR processes. | TMHARRIS | 2011-03-23
82 | FAST_LEAK | 2011-03-23 | FOREHAND | C100-RI-001 | C100-CAV-TUNE | R2 | 11 | 5 | 5 | Sometime between VTA testing and string assembly, the HOM A feedthrough developed a leak. |  | String to be disassambled and reworked. Replaced HOM A feedthrough with MPF-0033. The new feedthrough is 0.007" shorter in tip length. |  | HOM A filter will need to be re-tuned after the string is turned over to the cryomodule production group. | FOREHAND | 2011-03-23
83 | EP | 2011-03-23 | JDAVEN | C100-RI-038 | C100-CAV-EPOL | R1 | 39 | 3 | 12 | Miniscule current oscillation noticed. |  |  |  | FYI
84 | EP | 2011-03-28 | JDAVEN | C100-RI-040 |  | R1 | 40 | 3 | 12 | Miniscule current oscillation noticed. |  |  |  | FYI
85 | SCRATCH | 2011-03-30 | GEORGED | C100-RI-031 | C100-CAV-LAP | R2 | 37 |  |  | Pits are inside the first cell, divot located on first iris. The issue was noticed during lapping. | 08383F48-959C-9063-F5A696CA7EE13B72 |  |  | Cavity was inspected and tested with good RF performance in the VTA. | ABURRILL | 2011-07-15
86 | HPR | 2011-03-31 | TMHARRIS | C100-RI-020 | C100-CAV-HPR | R1 | 116 | 1 | 3 |  |  |  |  | ??
87 | HPR | 2011-03-31 | JDAVEN | C100_RI_017 | C100-CAV-HPR | R1 | 116 | 1 | 3 | Noticed a peculiar smell when unblanking the cavity in the HPR. Smell dissipated after initial discovery. |  |  |  | FYI
88 |  | 2011-04-02 | FOREHAND | C100-RI-027 | C100-CAV-ASSY2 | R1 | 52 |  |  | Cavity is not being baked prior to testing. This is being done as part of the cold leak Task Force to help identify the source(s) of our cold leaks.
89 |  | 2011-04-05 | JDAVEN | C100_RI_031 | C100-CAV-DEGR | R2 | 158 | 1 | 9 | James Davenport discovered a scratch during post-UHV rinsing/inspection that was not visible during the initial inspection due to aluminum oxide build up. FYI, scratches were shiney when found, suggesting that they may have been introduced after initial lapping and inspection and discovery in chem room. BC 04/05/11 |  | Andrew Burrill and Kurt Macha inspected the flange in question. |  | Cavity was sent back to Isaiah Daniels to be relapped. | JDAVEN | 2011-04-05
90 | OTHER | 2011-04-04 | OVERTONR | C100-RI-035 | C100-CAV-TUNE | R3 | 48 |  |  | Post CMM cavity needed to be tuned again. The traveler only has room for tuning cavity 3 times | 272D6AFB-D8B7-4B3D-D06C34C926A35310 | Cavity was tuned for HV Welding for 4th time and sent back to CMM | 272E1D73-C510-50BF-1C146C5382F8FFF9 | no resolution necessary | OVERTONR | 2011-04-04
91 |  | 2011-04-01 | OVERTONR | C100-RI-033 | C100-CAV-TUNE | R3 | 45 |  |  | Cavity needed to be straightened by CMM and is out of spec. | 27498AB5-B653-331D-875FD78AD4322D57 | Cavity was tuned for a 5th time | 2749AD1A-E66F-3672-1D15E751B9CD5683 |  | OVERTONR | 2011-04-01
92 |  | 2011-04-07 | FOREHAND | 0 | C100-CAV-TUNE | R3 | 38 | [6 empty] | delete me
93 | DEVIATIONS | 2011-04-08 |  | C100-RI-031 | C100-CAV-ASSY2 | R1 | 55 |  |  | Cavity evacuation slow pump down was set to .02" on micrometer instead of .002" leading to a faster than normal pumpdown to 150 mbar when rate was corrected |  |  |  | train the operators on the use of the slow pump down mechanism | ABURRILL | 2011-07-15
94 | DEVIATIONS | 2011-04-08 |  | C100-RI-034 | C100-CAV-ASSY2 | R1 | 56 |  |  | Cavity evacuation slow pump down was set to .02" on micrometer instead of .002" leading to a faster than normal pumpdown to 150 mbar when rate was corrected.
95 | LEAK | 2011-04-15 |  | C100-RI-031 | C100-CAV-VTRF | R3 | 56 |  |  | Following completion of the cavity testing the cavity vacuum rose from |  | Cavity needs dewar leak check following warm up as well as warm leak check to try and identify leak and determine if a HOM feedthru has broken. |  | Cavity was further evaluated in the landing and a leak was discovered to be in the FPC flange. | ABURRILL | 2011-07-15 | 5A2A1669-05E5-52D4-982747F0D4A9602E
96 | EP | 2011-04-18 | JDAVEN | C100-RI-044 | C100-CAV-EPOL | R1 | 45 | 1 | 1 | EP Cabinet will not power up or remain powered up. When it does maintain power long enough to boot up three alarms sound: 1. Storage Tank High Leak 2. Storage Tank Low Leak 3. Storage Tank Electrical Enclosure |  | The scheduled EP was postponed until Facilities could diagnose and solve the issues. C100-RI-044 was removed from the cabinet to avoid creating another "Smurf" cavity like C100-RI-032. |  | Demolition on Building 31 that took place on 16Apr2011 resulted in all the problems today. AJ Jacobson diagnosed the issue with the EP cabinet and put a temporary fix in place. | JDAVEN | 2011-04-18
97 | DEVIATIONS | 2011-04-18 | forehand | C100-RI-027 | C100-CAV-TUNE | R3 | 38 | 4 | 4 | Cavity was found to be be out of CMM spec after HV welding. Tabs were removed and replaced. |  | Extra beadpull taken after weld repair of HV tabs to verify no change. | 6A11EE1D-D17B-5C67-41C1EE23BC048BD8 | Cavity was out of field flatness spec before repair. Testing cavity will determine its use. Continue moving forward with processing as per F. Marhauser. | forehand | 2011-04-18
98 | RF_TUNING | 2011-04-20 |  | 0 | C100-CAV-VTRF | R3 | 58 |  |  | Cavity lock frequency of 1496.8750 MHz higher than specification of 1496.7000 MHz |  |  |  | Cavity is not in the He vessel, so this frequency is not really out of specification | ABURRILL | 2011-07-15
99 | DEVIATIONS | 2011-04-21 | DREYFUSS | ISI-0001 | C100-CAV-INSP-HMFT | R3 | 98 |  |  | This HOM was randomly picked to be part of additional thermal cycling and leak testing. |  | Feedthrough was assembled to a manifold with 5 other HOM's, evacuated and leak tested. After leak test the HOM was baked @ 120C in VTA bake box for 1 hour and leak tested again after bake.The HOM was then cryocycled in D2 in the VTA while under vacuum. Partial pressure of helium was traced during cooldowns. Crocycle consisted of 2x to 4K and then 1x to 2K. | 79863887-D583-8839-584186C134397265 | Feed through tested good and ready to use for production | DREYFUSS | 2011-04-21
100 | DEVIATIONS | 2011-04-21 | DREYFUSS | ISI-0002 | C100-CAV-INSP-HMFT | R3 | 99 |  |  | This HOM was randomly picked to be part of additional thermal cycling and leak testing. |  | Feedthrough was assembled to a manifold with 5 other HOM's, evacuated and leak tested. After leak test the HOM was baked @ 120C in VTA bake box for 1 hour and leak tested again after bake.The HOM was then cryocycled in D2 in the VTA while under vacuum. Partial pressure of helium was traced during cooldowns. Crocycle consisted of 2x to 4K and then 1x to 2K. | 798B1DAB-D38D-3FA8-1B4D3C884353E7FA | Feed through tested good and ready to use for production | DREYFUSS | 2011-04-21
101 | DEVIATIONS | 2011-04-21 | DREYFUSS | ISI-0003 | C100-CAV-INSP-HMFT | R3 | 100 |  |  | This HOM was randomly picked to be part of additional thermal cycling and leak testing. |  | Feedthrough was assembled to a manifold with 5 other HOM's, evacuated and leak tested. After leak test the HOM was baked @ 120C in VTA bake box for 1 hour and leak tested again after bake.The HOM was then cryocycled in D2 in the VTA while under vacuum. Partial pressure of helium was traced during cooldowns. Crocycle consisted of 2x to 4K and then 1x to 2K. | 798CC2E4-D300-E733-792356D4992CF2FC | Feed through tested good and ready to use for production | DREYFUSS | 2011-04-21
102 | DEVIATIONS | 2011-04-21 | DREYFUSS | RI-PRO-02 | C100-CAV-INSP-HMFT | R3 | 101 |  |  | This HOM was randomly picked to be part of additional thermal cycling and leak testing. |  | Feedthrough was assembled to a manifold with 5 other HOM's, evacuated and leak tested. After leak test the HOM was baked @ 120C in VTA bake box for 1 hour and leak tested again after bake.The HOM was then cryocycled in D2 in the VTA while under vacuum. Partial pressure of helium was traced during cooldowns. Crocycle consisted of 2x to 4K and then 1x to 2K. | 798E40D4-AA55-7F6A-FB8A90C66E0DDDAB | Feed through tested good and ready to use for production | DREYFUSS | 2011-04-21
103 | DEVIATIONS | 2011-04-21 | DREYFUSS | MPF-0032 | C100-CAV-INSP-HMFT | R3 | 73 |  |  | This HOM was randomly picked to be part of additional thermal cycling and leak testing. |  | Feedthrough was assembled to a manifold with 5 other HOM's, evacuated and leak tested. After leak test the HOM was baked @ 120C in VTA bake box for 1 hour and leak tested again after bake.The HOM was then cryocycled in D2 in the VTA while under vacuum. Partial pressure of helium was traced during cooldowns. Crocycle consisted of 2x to 4K and then 1x to 2K. | 798FF467-DD25-FFD4-E2B178E4E2029066 | Feed through tested good and ready to use for production | DREYFUSS | 2011-04-21
104 | DEVIATIONS | 2011-04-21 | DREYFUSS | MFP-0033 | C100-CAV-INSP-HMFT | R3 | 74 |  |  | This HOM was randomly picked to be part of additional thermal cycling and leak testing. |  | Feedthrough was assembled to a manifold with 5 other HOM's, evacuated and leak tested. After leak test the HOM was baked @ 120C in VTA bake box for 1 hour and leak tested again after bake.The HOM was then cryocycled in D2 in the VTA while under vacuum. Partial pressure of helium was traced during cooldowns. Crocycle consisted of 2x to 4K and then 1x to 2K. | 7992AB27-AA57-E366-DDB31722D7399803 | Feed through tested good and ready to use for production | DREYFUSS | 2011-04-21
105 | OTHER | 2011-04-25 | DANIELS | C100-RI-035 | C100-CAV-LAP | R2 | 42 |  |  | Changed lapping paper. Lapped cavity FPC flange with new lapping paper in this order.1.40 micron 268L 3M PN 60-0001-2023-2 2. 40 micron 266X 3M PN 60-0600-0219-7 3. 30 micron 266X 3M PN 60-0600-2018-9 4. 12 micron 266X PN 3M 60-0600-0217-1 5. 9 micron 266X PN 3M 60-0600-0216-3 |  | This is a new process with the lapping paper no action has been taken at this time other than using the new lapping paper. |  | N/A | DANIELS | 2011-04-25
106 | DEVIATIONS | 2011-04-25 | CASTAGNO | RI-PRO-024 | C100-CAV-INSP-HMFT | R3 | 127 |  |  | This HOM was randomly picked to be part of additional thermal cycling and leak testing performed after the incoming leak test. | 92656922-E4E9-C108-3D5FB0FBBF25F84F |  | 8E1771BB-EE30-3955-674ABC1350D65F4A
107 | DEVIATIONS | 2011-04-25 | CASTAGNO | RI-PRO-023 | C100-CAV-INSP-HMFT | R3 | 126 |  |  | This HOM was randomly picked to be part of additional thermal cycling and leak testing performed after the incoming leak test. | 9265F6A4-E1F5-4265-BB3E47A87696D9E1 |  | 8E16D495-BDED-7BD8-4987590F68CF7D76 | Test Successful
108 | DEVIATIONS | 2011-04-25 | CASTAGNO | RI-PRO-022 | C100-CAV-INSP-HMFT | R3 | 125 |  |  | This HOM was randomly picked to be part of additional thermal cycling and leak testing performed after the incoming leak test. | 92667059-CF78-98F5-D051F8ED9A4382DB |  | 8E162503-0EF0-E3C2-9873BC6F68EEC9B5 | Test Successful
109 | DEVIATIONS | 2011-04-25 | CASTAGNO | RI-PRO-008 | C100-CAV-INSP-HMFT | R3 | 113 |  |  | This HOM was randomly picked to be part of additional thermal cycling and leak testing performed after the incoming leak test. | 926703E3-0DD7-DCD6-E799AA9D5BA349CC |  | 8E159217-E5EF-559C-D245B68B12C78A7E | Test Successful
110 | DEVIATIONS | 2011-04-25 | CASTAGNO | RI-PRO-014 | C100-CAV-INSP-HMFT | R3 | 117 |  |  | This HOM was randomly picked to be part of additional thermal cycling and leak testing performed after the incoming leak test. | 9267AF37-C35B-9A60-E9A3E44EF2D51CEA |  | 8E148F82-9660-15A3-E54DAF9F642112F2 | Test Successful
111 | DEVIATIONS | 2011-04-25 | CASTAGNO | RI-PRO-016 | C100-CAV-INSP-HMFT | R3 | 119 |  |  | This HOM was randomly picked to be part of additional thermal cycling and leak testing performed after the incoming leak test. | 92683986-F99E-6DE0-79515401F951AC50 |  | 8E134A18-AA18-7FD6-0B39B80BBAED7732 | Test Successful
112 | DEVIATIONS | 2011-04-25 | CASTAGNO | RI-PRO-018 | C100-CAV-INSP-HMFT | R3 | 121 |  |  | This HOM was randomly picked to be part of additional thermal cycling and leak testing performed after the incoming leak test. | 9268B9BA-B725-4F8A-6D5D0E695605FE8C |  | 8E121D0C-EEE5-9AD4-309DD1B04EBA5A83 | Test Successful
113 | OTHER | 2011-04-26 | JDAVEN | C100-RI-035 | C100-CAV-DEGR | R2 | 169 | 1 | 4 | FPC flange had some Sharpie marks that couldn't be removed with Isopropanol, Acetone, Methanol, or light buffing with white scotch-brite. |  |  |  | Andrew Burrill and Steve Castagnola approved the flange for progress. Sharpie marks weren't going to effect seal surface. | JDAVEN | 2011-04-26
114 | EP | 2011-04-19 | JDAVEN | C100-RI-044 | C100-CAV-EPOL | R1 | 45 | 2 | 9 | Acid Flow meter not seeming to register onto the LabView System. |  |  |  | Computer was restarted, system functioning properly
115 | LEAK | 2011-04-27 | DREYFUSS | C100RI027 | C100-CAV-ASSY2 | R1 | 58 |  |  | Cavity cool leak developed at 150K during cooldown in dewar 8 immediately following a 120C bake. Cavity was given a warm dewar leak check and no leak detected. |  | Top and bottom beamline flanges removed, cavity given single cycle on HPR for retesting.
116 | HPR | 2011-04-28 | JDAVEN | C100RI035 | C100-CAV-HPR | R2 | 131 | 1 | 1 | Numerous HPR wand head change outs took place on 27Apr2011. James Davenport had to reposition the default wand head back to the original orientation of top fan jet facing post 4 and bottom fan jet facing post 2. |  |  |  |  | JDAVEN | 2011-04-28
117 | LEAK | 2011-04-28 |  | C100-RI-025 | C100-CAV-VTRF | R3 | 50 |  |  | Cavity vacuum was good, 5e-9 BUT no power to the cavity due to the valve on the top of the cavity being closed and the cavity likely having a cold leak. |  | Cavity valve was opened and the cavity had a slug of gas come out. Likely cold leak |  | Open valves before cooling down. This is now part of the cavity vacuum monitoring program
118 | EP | 2011-04-27 | JDAVEN | C100-RI-039 | C100-CAV-EPOL | R1 | 47 | 2 | 9 | Acid flow meter not registering correctly. Miniscule amounts of current oscillation. |  |  |  | AJ Jacobson has been notified about the acid flow meter situation. Current oscillation abscence is an ongoing issue with no apparent resolution. | JDAVEN | 2011-04-27
119 | HPR | 2011-05-03 | JDAVEN | C100-RI-036 | C100-CAV-HPR | R2 | 133 | 2 | 4 | Tony Reilly requested the 2nd HPR process be terminated prior to its completion. |  | New radial wedge beamtube hardware is being put into use. It is a possible remedy for the numerous cold leaks experienced during VTA testing. |  | Old style radial wedges will be removed, new radial wedges will be installed. Final HPR will take place 04May2011. | JDAVEN | 2011-05-03
120 | DEVIATIONS | 2011-05-04 | DREYFUSS | C100-RI-036 | C100-CAV-ASSY | R1 | 47 |  |  | cavity was stopped in middle of HPR to remove top beamline cap |  | Replaced top beamline cap with new S/S fasit plates on radial wedge clamp. |  | part of CLTF | DREYFUSS | 2011-05-04
121 | LEAK | 2011-04-30 |  | C100-RI-027 | C100-CAV-ASSY2 | R1 | 60 |  |  | C100-RI-027 developed cold leak in VTA dewar 8 at around 70 Torr. This was the second cold leak for this cavity. No RF testing was attempted.
122 | OTHER | 2011-04-25 | forehand | C100-RI-027 | C100-CAV-TUNE | R3 | 38 | 5 | 5 | Cavity did not test the first round due to cold leak |  | Re-checked HOM feed-throughs and tuned as necessary. | C6E885DA-ED1F-B86A-BCB4CCF042876CD6 | no resolution necessary | forehand | 2011-05-06
123 | RF_TUNING | 2011-05-09 | ABURRILL | C100-RI-036 | C100-CAV-TUNE | R4 | 50 |  |  | ISI-0005 HOM feedthru failed during HOM tuning in the landing following 120C bake. The braze between the N-type barrel and the feed-thru body broke, see attached photos. | D526FC68-A755-6F0C-6874DEB4CFD5C188 | H. Park to follow up with vendor |  | Feed-thrus have been evaluated using a torque wrench to look for other failures. Cause seems to be isolated to a few articles due to lack of braze material being used during the fabrication. Suggest a better contractual relationship with suppliers as JLab appears to be incurring excessive costs for vendor non-conformance. | ABURRILL
124 | DEVIATIONS | 2011-05-10 | DREYFUSS | C100-RI-032 | C100-CAV-ASSY2 | R2 | 65 |  |  | cavity is for HOM testing not for high power so no HPR and all one assembly |  |  |  | Test Successful | DREYFUSS | 2011-05-10
125 | EP | 2011-04-27 |  | C100-RI-039 | C100-CAV-EPOL | R1 | 47 | 2 | 9 | The acid flow meter was not registering correctly. |  |  |  | AJ Jacobson was notified. | JDAVEN | 2011-04-27
126 | EP | 2011-04-27 | JDAVEN | C100-RI-039 | C100-CAV-EPOL | R1 | 47 | 3 | 12 | Miniscule current oscillation. |  |  |  |  | JDAVEN | 2011-04-27
127 | EP | 2011-05-02 | JDAVEN | C100-RI-045 | C100-CAV-EPOL | R2 | 48 | 2 | 9 | Acid flow meter stopped registering periodically throughout the process. |  |  |  | Reiterated the issue to facilities. | JDAVEN | 2011-05-02
128 | EP | 2011-05-02 | JDAVEN | C100-RI-045 | C100-CAV-EPOL | R2 | 48 | 3 | 12 | Miniscule current oscillation. |  |  |  |  | JDAVEN | 2011-05-02
129 | EP | 2011-05-04 | JDAVEN | C100-RI-046 | C100-CAV-EPOL | R2 | 49 | 2 | 9 | Acid flow meter stopped registering periodically throughout the process. |  |  |  | Facilities notified previously. | JDAVEN | 2011-05-04
131 | EP | 2011-05-11 | JDAVEN | C100-RI-048 | C100-CAV-EPOL | R2 | 50 | 2 | 9 | Acid flow meter stopped registering periodically throughout process run. |  |  |  | Facilities notified previously. | JDAVEN | 2011-05-11
132 | DEVIATIONS | 2011-05-16 | DREYFUSS | C100-RI-016 | C100-CAV-ASSY | R1 | 51 |  |  | Test of new beamline fassit plates |  | switch from AL fassit plates to S/S fassit plates |  | part of CLTF investigation
133 | DEVIATIONS | 2011-05-16 |  | C100-RI-032 | C100-CAV-ASSY | R1 | 50 |  |  | Test of new beamline facet plates |  | switch from AL facet plates to S/S facet plates
134 | DEVIATIONS | 2011-05-16 | DREYFUSS | C100-RI-036 | C100-CAV-ASSY | R1 | 52 |  |  | Test of new beamline fascist plates |  | switch from AL fascist plates to S/S fascist plates
135 | EP | 2011-05-16 | FOLLKIEJ | C100-RI-047 | C100-CAV-EPOL | R2 | 51 | 2 | 9 | Acid flow meter did not function properly at all during the course of the entire run. |  |  |  | Facilities has been aware of this issue for going on 2 months now. | FOLLKIEJ | 2011-05-16
136 | DEVIATIONS | 2011-05-18 | CASTAGNO | C100-032 | C100-CAV-BAKE | R2 | 37 |  |  | This cavity will be baked @ 120C for 5 hours to simulate thermal stresses on the gaskets to see if heating the cavity contributes to cold leaks in the dewar. HOM's were the only components replaced since last time in the dewar. |  | Change bake time to 5 hrs and bake cavity
137 | EP | 2011-05-18 | JDAVEN | C100-RI-049 | C100-CAV-EPOL | R2 | 52 | 2 | 9 | Acid flow meter ceased to register an hour into polish process. |  |  |  | Re-emphasize the importance of a proper functioning acid flow meter registration to the Facilities group. | JDAVEN | 2011-05-18
138 | EP | 2011-05-23 | JDAVEN | C100-RI-050 | C100-CAV-EPOL | R2 | 53 | 2 | 9 | Acid flow meter did not register for the majority of the process run. |  |  |  | Facilities will address. | JDAVEN | 2011-05-23
139 | EP | 2011-05-23 | JDAVEN | C100-RI-050 | C100-CAV-EPOL | R2 | 53 | 3 | 12 | Minimal current oscillation was observed. |  |  |  |  | JDAVEN | 2011-05-23
140 | OTHER | 2011-05-24 | FOREHAND | C100-RI-032 | C100-CAV-TUNE | R3 | 44 | 4 |  | This cavity was used for RI HOM feed-through testing. | 232A116B-E197-D093-F38A66286FF3E830 | n/a |  | n/a | FOREHAND | 2011-05-24
141 | OTHER | 2011-05-25 | CASTAGNO | C100-RI-036 | C100-CAV-ASSY2 | R3 | 73 |  |  | Slow pumpdown took nearly 24 hrs to complete. Throttle valve was manually adjusted and monitored until block valve could be opened. |  | recalibrate throttle valve
142 | DEVIATIONS | 2011-05-25 | DREYFUSS | C100-RI-036 | C100-CAV-ASSY2 | R3 | 73 |  |  | removed lower beamline flange |  | Re HPR cavity |  | Reassembled cavity for testing | DREYFUSS | 2011-05-25
143 | DEVIATIONS | 2011-05-25 | JDAVEN | C100-RI-038 | C100-CAV-DEGR | R3 | 189 | 1 | 9 | When James Davenport was assisting Jim Follkie in caging C100-RI-038, it was discovered that the FPC flange was never lapped. |  |  |  | The cavity was returned to the lapping room. It then needed to be CMM'ed prior to progession along the use process flow. | JDAVEN | 2011-05-25
144 | DEVIATIONS | 2011-05-31 | CASTAGNO | C100-RI-038 | C100-CAV-ASSY2 | R3 | 75 |  |  | Prior to 2nd HPR, questionable marks were observed inside the beamline cap before 1st assembly. The 1st assembly proceeded as normal. Other cavities inside the clean room were inspected for these marks inside the beamtube on FPC end. All cavities in CR have marks inside the beam tube. |  | After discussion with Andrew B. we were questioning the alignment of the cavity in the HPR tool. Upper beam-cap was removed, and cavity was installed into HPR tool and aligned using shims from initial alignment. The alignment was visually verified (beamcap was not installed) and 2nd HPR resumed. Not sure where marks inside beamcap and cavity are comming from. The HPR tool is not the source of these marks. Not sure where these marks a coming from.
145 | EP | 2011-05-31 | JDAVEN | C100-RI-051 | C100-CAV-EPOL | R2 | 54 | 2 | 9 | Acid flow meter ceased to record. |  |  |  | AJ Jacobson inspected the unit 26May2011. He suggested it be replaced, but did some maintenance as a stop-gap. It did not prove to be successful. | JDAVEN | 2011-05-31
146 | EP | 2011-05-31 | JDAVEN | C100-RI-051 | C100-CAV-EPOL | R2 | 54 | 3 | 11 | Glycol supply for cavity cooling system is broken. Construction workers stepped on the thermostat joint in Nov2010; it was thought to be repaired. This was not the case. |  | Incoming water from the city to the heat exchanger was higher than process parameters allow. Could not cool the supply at all due to the absence of glycol inlet to system. |  | Lowered the process chiller temp from 37.5 to 35.5 to 33.5 then finally to 31.5 in an attempt to offset for the lack of lower cooling liquid. The cell temps ran higher due to this system malfunction. Phil Denny and Mike Sprouse assumed the responsibility of repairing the damage system. | JDAVEN | 2011-05-31
147 | EP | 2011-06-02 | JDAVEN | C100-RI-052 | C100-CAV-EPOL | R2 | 55 | 1 | 6 | WLS-IFC wireless receiver was not receiving any data from WLS-TC wireless transmitter in reference to the cavity thermocouples. |  | LabView was rebooted three times in an effort to correct situation. External battery was unhooked in attempt to rule out connection issues with WLS-TC. None of the efforts to resolve problem proved successful. James Davenport made the executive decision to reschedule C100RI052 EP process for 03June2011. The spare set of wireless data collection hardware was "borrowed" by another group and wasn't replaced with new hardware. |  | Tom Goodman slated to purchase two new complete sets of wireless data collection hardware, charging one complete set to the group that confiscated the spare set. EDIT: Tom Goodman possessed the spare WLC-TC. New unit was installed and system checked. Everything operational with wireless data collection system. | JDAVEN | 2011-06-02
148 | HPR | 2011-06-01 | JDAVEN | C100-RI-023 | C100-CAV-HPR | R2 | 148 | 1 | 1 | Jim Follkie initiated C100RI023 HPR two separate times but multiple power outages were effecting several areas of Test Lab. HPR cabinet lost power several times, preventing a complete process to be executed. |  | HPR Cabinet powered down for period of time to avoid continuous interruption of power supply. James Davenport powered cabinet back up 16:30 when building power had been continuous for nearly an hour. |  | Andrew Burrill instructed the Chem staff to abort any further HPR runs for the day, as the power situation was unpredictable at best. HPR will push over into 02June2011 work day. | JDAVEN | 2011-06-01
149 | ELNIK_FURNACE | 2011-06-02 | OVERTONR | F356224/07 | C100-CAV-INSP-WGD | R6 | 123 | 5 |  | power outage during end of run furnace was already cooling down |  | no action needed |  |  | OVERTONR | 2011-06-02
150 | ELNIK_FURNACE | 2011-06-02 | OVERTONR | F356224/02 | C100-CAV-INSP-WGD | R6 | 118 | 5 |  | Power outage during run furnace was already cooling down |  | no action needed |  |  | OVERTONR | 2011-06-02
151 | ELNIK_FURNACE | 2011-06-02 | OVERTONR | F890196/01 | C100-CAV-INSP-WGD | R6 | 125 | 5 |  | power outage during run furnace already cooling down |  | no action needed |  |  | OVERTONR | 2011-06-02
152 | ELNIK_FURNACE | 2011-06-02 | OVERTONR | F890195/04 | C100-CAV-INSP-WGD | R6 | 122 | 5 |  | power outage during run furnace already cooling down |  | no action needed |  |  | OVERTONR | 2011-06-02
153 | DEVIATIONS | 2011-06-07 | DREYFUSS | C100-RI-037 | C100-CAV-ASSY2 | R3 | 76 |  |  | cavity was tested passed but leaked we think. |  | Removed top beamline flange, Tophat and bottom beamline flange and rerinsed in HPR. |  | Put cavity back together and retest | DREYFUSS | 2011-06-07
154 | HPR | 2011-06-06 | TMHARRIS | C100-RI-037 | C100-CAV-HPR | R2 | 150 | 2 | 4 | C100RI037 is to serve as a canary cavity to qualify Production Cleanroom after the Glycol Incident on 02June2011. |  | Cavity performed poorly, early onset Q drop, field emission loaded. |  | Re-rinse and re-test | TMHARRIS | 2011-06-06
155 | HPR | 2011-06-06 | TMHARRIS | C100-RI-037 | C100-CAV-HPR | R2 | 150 | 2 | 4 | Test Lab power glitch caused the HPR to shut off prematurely. |  |  |  | Andrew Burrill instructed Teena Harris to run an additional loopback to ensure proper rinsing in the canary cavity. | TMHARRIS | 2011-06-06
156 | EP | 2011-06-02 | JDAVEN | C100-RI-052 | C100-CAV-EPOL | R2 | 55 | 2 | 6 | Glycol supply system had a disasterous leak. All glycol pressure lost. Repair is slated for 02June2011 before close of business. |  | C100RI052 EP cycle scheduled for 03June2011 (rescheduled from 02June2011) will be post-poned until 06June2011 to ensure all necessary repairs can be made with proper time frame to system check. |  | Facilities/Plant Department repaired broken bushing on the glycol supply line for the Cavity Cooling System heat exchanger. Thermostat was removed to prevent future incidents. | JDAVEN | 2011-06-02
157 | EP | 2011-06-06 | JDAVEN | C100-RI-052 | C100-CAV-EPOL | R2 | 55 | 2 | 9 | Acid flow meter ceased to register. |  |  |  | Facilities suggested a complete replacement of the acid flow meter when applicable. | JDAVEN | 2011-06-06
158 | EP | 2011-06-06 |  | C100-RI-052 | C100-CAV-EPOL | R2 | 55 | 3 | 12 | Minimal current oscillation observed. Drastic drop off from previous cycle. This is only 2nd cavity processed on this batch.
159 | EP | 2011-06-06 |  | C100-RI-052 | C100-CAV-EPOL | R2 | 55 | 3 | 12 | A power outage occurred at 8:39am. This caused the EP tool to power down, thereby interfering with the polish cycle. |  | The polish was interrupted for 11 minutes from 8:39-8:50. Once power was re-established to the cabinet and all the necessary precursory steps were taken, the polish resumed where it was ceased by the power outage. |  | Polish steps began at 8:52; total down time was 13 minutes from initial power bump to restart of the process. Investigation revealed the power blimp was caused by Dominion Va Power. | JDAVEN | 2011-06-06
160 | LEAK | 2011-06-08 | CASTAGNO | C100-RI-037 | C100-CAV-VTRF | R3 | 64 |  |  | Cold leak after cooldown. Cavity RF tested despite higher than usual cavity pressure (see attachment "C100-RI-037-VTA-COOL Processed Data.xlsx") | 70C52630-D263-F57E-0D83F21C57E3D1E8 | Individually leak tested each beam line seal and FPC seal with LN2 in VTA mez. Found leak @ FPC flange. |  | remove FPC, lap tophat, re-rinse and re-install.
161 | HPR | 2011-06-08 | TMHARRIS | C100-RI-029 | C100-CAV-HPR | R2 | 151 | 1 | 1 | Cavity being HPR again due to Glycol Incident on 02June2011.
162 | HPR | 2011-06-08 | TMHARRIS | C100-RI-023 | C100-CAV-HPR | R2 | 152 | 1 | 1 | Cavity being HPR again due to Glycol Incident on 02June2011.
163 | HPR | 2011-06-09 | JDAVEN | C100-RI-021 | C100-CAV-HPR | R2 | 153 | 1 | 1 | Cavity being HPR again due to Glycol Incident on 02June2011.
164 | HPR | 2011-06-09 | JDAVEN | C100-RI-021 | C100-CAV-HPR | R2 | 153 | 2 | 5 | Andrew Burrill requested the HPR process be terminated. Based on continued power outages |  |  |  | HPR process was shutdown at 6:38. Cavity was removed and left in Main Assembly area uncovered to prevent water from being trapped inside. | JDAVEN | 2011-06-09
165 | OTHER | 2011-06-02 | JDAVEN | C100RI029 | [5 empty] | Glycol Incident occurred while this cavity was in Main Assembly area. | 742E48CB-92DC-15CC-095E9393C3AF22FB |  | 742C4211-AC85-5E33-6EA07330BCE5318C | Cavity remained blanked off and in the Main Assembly room. | JDAVEN | 2011-06-02
166 | OTHER | 2011-06-02 | JDAVEN | C100RI021 | [5 empty] | Glycol Incident occurred while this cavity was in HPR. | 742D1D0C-B79E-A412-D2A78B4F3D87AC98 |  |  | Cavity removed from HPR, blanked off and left in Main Assembly room. | JDAVEN | 2011-06-02
167 | OTHER | 2011-06-02 | JDAVEN | C100RI023 | [5 empty] | Glycol Incident occurred while this cavity was uncovered in Class 10 Tent. Cavity had been HPR on 02June2011 for String Processing. | 742DAC1E-FF7A-349F-5CBA9D777A20EBD9 |  |  | Cavity blanked off and relocated to Main Assembly room. | JDAVEN | 2011-06-02
168 | OTHER | 2011-06-02 | JDAVEN | C100RI027 | [5 empty] | Glycol Incident occurred while this cavity was in Class 10 room located on large baker's rack. Cavity was blanked off and several feet away from drop zone.
169 | OTHER | 2011-06-02 | JDAVEN | C100RI033 | [5 empty] | Glycol Incident occurred while this cavity was in Class 10 room located on large baker's rack. Cavity was blanked off and several feet away from drop zone. | 742F7059-D005-8C34-25EF1CD89D6DB408
170 | OTHER | 2011-06-02 | JDAVEN | C100RI031 | [5 empty] | Glycol Incident occurred while this cavity was in Class 10 room located on large baker's rack. Cavity was blanked off and several feet away from drop zone. | 742FBEFC-C88A-DE0C-D7B77C916B06B8F9
171 | OTHER | 2011-06-02 | JDAVEN | C100RI035 | [5 empty] | Glycol Incident occurred while this cavity was in Class 10 room located on large baker's rack. Cavity was blanked off and several feet away from drop zone. | 74302197-0938-BAB2-888810FE7EAE2F61
172 | OTHER | 2011-06-02 | JDAVEN | C100RI034 | [5 empty] | Glycol Incident occurred while this cavity was in Vertial Attachment Area located on cavity tree. Cavity was blanked off and in entirely separate room from drop zone. | 74306A11-A854-B296-2840FCF3C20A8709 |  | 742B9A29-9E87-6DAF-BD6FBAA86CD15FEC
173 | OTHER | 2011-06-02 | JDAVEN | C100RI014 | [5 empty] | Glycol Incident occurred while this cavity was in Vertial Attachment Area located on cavity tree. Cavity was blanked off and in entirely separate room from drop zone. | 74316E4B-E57D-2F15-E41634EE1FB72CF5
174 | OTHER | 2011-06-02 | JDAVEN | C100RI028 | [5 empty] | Glycol Incident occurred while this cavity was in Vertial Attachment Area located on cavity tree. Cavity was blanked off and in entirely separate room from drop zone. | 74324521-C53A-9F81-A51577B62F98D039
175 | OTHER | 2011-06-02 | JDAVEN | C100RI022 | [5 empty] | Glycol Incident occurred while this cavity was in Vertial Attachment Area located on cavity tree. Cavity was blanked off and in entirely separate room from drop zone. | 7432E4BC-E989-7BB0-11A3904220B7CE7E
176 | OTHER | 2011-06-02 | JDAVEN | C100RI030 | [5 empty] | Glycol Incident occurred while this cavity was in Vertial Attachment Area located on cavity tree. Cavity was blanked off and in entirely separate room from drop zone. | 7433C441-E402-48B1-13D8692021F3F4E4
177 | OTHER | 2011-06-02 | JDAVEN | C100RI025 | [5 empty] | Glycol Incident occurred while this cavity was in Vertial Attachment Area located on cavity tree. Cavity was blanked off and in entirely separate room from drop zone. | 74358D16-ED69-84C4-B0D3548037F8BAF9
178 | DISCREPANCIES | 2011-06-16 | ABURRILL | C100-RI-024 | C100-CAV-VTRF | R3 | 67 |  |  | Cavity being held from use in production cycle due to HOM field tilt as identified by Haipeng. |  | Per Haipeng: After making up missing data on June 3, 2011 into the Traveler and analyzing the HOM data for the C100-RI-024 cavity (without helium vessel), I found previous identified this tilt HOM field profile's cavity marginally passed specification at ~2884MHz. With any handing process later, we have to carefully to watch the bead-pull and VTA data again for the final QA. Regards |  | Cavity taken out of use for C100 program, to be evaluated | ABURRILL | 2012-01-20
179 | DEVIATIONS | 2011-06-21 | JDAVEN | C100-RI-039 | C100-CAV-DEGR | R3 | 202 | 1 | 1 | C100RI039 is being degreased prior to an off-cycle "hot water" rinse in the Closed Chemistry Cabinet. The cavity will dry after the hot rinse then be baked along with C100RI040 and C100RI041.
180 | DEVIATIONS | 2011-06-21 | TMHARRIS | C100-RI-041 | C100-CAV-DEGR | R3 | 203 | 1 | 1 | C100RI041 is being degreased prior to an off-cycle "hot water" rinse in the Closed Chemistry Cabinet. The cavity will dry after the hot rinse then be baked along with C100RI040 and C100RI039.
181 | DEVIATIONS | 2011-06-21 | JDAVEN | C100-RI-040 | C100-CAV-DEGR | R3 | 204 | 1 | 1 | C100RI040 is being degreased prior to an off-cycle "hot water" rinse in the Closed Chemistry Cabinet. The cavity will dry after the hot rinse then be baked along with C100RI039 and C100RI041.
182 | DEVIATIONS | 2011-06-21 | TMHARRIS | C100RI039 | [5 empty] | C100RI039 was Hot Rinsed for an hour at the request of Andrew Burrill. It will dry overnight prior to a Bake the following day. |  |  |  |  | TMHARRIS | 2011-06-21
183 | DEVIATIONS | 2011-06-21 | TMHARRIS | C100RI041 | [5 empty] | C100RI041 was Hot Rinsed for an hour at the request of Andrew Burrill. It will dry overnight prior to a Bake the following day. |  |  |  |  | TMHARRIS | 2011-06-21
184 | DEVIATIONS | 2011-06-21 | TMHARRIS | C100RI040 | [5 empty] | C100RI040 was Hot Rinsed for an hour at the request of Andrew Burrill. It will dry overnight prior to a Bake the following day. |  |  |  |  | TMHARRIS | 2011-06-21
185 | EP | 2011-06-15 | FOLLKIEJ | C100-RI-054 | C100-CAV-EPOL | R2 | 57 | 2 | 9 | Acid flow meter ceased to register on LabView at 6:05.
186 | EP | 2011-06-15 | FOLLKIEJ | C100-RI-054 | C100-CAV-EPOL | R2 | 57 | 3 | 11 | Miniscule current oscillation.
187 | EP | 2011-06-15 |  | C100-RI-054 | C100-CAV-EPOL | R2 | 57 | 4 | 19 | DI water pressure alarm sounded at 10:40. Alarm proceeded to sound at every subsequent fill during Step 21/Fill & Dump cycle. |  |  |  | The process was not interrupted. All other DI water usage was terminated until the EP cycle had completed the usage of the resource. The cause for the alarms was a drastic drop in DI pressure at the DI bottles in the R.O. system. | FOLLKIEJ | 2011-06-15
188 | DISCREPANCIES | 2011-06-30 | GEORGED | C100-RI-039 | C100-CAV-LAP | R2 | 49 | 1 | 1 | Waviness and pits present in seal surface of HOM A. It was discovered at beginning of lapping step. | E118BA37-FBE9-B514-4077F92A547134B7 | Seal surface was repaired with 40 micron paper and scotchbrite. | E11965B9-96BA-2A41-6DC70AA7943730BB
189 | HPR | 2011-06-16 | TMHARRIS | C100-RI-023 | C100-CAV-HPR | R2 | 158 | 2 | 4 | DI-water supply hose ruptured during the initial start of the 2nd pass. Gate valve snagged on hose and wand travel caused breakage. |  | HPR was aborted for C100RI023. Facilities will repair damaged hose and inspect wand for any bending or damage. |  | C100RI023 was removed from the queue for 12GeV String #3. Cavity will be internally inspected for possible damage resulting from wand head. Path forward will be determined from that point by management. | TMHARRIS | 2011-06-16
190 | EP | 2011-07-06 | JDAVEN | C100_RI_056 | C100-CAV-EPOL | R2 | 59 | 2 | 8 | Servo Motor doubled RPM speed during several initial start-up attempts. Andrew Burrill was notified. |  | Process delayed by 1 hour 30 minutes while awaiting representation from Management and Facilities groups. James Davenport attempted to recreate the malfunction for Derrick Dail to observe/diagnose. Issue could not be replicated after four attempts. Andrew Burrill requested the EP continue without any further delay. |  | AJ Jacobson will inspect the servo and cables at first opportunity. | JDAVEN | 2011-07-06
191 | EP | 2011-07-06 | JDAVEN | C100-RI-056 | C100-CAV-EPOL | R2 | 59 | 3 | 12 | Inclement weather (serious electrical storms) caused power suply to Test Lab to flicker. All major devices effected. Process was interrupted from 9:32-9:39. |  |  |  | Process was able to return to normal operation after 9:39. | JDAVEN | 2011-07-06
192 | HPR | 2011-07-07 | JDAVEN | C100-RI-041 | C100-CAV-HPR | R2 | 169 | 2 | 5 | C100RI041 was in the Class 10 during the inclement weather power outage on 06-July-2011. |  |  |  | Cavity may have to be re-HPR'ed as per Steve Castagnola and Andrew Burrill. DI Water system is currently down due to inclement weather problem as well. | JDAVEN | 2011-07-07
193 | EP | 2011-07-14 | FOLLKIEJ | C100-RI-057 | C100-CAV-EPOL | R2 | 60 | 1 | 1 | Error message during the start up of LabView program. Process start delayed until situation solved. |  |  |  | Computer rebooted and error message did not appear again during start up process. | FOLLKIEJ | 2011-07-14
194 | EP | 2011-07-14 | FOLLKIEJ | C100-RI-057 | C100-CAV-EPOL | R2 | 60 | 1 | 1 | "Home" position sensor for carriage needed to be re-attached. |  |  |  |  | FOLLKIEJ | 2011-07-14
195 | EP | 2011-07-14 | FOLLKIEJ | C100-RI-057 | C100-CAV-EPOL | R2 | 60 | 4 | 19 | Error message appeared on LabView again. No data was recorded for the fill and dumps for this reason.
196 | OTHER | 2011-07-25 | MACHA | 0 | [5 empty] | A smell has been noticed near the CMM room entrance. This morning Chris said he was able to smell this odor in the class 10 room. Phil and I entered the clean room about 9:30 am and could not smell anything. I believe the odor smells like sewer decaying matter. It looks like a roof drain is piped near the CMM door. It also rained last night possibly causing movement of the odor. Phil was going to look into this.
197 | OTHER | 2011-07-26 |  | C100_RI_041 | C100-CAV-VTRF | R3 | 76 |  |  | Cavity lock frequency = 1496.8509 MHz over max allowed lock frequency of 1496.700 MHz (no helium vessel attached) |  |  |  | retuned and use | ABURRILL
198 | DEVIATIONS | 2011-07-20 |  | C100-RI-035 | C100-CAV-ASSY | R3 | 65 |  |  | HOM ISI-0002 was broken during the leak checking process following C100-RI-035 being dropped in the cleanroom. The Feedthru was replaced with ISI-0041 as noted in the assembly traveler |  | Feedthru is with CMM group for proper disposal.
199 | OTHER | 2011-07-27 |  | C100_RI_039 | C100-CAV-VTRF | R3 | 77 |  |  | Cavity lock frequency = 1496.8673MHz higher than max freq lock spec of 1496.700MHz. (Cavity does not have helium vessel) |  |  |  | No He vessel accounts for delta | ABURRILL | 2012-01-20
200 | EP | 2011-07-19 | JDAVEN | C100-RI-058 | C100-CAV-EPOL | R2 | 61 | 3 | 12 | Minimal current oscillation observed.
201 | EP | 2011-07-27 | JDAVEN | C100-RI-059 | C100-CAV-EPOL | R2 | 62 | 2 | 8 | Servo Motor would not engage properly. Cage did not rotate at all. |  | James Davenport attempted to remedy the issue from 5:45 until 6:55 to no avail. There were three steps taken in an attempt to solve the problem: 1. Cabinet powered down in to allot for manual rotation of cage in hopes of moving past a possible "dead" spot. 2. Power supply cables disconnected from servo motor and cleaned with electrical contact cleaner. 3. Servo motor electrical ports cleaned with the aforementioned contact cleaner. |  | The three prior mentioned steps were performed repeatedly for over an hour. It was decided to post-pone EP until Edward "AJ" Jacobson could provide a more knowledgeable assessment of the situation.
202 | HPR | 2011-07-21 | JDAVEN | C100-RI-023 | C100-CAV-HPR | R2 | 179 | 1 | 1 | C100RI023 was initially dropped by Teena Harris and Era Perry prior to the HPR during the caging process. |  | The cavity was passed out of the cleanroom and rechecked by CCM (Brian Carpenter) and Tuning (Roland Overton). Cavity passed all inspections free of damage or physical alteration from the fall. |  | use as is | ABURRILL
203 | EP | 2011-08-04 | FOLLKIEJ | C100-RI-060 | C100-CAV-EPOL | R2 | 63 | 3 | 12 | Minimal current oscillation.
204 | OTHER | 2011-08-05 |  | 0 | [5 empty] | Mortenson had to paint the floor with epoxy in the new addition east side of test lab. The worked started on 8/4/2011. Mortenson said it will be on and off for the next month with the epoxy work. No epoxy smell was noted inside test lab on 8/4/2011.
205 | EP | 2011-08-08 | JDAVEN | C100-RI-061 | C100-CAV-EPOL | R2 | 64 | 3 | 12 | Minimal current oscillation. |  |  |  | noted
206 | OTHER | 2011-08-09 | DREYFUSS | C100-RI-042 | C100-CAV-ASSY2 | R3 | 85 |  |  | smoke in the air due to fires outside smoke smell in cleanroom and particle counts on computer high | D0D5A58B-ED4B-232D-A371CC1941BE1671 | assembled cavity |  | fyi
207 | OTHER | 2011-08-09 | DREYFUSS | C100-RI-043 | C100-CAV-ASSY2 | R3 | 84 |  |  | smoke and smell from fires outside visibly in the air when turbo was turned on cavity was under vacuum on rough pump | D0F03DC6-A7E7-DB60-B96465BC99F3DEF4 | continued pump down of cavity |  | fyi
208 | HPR | 2011-08-09 | JDAVEN | C100-RI-044 | C100-CAV-HPR | R2 | 185 | 2 | 4 | At 16:45 Jim Follkie noticed on the HPR webpage that the pressure dropped in the HPR cycle between 14:30-14:45. James Davenport instructed Jim Follkie to terminate the cycle and remove the cavity. |  | Upon arriving inside the Cleanroom, Jim discovered the braided metal supply hose to the HPR wand had ruptured, sending water everywhere. The process hadn't appeared to terminate despite the lack of pressure to the wand. The table was still rotating and the wand was still progressing down. |  | Derrick Dail replaced the damaged line on 09Aug2011. James Davenport tested the system 10Aug2011 with no visual evidence of any further damage. The pump pressure did have to be adjusted back to normal operating parameters as it had climaxed to 1600psi during the test run. | JDAVEN | 2011-08-10
209 | HPR | 2011-08-10 | TMHARRIS | C100-RI-044 | C100-CAV-HPR | R2 | 186 | 2 | 4 | The pump pressure had to be adjusted due to the events on 09Aug2011. Derrick Dail adjusted the pressure from the pump room while Teena Harris verified the correction from inside the Cleanroom at the HPR maintenance access. |  |  |  | Pump pressure back within normal operating parameters. | TMHARRIS | 2011-08-10
210 | EP | 2011-08-10 | JDAVEN | C100-RI-063 | C100-CAV-EPOL | R2 | 65 | 1 | 6 | WLS-TC was dead. Attempted to replace with spare unit only to discover spare unit was dead as well; spare was a supposedly refurbished part supplied by vendor. |  |  |  | EP postponed until ILC welding unit could be borrowed via Ari Palcheski. Steve Castagnola tabbed to purchase two brand new units by Larry King. EP begin 5 hours after initial start attempt. WLS-TC only lasted from 02Aug-10Aug, just over a week. | JDAVEN | 2011-08-10
211 | EP | 2011-08-10 | JDAVEN | C100-RI-063 | C100-CAV-EPOL | R2 | 65 | 2 | 10 | The acid out thermocouple is going bad. Reading is sporadic and no representative of the actual acid out-going temperature. |  |  |  | Replace Type K thermocouple attached to the acid out line (brown) inside the EP cabinet 11Aug2011. | JDAVEN
212 | OTHER | 2011-08-11 | CASTAGNO | C100-RI-040 | C100-CAV-ASSY2 | R3 | 86 |  |  | Strong solvent smell is noticeable in the south end of the Test Lab. |  |  |  | fyi
213 | OTHER | 2011-08-11 | CASTAGNO | C100-RI-044 | C100-CAV-ASSY2 | R3 | 87 |  |  | Strong solvent smell is noticeable in the south end of the Test Lab. |  |  |  | fyi
214 | OTHER | 2011-08-11 | CASTAGNO | C100-RI-042 | C100-CAV-BAKE | R2 | 43 |  |  | Circuit that the bake box was plugged into tripped at 115 deg. C. Old bake box was used for this bake. |  | Restart bake using a different outlet. |  | fyi
215 | OTHER | 2011-08-12 | CASTAGNO | 0 | [5 empty] | Smoke from a local wildfire could be detected inside the Test Lab high bay area as well as inside the production clean room. |  |  |  | fyi
216 | OTHER | 2011-08-11 | CASTAGNO | 0 | [5 empty] | Smoke from a local wildfire could be detected inside the Test Lab high bay area as well as inside the production clean room. |  |  |  | fyi
217 | OTHER | 2011-08-11 | CASTAGNO | 0 | [5 empty] | Strong solvent smell (Epoxy?) is noticeable in the south end of the Test Lab. |  |  |  | fyi
218 | DEVIATIONS | 2011-08-19 | FOREHAND | C100-RI-043 | C100-CAV-TUNE | R4 | 58 | 5 | 5 | HOM filters could not be checked before VTA test due to amplifier problems |  |  |  | no resolution necessary | FOREHAND | 2011-08-19
219 | DEVIATIONS | 2011-08-19 | FOREHAND | C100-RI-044 | C100-CAV-TUNE | R4 | 59 | 5 | 5 | HOM filters could not be checked before VTA test due to amplifier problems |  |  |  | no resolution necessary | FOREHAND | 2011-08-19
220 | OTHER | 2011-08-18 | CASTAGNO | 0 | [5 empty] | Smoke from a local wildfire could be detected inside the Test Lab high bay area as well as inside the production clean room. |  |  |  | fyi
221 | OTHER | 2011-08-19 |  | 0 | [5 empty] | Smoke from a local wildfire could be detected inside the Test Lab high bay area as well as inside the production clean room. |  |  |  | fyi
222 | OTHER | 2011-08-23 | CASTAGNO | C100-RI-043 | [5 empty] | After cavity was qualified in the VTA, cavity was moved to CR. When cable connector was removed from HOM in pos.A (RI-PRO-034) the center conductor detached from the HOM. |  | Remove broken HOM from cavity. Replace HOM with RI-PRO-018.
223 | EP | 2011-07-27 | JDAVEN | C100-RI-059 | C100-CAV-EPOL | R2 | 62 | 2 | 8 | "AJ" Jacobson, Derrick Dail, Jim Follkie, and Tom Goodman did extensive work on the servo motor issue. They discovered several wires shorted out and replaced the amp. |  |  |  | Andrew Burrill finished tying up exposed wires. Wire issue corrected and test cycle successful. | JDAVEN | 2011-07-29
226 | SCRATCH | 2011-09-01 | JDAVEN | C100-RI-070 | C100-CAV-DEGR | R3 | 238 | 1 | 4 | FPC flange is jacked up. | 796F88E9-F5FE-5B1F-5011544E395E1B5D | un-jack |  | fyi
227 | EP | 2011-09-01 | JDAVEN | C100-RI-067 | C100-CAV-EPOL | R2 | 69 | 2 | 8 | Servo Motor would not engage. Every time the program was started the cabinet would power down. Process postponed until servo issue is diagnosed and remedied. |  | Phil Denny, Steve Dutton, Jim Follkie, AJ Jacobson, and Tom Goodman combined forces to diagnose the problem with the servo motor from 01Sep2011-06Sep2011. The amp was swapped out with a new one. The command cable was stripped and checked for any short circuits. The leads for the command cable were soddered inside the control panel. |  | Tom Goodman and AJ Jacobson addressed the issue on 06Sep2011. They swapped the old amp back into the control cabinet. It was stated that "the problem fixed itself". Three separate successful servo tests were performed using the Maintenance servo motor test feature. EP of C100RI067 slated for 07Sep2011. | JDAVEN | 2011-09-06
228 | EP | 2011-09-07 | JDAVEN | C100-RI-067 | C100-CAV-EPOL | R2 | 69 | 2 | 8 | Servo Motor would not engage in process run. Process postponed until facilities personnel can address issue. |  | Servo would not engage in process cycle, but would run in servo motor test feature in maintenance screen. Servo motor caused cabinet to shut-off in two separate test runs. Phil Denny and Tom Goodman addressed the issue from 8:45 until 9:30. No nitrogen was being fed to the cabinet during process initialization. During the previous days attempt to remedy the servo motor issue (D3SeqNum 227), the exhaust pressure gauge control line was inadvertently loosened. This caused the house nitrogen solenoid to shut-off; without the nitrogen supply the process would not allow the servo to engage. |  | The exhaust control line was tightened and the servo motor engaged properly for the process to resume. EP proceeded at 9:37 after a four and a half our delay. | JDAVEN | 2011-09-07
229 | EP | 2011-09-07 | JDAVEN | C100-RI-067 | C100-CAV-EPOL | R2 | 69 | 4 | 19 | Cabinet shut-off during Step 21 Fill & Dump cycle. Power returned and the process was restarted at Step 24 Resistivity Rinse. |  | Assumption is that the servo motor shorted out again, causing the cabinet to power down. There was no apparent release of water into the sump as no visual increase in sump temperature was noted on the LabView. | 0856CCFD-B568-83AA-F40B9B901C1C32D9 | Power returned to cabinet and the process was restarted at Step 24 Resistivity Rinse. Acid change out was previously scheduled for 08Sep2011 prior to the instance so plans were not altered. | JDAVEN | 2011-09-07 | 084FE2F5-EAE6-6B80-D99A201856F043C1
230 | HPR | 2011-09-12 | FOLLKIEJ | C100-RI-039 | C100-CAV-HPR | R2 | 206 |  |  | When I took the cavity out I expected it to be finished with its cycle. It was still running and I'm not sure why? It is possible the start time was posted incorrect? It is possible a third cycle started, but this is believed to not be true. |  | The group met and it was decided to progress the cavity forward to VTA test. I attempted to look inside the cavity for blue strip. With limited viewing it looked like blue stripe was unlikely. |  | Found an incorrect time entered into spreadsheet | FOLLKIEJ | 2011-09-12
231 | LEAK | 2011-09-13 | FISCHER | C100-RI-027 | C100-CAV-INST-HELV | R1 | 24 |  |  | After welding assembly was leak checked. Leak found in the HOM end hv bellows. Repair was performed on sight. Unit re-leak checked.Leak check was good. |  | Leak repaired. Leak check after repair deemed to be good. |  | Unit should have been leak checked after last VTA test. | WORLAND | 2011-09-13
232 | DEVIATIONS | 2011-09-13 | WILCOX | C100-RI-050 | C100-CAV-INST-HELV | R2 | 51 |  |  | Chip in bellows 4840-161 Talked with John Fischer and John Hogan. MUST leak check after cold shock to try and prevent a leaky unit from entering a string. | EBAEFA64-0760-AF65-C366FC589408FB54 | Leak Check Helium Vessel AFTER VTA test. Per John Fischer and John Hogan | E66B0D67-E1BF-9C86-4B27EA6583BCDAF1 | Uploaded data on step 5. Sent J. Fischer, J Hogan, S. Castagnola, and M. McCrea notification via Email | WILCOX
233 | ELNIK_FURNACE | 2011-09-14 | OVERTONR | F061359-11 | C100-CAV-INSP-WGD | R6 | 140 |  |  | wave guide was baked twice |  | second traveler started to capture data |  | visual done on the plating to make sure no blisters/flaking appeared, waveguide okay | GEORGED | 2011-09-15
234 | ELNIK_FURNACE | 2011-09-14 | OVERTONR | F063134-07 | C100-CAV-INSP-WGD | R6 | 136 |  |  | this wave guide was baked twice |  | second traveler started to capture data |  | visual on the plating to make sure no blisters/flaking appeared, waveguide okay | GEORGED | 2011-09-15
235 | EP | 2011-08-11 | JDAVEN | C100-RI-064 | C100-CAV-EPOL | R2 | 66 | 1 | 1 | Pior to installing C100RI064 into the EP carriage on 11Aug11, the copper disk was removed from the end group assembly and thoroughly buffed. Two helicoils were replaced in an attempt to rectify a possible leak noticed during EP of C100RI063. |  |  |  | fyi
236 | EP | 2011-08-11 | JDAVEN | C100-RI-064 | C100-CAV-EPOL | R2 | 66 |  |  | AJ Jacobson replaced the K type thermocouple lines for the "acid in" and "acid out" data. This was done as a possible solution to the issues noticed during C100RI063 with acid out temps readings. |  |  |  | fyi
237 | EP | 2011-09-12 | JDAVEN | C100-RI-068 | C100-CAV-EPOL | R2 | 70 | 3 | 12 | @9:47 - James Davenport heard a "click". He immediately began inspecting cavity and carriage set-up. No evidence of a "click" cause. @9:49 - Noticed an immediate drop in current oscillation/amps. Cell temps also took a dramatic dive, indicating an absence of heat creation. Rongli Geng suggested to jog the process forward w/o attempting to return current. Contact will be inspected once process is done. |  | NOTE: Cavity sat for an additional hour between process ending and initial removal of cavity from carriage due to desire of Rongli Geng to inspect/omh electrical system and contacts. |  | @13:16 - Geng/Davenport omhed out the electrical contact points on the cavity/carriage. Everything fell within process parameter tolerances. Davenport inspected the brake assembly. Nothing off normal at any of the inspected areas. Phil Denny, Kirk Davis, Tom Goodman, and Tony Reilly tabbed to inspect power supply generator. |  | 2011-09-12
238 | EP | 2011-09-09 | JDAVEN | C100-RI-068 | C100-CAV-EPOL | R2 | 70 |  |  | AJ Jacobson, Tom Goodman, and Steve Dutton replaced the old servo power supply/command cables. New cables routed through the front of the maintenance access door and through notices on face of cabinet. Possibility of never unhooking cables to eliminate corrosion build-up on contacts. |  |  |  |  | JDAVEN | 2011-09-09
239 | EP | 2011-09-15 | JDAVEN | C100-RI-068 | C100-CAV-EPOL | R2 | 71 | 1 | 6 | Cavity resistivity test pre-procedure revealed high omh for cavity. Copper Disk omhed 1.0, Cathode rod omhed 1.0, anode leads omhed 1.0, cage omhed 1.0. Cavity/ground straps omhed 3.2. Attempted rotating cage and checking again. Also disassembled brake and rebuffed anode leads. |  |  |  | Jim Follkie noticed corrosion build-up on middle ground strap at cage connection. Moved ground strap an inch over and the resistivity went back to within process parameters. | FOLLKIEJ | 2011-09-15
240 | EP | 2011-09-15 | JDAVEN | C100-RI-068 | C100-CAV-EPOL | R2 | 71 | 5 | 21 | DI-water spray gun was severed during carriage removal from cabinet. |  |  |  | Derrick Dail repaired hose once cavity and hardware were neutralized from carriage assembly. | JDAVEN | 2011-09-15
241 | DEVIATIONS | 2011-09-23 | CASTAGNO | C100-039 | C100-CAV-ASSY2 | R3 | 101 |  |  | VTA test resulted in abnormal FP behavior. |  | Steve C. removed field probe and used white Scotch Brite on the interior of FP flange tube (cavity) to remove residue/staining. Installed field probe 0010-062 and set to 1.091". Cavity was then rerinsed, dried for 30 minutes and assembled.
242 | DEVIATIONS | 2011-09-27 | CASTAGNO | C100-RI-049 | C100-CAV-BAKE | R2 | 55 |  |  | Bake stopped as soon as cavity temp reached 125C. |  | Overtemp switch tripped @ 128C. Replaced with 141C overtemp switch. |  | Restart bake with new overtemp switch. | CASTAGNO | 2011-09-27
243 | EP | 2011-09-21 | JDAVEN | C100-RI-070 | C100-CAV-EPOL | R2 | 73 | 1 | 1 | WLS-TC dead. Process start delayed until replacement is installed. |  |  |  | Process delayed from 4:55-5:45 while James Davenport changed out dead WLS-TC for brand new one. Ran test on InstaCal to verify new unit was operational. Process commenced. | JDAVEN | 2011-09-21
244 | EP | 2011-09-21 | JDAVEN | C100-RI-070 | C100-CAV-EPOL | R2 | 73 | 5 | 21 | Noticed the cathode guide rod was closer than usual during removal process. Steps were taken to prevent guide rod to iris contact. |  |  |  | Checked mounting bracket for any anomaly. Nothing discovered. Only action item that was off-normal was the left end group was rinsed prior to removal of cavity. Additional liquid in the system caused cavity to sag, thereby creating an off-normal level/angle for the guide rod to travel. This was solidified as a fact when the cavity was removed; there was an excess of liquid dumped from the cavity when it was held vertical for low pressure rinsing. James Davenport recommended to Andrew Burrill that the cavity go through an optical inspection to check for potential damage. It was emphasized to NEVER rinse prior to cavity removal again. | JDAVEN | 2001-09-21
245 | OTHER |  | CJOYNER | C100-RI-052 | C100-CAV-ASSY2 | R3 | 105 |  |  | The Cavity foreline throttle valve was open too far .022 valve should have been .002 |  | Due to the valve was not set correctly. Cavity C100-RI-052 was pumped down alittle too fast. I took 1 hour |  | When cavity reached | DREYFUSS | 2011-09-28
246 | ELNIK_FURNACE | 2011-10-05 | OVERTONR | C100RI080 | C100-CAV-HEAT | R1 | 82 |  |  | Temperature ramp up unstable ,stayed at temperature longer then program set for | 4695F349-9218-BF86-3C95F0EDF1DC09A5 | will investigate when furnace is open
247 | ELNIK_FURNACE | 2011-10-05 | OVERTONR | C100RI081 | C100-CAV-HEAT | R1 | 83 |  |  | Temperature ramp up unstable ,stayed at temperature longer then program set fo | 46B666CE-A929-E68E-052BDCC6BEF31654 | will investigate when furnace is open |  | nothing found
248 | OTHER | 2011-10-04 | TMHARRIS | C100-RI-038 | C100-CAV-HPR | R2 | 244 |  |  | Cavity has a smell |  | cavity was hpr |  | cavity passed
249 | OTHER | 2011-10-03 | TMHARRIS | C100-RI-033 | C100-CAV-HPR | R2 | 237 |  |  | cavity interior has an odor |  | cavity was hpr |  | ok
250 | OTHER | 2011-10-05 | TMHARRIS | C100-RI-035 | C100-CAV-HPR | R2 | 240 |  |  | cavity interior has an odor |  | cavity was hpr |  | fyi
251 | LEAK | 2011-09-23 |  | C100-RI-048 | C100-CAV-VTRF | R3 | 99 |  |  | Superfluid leak developed before testing. Cavity tested and met spec but barely |  | ID leak if possible, reprocess |  | multiple cavity cold leaks found on FPC flange using LN2
252 | LEAK | 2011-09-29 |  | C100-RI-049 | C100-CAV-VTRF | R3 | 102 |  |  | Cavity developed at 3e-5 Torr cold leak when passing thru lambda. High power test aborted |  | ID leak, repair and retest
253 | LEAK | 2011-10-01 |  | C100-RI-051 | C100-CAV-VTRF | R3 | 103 |  |  | Cavity developed a superfluid cold leak, test was carried out keeping an eye on vacuum level. Cavity met spec |  | ID leak, store for string assembly
254 | OTHER | 2011-10-06 | JDAVEN | C100-RI-047 | C100-CAV-HPR | R2 | 241 |  |  | This cavity not used for the rebuild of String 4 because it was of the wrong directional orientation. Instead C100RI048 was used.
255 | DEVIATIONS | 2011-10-07 | CASTAGNO | VTA-TH-02 | C100-CAV-LAP-THAT | R1 | 53 |  |  | Tophat was lapped on lapping machine using 360 grit media. No other hand lapping was performed. This will be a test to see if a slightly coarser surface finish will still seal in the VTA.
256 | DETOURS | 2011-10-08 | CASTAGNO | C100-RI-027 | C100-CAV-ASSY | R3 | 88 |  |  | This cavity was removed from string 4 due to HV bellows leak. VTA test is to see how clean the cavity is after string bleedup. |  | Assembled tophat and beamline flanges to the cavity after removal from string. No HPR was performed. |  | failed leak check, test aborted
257 | DEVIATIONS | 2011-10-10 | JDAVEN | C100-RI-057 | C100-CAV-HPR | R2 | 251 | 1 | 1 | James Davenport assisted on the first assembly of the cavity due to Assembly Group being short handed in personnel for the day.
258 | HPR | 2011-10-11 | JDAVEN | C100-RI-049 | C100-CAV-HPR | R2 | 252 | 2 | 5 | I noticed C100RI049 was uncovered outside of the Class 10 tent, sitting in the corner of the Class 100 area. A great deal of work took place around it during this time; the work consisted of two cavities being disassembled and uncaged. C100RI049 was then pushed into the Class 10 uncovered.
259 | DISCREPANCIES | 2011-10-12 | CASTAGNO | C100-RI-056 | C100-CAV-ASSY2 | R3 | 110 |  |  | This cavity took ~48 hours for vacuum to go below the e-6 range. |  | Leak tested cavity after 48 hours of pumping (press=8 e-8 mbar) Cavity is leak tight. |  | Progress with bake and high power VTA qualification test. | CASTAGNO | 2011-10-13
260 | DEVIATIONS | 2011-10-13 | JDAVEN | C100-RI-070 | C100-CAV-DEGR | R3 | 273 | 1 | 1 | C100RI070 had a post-EP optical inspection. Ari Palchezki noticed a stain just outside Iris#1 (below the FPC beam tube flange). James Davenport, assisted by Teena Harris, hydrofluoric acid q-tipped the area to remove the stain. Unaided visual inspection did not produce the presence of the stain after the HF process was administered. |  |  |  | James Davenport ultrasonic cleaned the cavity after the acid process as in standard practice post acid contact with cavities. Further instruction from Andrew Burrill is needed for the next process step for the cavity.
261 | LEAK | 2011-10-07 | ABURRILL | C100-RI-054 | C100-CAV-VTRF | R3 | 107 |  |  | Cavity developed a cold leak during 4K fill. Cavity cooldown was aborted. |  | FPC flange was lapped using modified procedure including the use of scotchbrite. The finish is shinny compared to the standard lapped finish. This may have something to do with the leak. |  | Disassemble and inspect FPC flange. RE-lap and reprocess
262 | SCRATCH | 2011-10-17 | FOREHAND | C100-RI-061 | C100-CAV-TUNE | R4 | 79 | 4 | 4 | Cavity beam tube on the FPC end was scratched during the measurement. Scratch will need to be polished before the cavity can be tested. |  |  |  | scratch polished, cavity proceeded to testing. Eventually passed.
263 | DEVIATIONS | 2011-10-19 | JDAVEN | C100-RI-061 | C100-CAV-HPR | R2 | 263 | 1 | 1 | The normal process flow of allowing 8-10 hours of drying time was not observed. Second HPR followed immediately after same day first HPR and first assembly. |  |  |  | fyi
264 | DEVIATIONS | 2011-10-20 |  | C100-RI-063 | C100-CAV-HPR | R2 | 265 | 1 | 1 | The normal process flow of allowing 8-10 hours of drying time was not observed. Second HPR followed immediately after same day first HPR and first assembly. |  |  |  | fyi
265 | DEVIATIONS | 2011-10-24 | CASTAGNO | VTA-TH-07 | C100-CAV-LAP-THAT | R1 | 65 |  |  | Tophat lapped with lapping machine only. Used 600 grit abrasive powder. Lapping paper was not used.
266 | EP | 2011-11-01 | FOLLKIEJ | C100-RI-079 | C100-CAV-EPOL | R2 | 82 |  |  | At 9:34 am the current dropped from 75 amps to less than 1 amp. The voltage stayed the same at 13.65 volts. The power supply was powered down using Lab View and then powered back up. The current came back briefly then dropped out again. A spark and a "pop" sound was also witnessed coming from the center cell ground strap. |  | The ground strap for the center cell was replaced and normal process resumed. |  | All 5 gound strap were replaced the following day | FOLLKIEJ | 2011-10-27
267 | LEAK | 2011-11-07 | WILCOX | C100-RI-072 | C100-CAV-INST-HELV | R2 | 72 | 3 | 15 | Unit had a -8 range leak on a weld that joins the head 4820-156 to the bellows 4840-171. |  | Sent unit back to weld shop for weld repair and then rechecked. Unit passed leak check | EF6304CE-B361-3AE4-9A04A81BFBEB8A43 |  | WILCOX | 2011-11-07
268 | EP | 2011-11-09 | JDAVEN | C100-RI-083 | C100-CAV-EPOL | R2 | 86 | 4 | 17 | James Davenport had a momentary mental lapse and terminated the polish process just short of the desired 30micron target (29.50micron). |  |  |  | The polish process was re-initiated to allot for the proper material removal. | JDAVEN | 2011-11-09
269 | HPR | 2011-11-10 | JDAVEN | C100-RI-071 | C100-CAV-HPR | R2 | 292 | 1 | 2 | Cavity received 1st assembly an hour after completing 1st rinse. 2nd rinse to place same day. |  |  |  |  | JDAVEN | 2011-11-10
270 | LEAK | 2011-11-14 | DREYFUSS | C100-RI-071 | C100-CAV-ASSY2 | R3 | 128 |  |  | Cavity did not pump down over a two day period found a leak on HOM B gasket | 73DDE5B8-C430-C1D0-7E3C7D9402AC7089 | Turned on turbo to get to 4 Mbar to find leak removed cavity from stand replaced gasket and sent for re-rinse |  | Re-rinse and pump cavity back down | DREYFUSS
271 | HPR | 2011-11-15 | JDAVEN | C100-RI-071 | C100-CAV-HPR | R2 | 295 | 1 | 1 | Production Cleanroom resin bottles had been replaced in the early a.m. This is the first HPR after the clean-up cycle was executed for an hour. |  |  |  | cavity tested fine, admin limited
272 | DEVIATIONS | 2011-11-18 | CASTAGNO | VTA-TH-06 | C100-CAV-LAP-THAT | R1 | 71 |  |  | Tophat was lapped with lapping machine only. 14 micron AlOx. No hand lapping was performed.
273 | DEVIATIONS | 2011-11-21 | CASTAGNO | VTA-TH-08 | C100-CAV-LAP-THAT | R1 | 72 |  |  | Tophat was lapped with lapping machine only. 14 micron AlOx. No hand lapping was performed.
274 | OTHER | 2011-11-29 | FOREHAND | C100-RI-079 | C100-CAV-TUNE | R4 | 97 | 4 | 4 | The no-touch beadpull loading pipe touched the iris of the cavity while being loaded for string insertion. The front of the pipe has a nylon insert, which is the piece that came in contact with the cavity. |  | The cavity continued to receive all the "post welding" measurements. We sent the cavity for an internal inspection before sending it to the next step in the process. |  | The internal inspection revealed no evidence of damage to the cavity at all. The five irises surrounding the iris that was touched were all inspected. Cavity can move through the process as usual. | FOREHAND | 2011-11-29
275 | LEAK | 2011-12-02 | DREYFUSS | F061359/11 | [5 empty] | The window waveguide assembly leaked in clean room |  | Removed window and tried new gasket re-leak checked still leaking |  | remove waveguide from clean room | DREYFUSS | 2011-12-02
276 | LEAK | 2011-12-02 | DREYFUSS | F067764/08 | [5 empty] | The window waveguide assembly leaked in clean room |  | Removed window and tried new gasket re-leak checked still leaking |  | remove waveguide from clean room | DREYFUSS | 2011-12-02
277 | LEAK | 2011-12-02 | DREYFUSS | F356223/04 | [5 empty] | The window waveguide assembly leaked in clean room |  | Removed window and tried new gasket re-leak checked still leaking |  | remove waveguide from clean room | DREYFUSS | 2011-12-02
278 | LEAK | 2011-12-02 | DREYFUSS | F067764/16 | [5 empty] | The window waveguide assembly leaked in clean room |  | Removed window and tried new gasket re-leak checked still leaking |  | remove waveguide from clean room | DREYFUSS | 2011-12-02
279 | OTHER | 2011-12-04 | JDAVEN | C100-RI-079 | C100-CAV-DEGR | R3 | 304 | 1 | 4 | Buffed field probe input flange with white Scotch Brite to remove some residual staining from the EP process. |  |  |  | cavity field emission loaded on first test
280 | HPR | 2011-12-06 |  | C100-RI-079 | C100-CAV-HPR | R2 | 321 | 2 | 5 | The cavity was removed 45 minutes after the official end time due to shortage of personnel. Jim Follkie and Era Perry were occupied with Crane Training. Teena Harris and James Davenport were involved with EP of Ichiro-8. |  |  |  | cavity field emission loaded | TMHARRIS | 2011-12-06
281 | HPR | 2011-12-07 | JDAVEN | C100-RI-080 | C100-CAV-HPR | R2 | 323 | 1 | 1 | Cavity inserted into the HPR while a C100 string was being pushed out of Cleanroom. |  |  |  | cavity field emission loaded
282 | LEAK | 2011-12-08 | MACHA | F061359/11 | [5 empty] | Attempted to recover a waveguide that leaked at the window seal. I polished on the knife edge with wet stones and cratex to repair scratches and imperfections. We leak tested the waveguide again with a window and copper seal. The knife edge seal still leaked. |  | The waveguide is still labeled as do not use and it will be put back into inventory. The next step to try to recover is probably re-machining the knife edge? |  | WG re-machined and placed back in system.
283 | LEAK | 2011-12-09 | DREYFUSS | C100-RI-080 | C100-CAV-ASSY2 | R3 | 139 |  |  | top beamline radial wedge clamp was never tightened down |  | removed from test stand tightened radial wedge clamp and re-rinsed |  | Re-pump down and leak check cavity field emission loaded | DREYFUSS | 2011-12-09
284 | DEVIATIONS | 2011-12-09 | DREYFUSS | C100-RI-080 | C100-CAV-ASSY2 | R3 | 139 |  |  | Extra person in clean room doing a particle study standing over shoulders taking pictures and asking questions while assembly was happening could have contributed by causing some distraction |  | Leak in top beamline |  | Rework | DREYFUSS | 2011-12-09
285 | DEVIATIONS | 2011-12-09 | DREYFUSS | C100-RI-081 | C100-CAV-ASSY | R3 | 115 |  |  | Extra person in clean room doing a particle study standing over shoulders taking pictures and asking questions while assembly was happening. |  |  |  | Cavity passed, no radiation. Cavity on same track as C100-80 which was field emission loaded. | DREYFUSS | 2011-12-09
286 | DEVIATIONS | 2011-12-09 | DREYFUSS | C100-RI-079 | C100-CAV-ASSY2 | R3 | 138 |  |  | Extra person in clean room doing a particle study standing over shoulders taking pictures and asking questions while assembly was happening. |  |  |  | field emission loaded. | DREYFUSS | 2011-12-09
287 | HPR | 2011-12-09 | TMHARRIS | C100-RI-080 | C100-CAV-HPR | R2 | 326 | 1 | 1 | Cavity had to be re-HPRed due to a warm leak in top beam tube connection discovered during initial post-final assembly leak check. |  |  |  | field emission loaded test
288 | LEAK | 2011-12-09 | MACHA | C100-RI-078 | [5 empty] | C100-78H developed a leak during cool-down for VTA test. The cavity pressure data file is attached. The cavity was pulled from D-4 and leak testing with nitrogen cold shocking has begun. | 6C0B2FA4-9B8B-0D0A-BF61C9D6FE5F14A4 | Cold shocked all the metal seal flanges and sprayed with helium. (Top-hat, Bottom Beam-line, Top BL, HOMs, FP and Top-hat adapter seal) No leak detected on first cold shock test. |  | No resolution
289 | HPR | 2011-12-13 | JDAVEN | C100-RI-064 | C100-CAV-HPR | R2 | 331 | 1 | 1 | The cavity was installed into the HPR cabinet 12Dec2011 for some reason by Era Perry (based on logical conclusion from scheduled overtime staff and handwriting in logbook). The bottom beamtube was left exposed to the trickle flow of the HPR wand head all night. |  | James Davenport proceeded as the production schedule recommended after careful inspection of bottom beamtube interior area, flange surface, and exposed HOM hooks. No blue discoloration present. Will contact supervisor and managerial level staff to coordinate a path forward after HPR.
290 | HPR | 2011-12-14 | JDAVEN | C100-RI-079 | C100-CAV-HPR | R2 | 334 | 1 | 1 | Cavity failed Dewar test due to field emissions. Bottom beam tube flange opened up for repeat of final HPR.
291 | HPR | 2011-12-15 | JDAVEN | C100-RI-078 | C100-CAV-HPR | R2 | 337 | 1 | 1 | Cavity failed Dewar Test due to field emission. FCP tophat and HOM beam tube assembly hardware removed to allow for re-HPR. |  | Flanges on cavity re-inspected. Scratches found on fpc flange. AD 1-4-12 | 9C488369-958F-7299-4BE0CD5D72E65B04 | Flange will be reworked in lapping room. AD 1-4-12
292 | CAVITY_COVERING | 2011-12-20 | CJOYNER | C100_RI_078 | C100-CAV-ASSY2 | R3 | 144 |  |  | Before the 2assy. took place in the class 10 area. Cavity C100-RI-078. was flipped in the back-tech uncovered. |  | Once the techs. noticed that the cavity was uncovered.It was covered immediately. |  | Covered C100-RI-078 cavity asap. Then continued with 2nd assy. Cavity Field emission loaded | CJOYNER | 2011-12-15
293 | OTHER | 2011-12-21 | FOREHAND | C100-RI-084 | C100-CAV-ASSY2 | R3 | 147 | 1 | 4 | Original pumpdown of cavity took place with the right angle bleed up valve open. This was not discovered until the leak test was about to take place. The pressure improved by almost an order of magnitude shortly after closing this valve. |  |  |  | n/a | FOREHAND | 2011-12-21
294 | HPR | 2012-01-13 | TMHARRIS | C100-RI-084 | C100-CAV-HPR | R2 | 345 | 1 | 1 | Heavy field emission during VTA test. reHPR for retest. |  |  |  | cavity passed
295 | HPR | 2012-01-18 | JDAVEN | C100-RI-075 | C100-CAV-HPR | R2 | 349 | 2 | 5 | James Davenport neglected to uncover the top beam tube flange when removing flange blanks. |  |  |  | Andrew Burrill advised for all three cavities to reHPR. | JDAVEN | 2012-01-19
296 | HPR | 2012-01-18 | JDAVEN | C100-RI-080 | C100-CAV-HPR | R2 | 350 | 2 | 5 | James Davenport neglected to uncover the top beam tube flange when removing flange blanks. |  | String #9 cavities |  | Andrew Burrill advised for all three cavities to reHPR. | JDAVEN | 2012-01-19
297 | HPR | 2012-01-18 |  | C100-RI-074 | C100-CAV-HPR | R2 | 351 | 2 | 5 | James Davenport neglected to uncover the top beam tube flange when removing flange blanks. |  | String #9 cavities |  | Andrew Burrill advised for all three cavities to reHPR. | JDAVEN | 2012-01-19
298 | HPR | 2012-01-19 |  | C100-RI-075 | C100-CAV-HPR | R2 | 353 | 1 | 1 | Andrew Burrill instructed all three cavities HPR on 18Jan2012 to be reHPR due to James Davenport's error in removing blanks.
299 | HPR | 2012-01-19 | JDAVEN | C100-RI-080 | C100-CAV-HPR | R2 | 354 | 1 | 1 | Andrew Burrill instructed all three cavities HPR on 18Jan2012 to be reHPR due to James Davenport's error in removing blanks. |  | String #9 cavities
300 | HPR | 2012-01-19 | JDAVEN | C100-RI-075 | C100-CAV-HPR | R2 | 353 | 2 | 4 | The DI-water level in the RO tank dropped to 40% due to heavy water usage in the Production Chemroom (both R&D HPR and excessive acid bench chemistry work). |  |  |  | Phil Denny ensured the tank reservoir is acceptable until 15% is reached. Andrew Burrill advised to proceed with HPR as scheduled. | JDAVEN | 2012-01-19
301 | HPR | 2012-01-19 | JDAVEN | C100-RI-074 | C100-CAV-HPR | R2 | 355 | 1 | 1 | Andrew Burrill instructed all three cavities HPR on 18Jan2012 to be reHPR due to James Davenport's error in removing blanks.
302 | HPR | 2012-01-20 | TMHARRIS | C100-RI-076 | C100-CAV-HPR | R2 | 356 | 1 | 1 | Cavity had to be reHPR due to previous three cavities being reHPR and a loss of a day in the schedule.
303 | HPR | 2012-01-24 | JDAVEN | C100-RI-062 | C100-CAV-HPR | R2 | 361 | 2 | 4 | LEWA pump/HPR have been acting inconsistent for nearly 4 weeks. Used this opportunity in an effort to re-quality variables of HPR process. |  | This process was stopped/restarted several times in an effort to adjust all variables of HPR process. Once the baseline was re-established, the process was restarted one final time for uninterrupted run. |  | AJ Jacobson and Derrick Dail re-qualified following HPR variables: wand pressure(should be between 1200-1300psi) and inlet resistivity (>17.90Momphs). | JDAVEN | 2012-01-24
304 | HPR | 2012-01-25 | JDAVEN | C100-RI-009 | C100-CAV-HPR | R2 | 362 | 2 | 4 | Unexplained jump in pressure from 1250psi to upwards of 1700psi between the times of 9:35-9:50. |  |  |  | Derrick Dail adjusted LEWA pump output to reduce increased pressure. James Davenport will monitor rest of the process closely. | JDAVEN | 2012-01-25
305 | HPR | 2012-01-27 | JDAVEN | C100-RI-086 | C100-CAV-HPR | R2 | 363 | 2 | 4 | Continued to experience issues with the LEWA pump pressure. This problem has not been resolved to date, going on 3+ weeks now. Efforts to this point have been stop-gap at best in regaining a operational baseline for the work day at hand and restarting all over the following work day. |  |  |  | previous "repair" of LEWA pump has caused all previous pressure fluctuation issues. Issue addressed and corrected.
306 | DEVIATIONS | 2012-01-27 | FOREHAND | C100-RI-086 | C100-CAV-TUNE | R4 | 105 | 4 | 4 | This cavity had HOM feed throughs and field probes measured and calibrated without having the He vessel welded. The cavity will only be used as a spare or for future R & D. |  | F0 = 1494.013 FP = 0010-096 Tip = 1.090 QextFP = 1.05e12 HOMA = ISI-0082 HOMB = ISI-0083 |  | N/A | FOREHAND | 2012-01-27
307 | LEAK | 2012-02-01 | DREYFUSS | C100-RI-086 | C100-CAV-ASSY2 | R3 | 153 |  |  | cavity has leak in top hat |  | Remove top hat and re-rinse cavity
308 | OTHER | 2012-02-01 | CJOYNER | C100-RI-009 | C100-CAV-BAKE | R2 | 93 |  |  | Cavity bake went off line. |  | Reconnect the data cable to the bake box.Bake program is ramping down. |  | Stopped bake chart and let the bake box, cool down |  | 2012-02-01
309 | OTHER | 2012-03-22 | FOREHAND | C100-RI-039 | C100-CAV-TUNE | R4 | 56 | 1 | 1 | This cavity did not follow normal processing path. Low temp bake took place in elnik furnace prior to HV welding. The cavity was then set up for VTA testing. The results of this test would help determine if all future cavities might be processed in this order. |  | The cavity needed to be re-tuned after low temp baking. The cavity was then set up with field probe and HOM feed-throughs. FP ID = 0010-048 length = 1.082" HOMA = ISI-0016 HOMB = ISI-0018 Tech. = overtonr | 553AC359-D39D-6B69-C179AF67BA95EAFD | Decided not to continue this process with other cavities. No resolution necessary | FOREHAND | 2012-03-22
310 | SCRATCH | 2012-04-25 | ABURRILL | My Head | [5 empty] | test |  | test |  | test

## C100R-D3
Status Report for C100R-D3 on 10-Jun-20
TRAV_SEQ_NUM | D3CLASSIFICATION | D3OPENEDTIME | CATEGORIZETECHNICIAN | SERIALID | TRAVELERID | REVNO | SEQNO | PAGENO | STEPNO | D3DESCRIPTION | DESCRIPTIONFILES | D3ACTIONS | ACTIONFILES | D3RESOLVE | RESOLVETECHNICIAN | D3_RESOLVED_TIME | RESOLVEFILES
1 | OTHER | 2019-02-01 | DREYFUSS | C100R-CAV-ASSY | C100R-CAV-ASSY | R2 | 15 | 1 | 1 | This cavity was fully disassembled and is a re-test due to VTA failure of high field emission |  | re-rinsed cavity in HPR |  | Reassembled cavity for new VTA test | DREYFUSS | 2019-02-01
2 | OTHER | 2019-02-04 | DREYFUSS | C100R-CAV-ASSY | C100R-CAV-ASSY | R2 | 16 | 1 | 1 | This cavity was fully disassembled and is a re-test due to VTA failure of high field emission |  | re-rinsed cavity in HPR |  | Reassembled cavity for new VTA test | FOREHAND | 2019-02-04
3 | LEAK | 2019-02-08 | DREYFUSS | F067764_15 | C100R-CST-ASSY-WGD | R1 | 3 | 1 | 2 | leaker | 0792D6B4-D181-C68F-4DDE875413A8CCF3 | change gasket |  |  | DREYFUSS | 2019-02-05
4 | LEAK | 2019-02-05 | DREYFUSS | F063951_06 | C100R-CST-ASSY-WGD | R1 | 4 | 1 | 2 | leak | 07A99D45-E67C-3C36-23C9A6CF7FDBCF2B | replace gasket |  |  | DREYFUSS | 2019-02-08
5 | LEAK | 2019-02-08 | DREYFUSS | F061359_10 | C100R-CST-ASSY-WGD | R1 | 5 | 1 | 2 | leak | 07CAF035-E421-8928-AAA4BD6D97051792 | replaced gasket
6 | LEAK | 2019-02-07 | DREYFUSS | F061584_04 | C100R-CST-ASSY-WGD | R1 | 6 | 1 | 2 |  | 07E403DE-FADB-F143-AC4052AC5C8F68AA |  |  |  | DREYFUSS | 2019-02-08
7 | LEAK | 2019-02-08 | DREYFUSS | F067764-05 | C100R-CST-ASSY-WGD | R1 | 7 | 1 | 2 | leak | 1321EDBA-09B7-E1CB-D877D545744DDC3F | replaced gasket |  | re-leak checked | DREYFUSS | 2019-02-08
8 | LEAK | 2019-02-12 | DREYFUSS | F067764-15 | C100R-CST-ASSY-WGD | R1 | 3 | 1 | 2 |  | E18D9D71-DD3E-34A8-FE158434E07A728D |  |  |  | DREYFUSS | 2019-02-12
9 | LEAK | 2019-02-12 | DREYFUSS | F067764-05 | C100R-CST-ASSY-WGD | R1 | 7 | 1 | 2 |  | E1A32D9D-B699-FF7B-8FB794C09F45F346 |  |  |  | DREYFUSS | 2019-02-12
10 | LEAK | 2019-02-12 | DREYFUSS | F061359-10 | C100R-CST-ASSY-WGD | R1 | 5 | 1 | 2 |  | E1AE8977-0E31-B92F-BE30AC1D90A7A62B |  |  |  | DREYFUSS | 2019-02-12
11 | LEAK | 2019-02-12 | DREYFUSS | F061584-04 | C100R-CST-ASSY-WGD | R1 | 6 | 1 | 2 |  | E1BA3786-BD64-ADCE-428210C9C1170511 |  |  |  | DREYFUSS | 2019-02-12
12 | LEAK | 2019-02-13 | DREYFUSS | F898271_01 | C100R-CST-ASSY-WGD | R1 | 9 | 1 | 2 |  | 13914A97-BB9A-A6EC-C2AAD686CF7133F7 |  |  |  | DREYFUSS | 2019-02-13
13 | LEAK | 2019-02-15 | DREYFUSS | F890195_03 | C100R-CST-ASSY-WGD | R1 | 11 | 1 | 2 | failed leak check with cyromodule group torque pattern up to 120 inch pounds | 6F6B2ED8-FA8B-57D0-598ED7BB11623681 |  |  |  | DREYFUSS | 2019-02-15
14 | LEAK | 2019-02-21 | FOREHAND | G096314/03 | C100R-CST-ASSY-WGD | R1 | 16 | 1 | 2 |  | A2FFBD9C-D15F-BE55-3B69D49D45D316B8 |  |  |  | DREYFUSS | 2019-02-21
15 | LEAK | 2019-02-21 | DREYFUSS | F890195/03 | C100R-CST-ASSY-WGD | R1 | 11 | 1 | 2 |  | A43670FC-FB66-513A-F44C68E71BC045A8 |  |  |  | DREYFUSS | 2019-02-21
16 | LEAK | 2019-02-21 | DREYFUSS | F897109/01 | C100R-CST-ASSY-WGD | R1 | 14 | 1 | 2 | [5 empty] | DREYFUSS | 2019-02-16
17 | DEVIATIONS | 2019-02-21 | DREYFUSS | F898271/01 | C100R-CST-ASSY-WGD | R1 | 10 | 1 | 2 | This cavity had only 10 bolts in flange and was leak toght we added the other 8 bolts and torqued according to meeting notes and procedure and window stayed leak tight |  |  |  |  | DREYFUSS | 2019-02-21
18 | LEAK | 2019-02-22 | DREYFUSS | F063952-02 | C100R-CST-ASSY-WGD | R1 | 8 | 1 | 2 | [5 empty] | DREYFUSS | 2019-02-22
19 | LEAK | 2019-02-22 | DREYFUSS | F061359-10 | C100R-CST-ASSY-WGD | R1 | 12 | 1 | 2 |  | DB8B848B-AC20-A20B-35D21C87E43CBF8C |  |  |  | DREYFUSS | 2019-02-22
20 | LEAK | 2019-02-22 |  | F890195/03 | C100R-CST-ASSY-WGD | R1 | 13 | 1 | 2 |  | DB998AAA-DFD3-9F51-8C0A33121804F884 |  |  |  | DREYFUSS | 2019-02-22
21 | LEAK | 2019-02-22 | DREYFUSS | F890195/03 | C100R-CST-ASSY-WGD | R1 | 18 | 1 | 2 | leaked first time didnt leak with 10 bolts in then re-leaked after adding 8 bolts in then leaked again after disassembly and reassembly with all bolts and alcohol | DBF941A4-D34E-59DD-0DF9ADF978803190 |  |  |  | DREYFUSS | 2019-02-22
22 | OTHER | 2019-02-26 | FOREHAND | C100R-01 | C100R-CST-ASSY | R1 | 1 | 2 | 5 | Due to pump configuration in the clean room, the cavity is being assembled from downstream(cav8) to upstream (cav1) instead of as stated in traveler. |  | There will be several differences between this assembly and how the traveler states to complete each task. Notes will be taken during the assembly of the string in order to update the traveler to match the assembly in the new facility and with updated techniques. |  | No resolution necessary. This is for information only. | FOREHAND | 2019-02-26
23 | OTHER | 2019-02-26 | FOREHAND | C100R-01 | C100R-CST-ASSY | R1 | 1 | 6 | 17 | The VAT valve will be installed using a single lollipop adapted with a holder to place the valve at 45 degrees. Realized there was not enough clearance to put backing rings in place after the cavity and valve were moved together. |  | Cavity had to be pulled back from the valve enough to put both backing plates in place. After this was complete we installed four nuts and snugged them by hand. We now realized there was not enough room for two of the nuts to be started onto the studs. Called manager in to discuss all options. |  | Came to the conclusion that we would have to back the cavity away again to allow the nuts to clear the HOM flange interference. The cavity was held firmly in place while sliding the backing plates off enough to complete this task. The cavity was then pulled back and the nuts were installed. | FOREHAND | 2019-10-22
24 | DISCREPANCIES | 2019-03-01 |  | C100R-01 | C100R-CST-ASSY | R1 | 1 | 4 | 11 | Wave guide F061534_03 is the wave guide with the serpentine gasket imprint on the small flange |  |  |  |  | DREYFUSS | 2019-03-01
25 | OTHER | 2019-03-01 | DREYFUSS | C100R-01 | C100R-CST-ASSY | R1 | 1 | 4 | 9 | found sliver of metal around bolt hole before installing gasket-ed blank | 3EEEB2EA-EDC6-AF7C-43F365DE9D91CFD3 | Called Kirk Davis |  | Removed sliver from flange and resumed assembly | DREYFUSS | 2019-03-01
26 | DISCREPANCIES | 2019-03-01 | DREYFUSS | C100R-01 | C100R-CST-ASSY | R1 | 1 | 3 | 6 | found water in the helium vessel | 49F6CFB8-D0EE-E3BF-C2F1D5A8960ED1E8 |  |  |  | FOREHAND | 2019-03-01
27 | DEVIATIONS | 2019-02-26 | ASHLEYA | C100-RI-039 | C100R-CAV-CHEM-HPR | R1 | 24 |  |  | Tape residue found on cavity. |  | Removed with acetone and red bag isopropyl wipes. |  | No more tape residue present. | ASHLEYA

## F100-D3
Status Report for F100-D3 on 10-Jun-20
TRAV_SEQ_NUM | D3CLASSIFICATION | D3OPENEDTIME | CATEGORIZETECHNICIAN | SERIALID | TRAVELERID | REVNO | SEQNO | PAGENO | STEPNO | D3DESCRIPTION | DESCRIPTIONFILES | D3ACTIONS | ACTIONFILES | D3RESOLVE | RESOLVETECHNICIAN | D3_RESOLVED_TIME | RESOLVEFILES
1 | DETOURS | 2012-03-02 | ROYB | TEST | F100-CM-DISA | R1 | 1 |  |  | TEST |  | TEST |  | TEST | BENNETTR | 2012-03-02
2 | OTHER | 2013-06-26 | FOREHAND | 0 | F100-CAV-BAKE | R1 | 3 | 1 | 1 | Cavity bake box number two failed to reach the temperature of 120C several times. we abandoned the run and moved the cavity to bakle box number 1. Due to lost time and schedule restraints, we cut the run from 24hrs to 18hrs. |  | No glitches in bake box 1. |  | n/a | FOREHAND | 2013-06-26
3 | DEVIATIONS | 2013-07-01 | FOREHAND | JL016 | F100-CAV-BAKE | R1 | 4 | 1 | 1 | Cavity isolation valve was closed during the bake. Valve was thought to be open until discovered closed after bake was complete. |  | Slowly opened valve while still pumping on the connection hose and watched the pressure to see how high it would go. The pressure never climbed above the mid -4 torr range. |  | Moving forward without any action due to the fact that the pressure was still reasonable. | FOREHAND | 2013-07-01
4 | DETOURS | 2013-07-29 | CJOYNER | 0F100HOM04 | F100-CAV-ASSY | R1 | 7 | 1 |  | Cavity f-100 the hom B ... bead wield broke |  | I was taken out of clean room and replaced |  | Another hom was sent back into the clean room | CJOYNER | 2013-07-29
5 | DETOURS | 2013-09-17 | FOREHAND | JL-012 | [5 empty] | Cavity failed warm dewar leak check after baking in bake box.Attached file details leak investigation | C821ABE7-CD05-B59F-C38E3DD995C10906 | cavity was sent back to clean room for dis assembly, visual inspection, HPR and re-assembly |  | see above | FOREHAND | 2013-09-17
6 | DISCREPANCIES | 2013-09-17 | FOREHAND | JL-012 | [5 empty] | Field probe flange leaked after assembly. |  | Cavity was removed from clean room for flange repair. | C95F9459-0D98-2E52-5F1E921C0E1CE987 | see above. | FOREHAND | 2013-09-17
7 | DETOURS | 2013-10-01 | DREYFUSS | JL013 | F100-CAV-ASSY2 | R1 | 20 |  |  | for this test cavity was put together with old beamline blanks and a top hat was also assembled on the beamline with 316L harware and the belvelle washers where installed on the nut end not the bolt head end, also the copper spacer gasket was used on the top hat to the FPC flange.
8 | DETOURS | 2013-10-21 | FOREHAND | JL012 | F100-CAV-RFIN | R1 | 2 | 2 | 2 | The field probe flange on JL-012 was machined-due to surface issues- earlier in production. The machining changed the depth, parallelism, ID and OD of the seal groove on the flange. After failed attempts to get the cavity to seal with the standard aluminum gasket, we decided to use a stainless steel ring with indium on both sides. |  | We made 316L rings to fit the seal groove of the field probe flange with the intention of sealing them with 0.025" indium on both sides. This requires re-setting the field probe tip in order to maintain the same Qext of 1E12. |  | Field probe feed-through was installed on cavity with indium on both sides of the SS ring and the flange was torqued to 40 in/lbs. The calibrated RF measurement showed the Qext to be at approx. 9.44e11 with the tip length set at 1.105". This is considered to be the final setting for the field probe tip unless the VTA testing results do not coincide with these results. | FOREHAND | 2013-10-21
9 | DETOURS | 2013-11-05 | DREYFUSS | JL012 | F100-CAV-ASSY | R1 | 16 |  |  | Field probe flange was found to be out of tolerance and would not seal |  | Engineers came up with a steel ring and .025 indium on both sides of the ring |  | Ring and indium where used for assembly cavity was now leak tight. | DREYFUSS | 2013-11-05
10 | DETOURS | 2013-11-05 | DREYFUSS | JL015 | F100-CAV-ASSY | R1 | 18 |  |  | Field probe and HOM flanges where found to be out of tolerance and would not seal |  | Engineers came up with a steel ring and .025 indium on both sides of the ring |  | Rings and indium where used for assembly cavity was now leak tight. | DREYFUSS | 2013-11-05
11 | DETOURS | 2013-11-05 | DREYFUSS | JL011 | F100-CAV-ASSY | R1 | 17 |  |  | White milky stains in beam tube both ends took pics sent to Steve Castagnola.Cavity was also left in walk way not in a cubicle to dry over night unknown foot traffic past cavity at night. Also cavity was pumped on thru the night with a leak in the field probe re-torqued field probe for third time got a quarter turn on one bolt no turn on other seven bolts pressure started dropping immediately. Also found pit marks and what looks like scratches in beam tube field probe hom side. |  | Cavity was tested any ways and failed |  | Cavity was taken apart and is now getting an internal inspection | DREYFUSS | 2013-11-05
12 | DEVIATIONS | 2013-11-11 | FOREHAND | JL011 | F100-CAV-ASSY | R1 | 19 | 2 | 3 | We were not ready to install tophat before the cavity had to go in the rinse cabinet. |  | Tophat will be installed during the final assembly instead of first assembly. |  | Tophat will be installed during the final assembly instead of first assembly. | FOREHAND | 2013-11-11
13 | DISCREPANCIES | 2013-12-03 | DREYFUSS | F100_01 | F100-CST-ASSY-ASO1 | R1 | 2 |  |  | this is a list of discrepancies that where found on the cavities before string assembly. JL015: nothing found JL016: FPC beamline some pitting and light stains JL013: light stains on both beamlines and FPC flange JL012: nothing found JL010:light staining (foggy white)on FPC JL014: stains on FPC flang and beamline on FPC end JL017: stains on FPC flange and beamline FPC end | A300BB0A-ACD1-0EE8-6B4DC149435F541D | meeting happened took pictures | A3116626-CE70-3B2C-1C69F4AEC2C14BF7 | put string together | DREYFUSS | 2013-12-03 | A31F64F4-FB25-AF6B-0F0EBC05759E4D5E
14 | DISCREPANCIES | 2013-12-12 | DREYFUSS | F100_01 | F100-CST-ASSY-ASO1 | R1 | 2 |  |  | JL0014 and JL016 connection was uncovered with gasket put up to open beam line gasket did not fit so cavities re-covered then uncovered again to test new gasket the re-covered until uncovered again to do assembly with new gasket. |  | find a gasket with a smaller diameter to fit the flange |  | found a gasket with a smaller diameter to fit the flange installed gasket and assembled cavities together. | FOREHAND | 2013-12-03
15 | DEVIATIONS | 2013-12-12 | DREYFUSS | F100_01 | F100-CST-ASSY-ASO1 | R1 | 2 |  |  | Gasket 2A was switched with 4B so 2A should be where 4B is supposed to be. |  |  |  | put string together no affect on string | FOREHAND | 2013-12-06

## R100-CAV-D3
Status Report for R100-CAV-D3 on 10-Jun-20
TRAV_SEQ_NUM | D3CLASSIFICATION | D3OPENEDTIME | CATEGORIZETECHNICIAN | SERIALID | TRAVELERID | REVNO | SEQNO | PAGENO | STEPNO | D3DESCRIPTION | DESCRIPTIONFILES | D3ACTIONS | ACTIONFILES | D3RESOLVE | RESOLVETECHNICIAN | D3_RESOLVED_TIME | RESOLVEFILES
1 | DEVIATION | 2010-07-12 | jdaven | R100_02 | R100-CAV-DEGR | R1 | 9 | 5 | 24 | R100-02 did not progress directly to the HPR step because Roland Overton and Aaron DeKerlegand needed to collect some data on 13-Jul-2010 |  | Cavity wasn't HPR immediately after the Post-EP degrease. |  | Once Tuning and CMM data are collected cavity will be returned to Chemistry for another degrease and the HPR. | jdaven | 2010-07-12
2 | DETOUR | 2010-07-14 | dreyfuss | R100_02 | R100-CAV-DEGR | R1 | 10 |  |  | Black film on top hat coming off on gloves appears to be a film from the welder |  | re acid etched whole top hat and relapped flange |  | put top hat onto cavity | dreyfuss | 2010-07-14
3 | DEVIATION | 2010-07-14 | jdaven | R100_03 | R100-CAV-DEGR | R1 | 12 | 1 | 9 | R100-03 was delivered a full day late according to the schedule. CMM and lapping did not meet the deadlines for deliver to Chemistry group. |  | The cavity was forced dried with Methanol and Nitrogen blowing to expedite the process in order to stay on schedule. Rongli Geng verified this method as a valid solution citing several occurrences of the practice for ILC EPs. Nitrogen flow continued for two hours through cavity. |  | R100-03 is dry and ready for EP setup. | jdaven | 2010-07-14
4 | DEVIATION | 2010-06-09 | jdaven | R100_01 | R100-CAV-EPOL | R1 | 1 | 5 | 24 | Peter Kneisel ignored the schedule and occupied the 100gal Ultrasonic causing R100-01 to sit for a half hour between EP disassembly and UHV cleaning. |  | Cavity didn't go directly into ultrasonic after removal from EP machine. |  | Hopefully in the future the schedule will be observed in order to expedite the work day for EP technicians. | jdaven | 2010-06-09
5 | DEVIATION | 2010-07-12 | jdaven | R100_02 | R100-CAV-EPOL | R1 | 2 | 2 | 8 | EP Process start time delayed due to several issues: 1. Chiller maintenance technicians from 09July2010 neglected to power the system back up. 2. LabView Computer had a failed RAM stick requiring David Bliss to address and resolve. 3. Servo Motor wouldn't engage to rotate cavity and then appeared to reverse voltage. |  | 1. Derrick Dail returned power back to the Chiller System. 2. David Bliss diagnosed the malfunctioning RAM stick and removed it. 3.Aaron Jacobson checked the Servo Motor. James Davenport cut power off to the cabinet, disconnected the power supply cables, and went through an alternating round of attempts to engage motor. |  | EP process delayed by a total of 5 hours. 1. Chiller at normal operation. 2. David Bliss ordered a replacement RAM stick; awaiting delivery for installation. 3. No real resolution. Replacements are on order for the motor and both power cables. | jdaven | 2010-07-12
6 | DEVIATION | 2010-07-15 | jdaven | R100_03 | R100-CAV-EPOL | R1 | 3 | 2 | 8 | Servo Motor wouldn't engage to rotate cavity. |  | James Davenport cut power off to the cabinet, disconnected the power supply cables, and went through an alternating round of attempts to engage motor. |  | EP process delayed by 30 minutes. Replacement parts still on order. | jdaven | 2010-07-15
7 | DEVIATION | 2010-07-27 | castagno | R100_04 | R100-CAV-ASSY2 | R1 | 4 | 1 | 2 | While tightening the fasteners that retain the valve, burst disc, and cavity beampipe extension, one of the studs broke. |  | Due to physical limitations of the flange configuration, the broken stud could not be removed and replaced. All fasteners associated with the joint had to be loosened and the valve/burst disc removed from the extension tube/cavity assy. The stud was replaced and the sub assy was cleaned with N2 and installed with new gaskets to the cavity beam pipe extension in the class 10 area. |  | Progress cavity as per normal processing. | castagno | 2010-07-27
8 | DEVIATION | 2010-07-15 | jdaven | R100_03 | R100-CAV-EPOL | R1 | 3 | 5 | 24 | R100-03 did not progress directly to the HPR step because Roland Overton and Aaron DeKerlegand needed to collect some data on 15-Jul-2010. |  | Cavity wasn't HPR immediately after the Post-EP degrease. |  | Once Tuning and CMM data are collected cavity will be returned to Chemistry for another degrease and the HPR. | jdaven | 2010-07-15
9 | DEVIATION | 2010-07-23 | jdaven | R100_04 | R100-CAV-EPOL | R1 | 4 | 5 | 24 | R100-04 did not progress directly to the HPR step because Danny Forehand and Aaron DeKerlegand needed to collect some data on 24-Jul-2010 |  | Cavity wasn't HPR immediately after the Post-EP degrease. |  | Once Tuning and CMM data are collected cavity will be returned to Chemistry for another degrease and the HPR. | jdaven | 2010-07-23
10 | DEVIATION | 2010-07-29 | jdaven | R100_05 | R100-CAV-EPOL | R1 | 5 | 5 | 24 | R100-05 did not progress directly to the HPR step because Roland Overton and Aaron DeKerlegand needed to collect some data on 30-Jul-2010 |  | Cavity wasn't HPR immediately after the Post-EP degrease. |  | Once Tuning and CMM data are collected cavity will be returned to Chemistry for another degrease and the HPR. | jdaven | 2010-07-29
11 | DISCREPANCY | 2010-08-01 | castagno | R100_05 | R100-CAV-ASSY2 | R1 |  |  | 4 | Strange odor coming from exhuast port of rough pump @ ~10 mbar. |  | None |  | Continue normal cavity processing steps. | castagno | 2010-08-01
12 | DEVIATION | 2010-07-31 | jdaven | R100_05 | R100-CAV-HPR | R1 | 11 | 1 | 5 | Rongli Geng wanted to reHPR Ichiro 7 due to a leak during pump down. He removed R100-05 from the HPR. |  | It is unclear if Rongli covered the cavity before removal. The cavity was left in the main assembly area by the double doors. |  | Any field emission during VTA test can be directly linked to this break from protocol. | jdaven | 2010-07-31
13 | DEVIATION | 2010-08-02 | jdaven | R100_06 | R100-CAV-HPR | R1 | 13 | 1 | 5 | R100-06 could not be installed into Class 10 Area because Rongli Geng was occupying it with PKU-3 (was awaiting final assembly). |  | PKU-3 was not originally on the planning schedule. It was forced into the schedule at the 8:30 meeting. |  | R100-06 placed near front of Main Assembly room on a lift cart uncovered to dry over night. | jdaven | 2010-08-02
14 | LEAK | 2010-08-04 | castagno | R100_06 | R100-CAV-ASSY2 | R1 | 6 | 1 | 4 | Total cavity pressure would not go below 2.9e-5 mbar after four hours on the turbo pump. |  |  | D2F3841A-D424-5D02-6F0156BA67C0F746 |  | castagno | 2010-09-02
15 | MISTAKE | 2010-05-12 | stipton | R100_03 | R100-CAV-THKN | R1 | 15 | 1 | 13 | WHEN TAKING THE MEASUREMENTS A POINT WAS MISSED. THEREFORE ONLY 29POINTS WERE TAKEN INSTEAD OF 30. THE DATA HAS BEEN EDITED FOR ERRORS |  | EDITED THE DATA FOR ERRORS |  |  | stipton | 2010-05-14
16 | DEVIATION | 2010-08-11 | tmharris | R100_08 | R100-CAV-DEGR | R1 | 16 | 1 | 1 | R100-08 was delivered a full day late according to the schedule. CMM and lapping did not meet the deadlines for delivery to Chemistry group. |  | The cavity was forced dried with Methanol and Nitrogen blowing to expedite the process in order to stay on schedule. Rongli Geng verified this method as a valid solution citing several occurrences of the practice for ILC EPs. Nitrogen flow continued for two hours through cavity. |  | R100-08 is dry and ready for EP setup. | tmharris | 2010-08-11
17 | DEVIATION | 2010-08-12 | jdaven | R100_08 | R100-CAV-HPR | R1 | 19 | 1 | 5 | R100-08 was only HPR'ed without assembles due to scheduled Power Shutdown for the removal of Test Lab south wall. It will be processed when power outage is over. |  | Cavity was rinsed and set over night for drying. |  | Cavity blanked off and removed from cleanroom for storage/tuning/rf-frequency check/thickness measurements during Shutdown. | jdaven | 2010-08-12
18 | OTHER | 2010-09-01 | castagno | R100_01 | R100-CAV-ASSY | R1 | 10 | 1 | 2 | Some kind of dirt/loose particles were observed in the seal groove of the FPC beam line flange. |  | Cavity was oriented with the FPC beam line flange down and parallel to the floor. The seal groove was wiped with and IPA soaked wiper. Loose particles and a black/gray residue was left behind on the wiper. This process was repeated again. No residue was observed on the IPA soaked wiper after the second cleaning. The aformentioned process was repeated on the other beam line flange but no residue or loose particles were observed after wiping. |  | Continue with cavity assembly. Perhaps we should perform a solvent wipe of each flange prior to 1st HPR? | castagno | 2010-09-01
19 | OTHER | 2010-09-07 | castagno | R100_06 | [5 empty] | When this cavity was being disassembled after VTA testing, the center conductor came off of the HOM feedthru (S/N 47726-0005)when the right angle adapter was removed. SC |  |  |  | replace HOM feed-thru |  | 2010-11-19
20 | SLOW_LEAK | 2010-09-09 | grenoble | R100_08 | R100-CAV-VTRF | R1 | 9 |  |  | Superfluid leak in R100-8 cavity. |  |  |  | FPC flange re-lapped with original 9 um paper, No leak |  | 2010-11-19
21 | RF_TEST_ABORTED | 2010-09-09 | grenoble | R100_08 | R100-CAV-VTRF R1 | R1 | 9 | 5 |  | R100-8 cavity has a super fluid leak. Unable to fully test cavity. Field emitter discovered during high power RF test in VTA. |  | Remove lower flange from cavity and perform another HPR in main clean room. Retorque all fasteners to try to eliminate the cold leak.
22 | DEVIATION | 2010-09-15 |  |  | R100-CAV-VTRF | R1 | 9
23 | BCP | 2010-09-20 | KDAVIS | R100_7 | R100-CAV-VTRF | R1 | 16 |  |  | R100_7b Vertical Test could not be performed using production RF system due to standing wave misbehavior. Input coupling was low using production system 4e9, Q0 was high using production system 7e10. |  | Performed test using outboard 250W RF amplifier. |  | Cavity Q0 is more believable using outboard amp (2e10). | KDAVIS | 2010-09-20
24 | OTHER | 2010-09-21 | castagno | R100_08 | R100-CAV-ASSY2 | R1 | 24 |  |  | Valve at bottom of cavity was not opened prior to the slow pumpdown. This resulted in the cavity not being evacuated during the slow pumpdown. The cavity was also not under vacuum during the VTA test. |  | Bleed up the test stand. Open cavity valve and reinitialize slow pump down procedure. Resume travelers and retest in VTA.
25 | DEVIATION | 2010-08-02 | jdaven | R100_06 | R100-CAV-EPOL | R1 | 25 | 3 | 12 | COMEDY OF ERRORS: 1. James neglected to hook-up cavity cooling supply hose to system 2. After restarting process after D3#1 was corrected, James neglected to turn tank temp control back on resulting in a drastic spike in sump, CWin, CWout, AcidIN, & AcidOUT temps. 3. Cooling system running to hot then too cold, then no incoming supply. |  | 1. Shut off process in order to attach supply hose @ 8:10. 2. Shut off current @8:45 in order to allow for sump to return to lower temp. 3. Chased the desired temp for 80 minutes. |  | All three variables resulted in nearly 2 hours of lost efficient work time in process. The cooling system supply temp is checked at the wall connect prior to engaging the system now. | jdaven | 2010-08-02
26 | BCP | 2010-09-30 | overtonr | R100_08 | R100-CAV-VTRF | R1 | 18 |  |  | When preparing to remove test stand from dewar found rough valve on test stand closed. Cavity and turbo valves were open and vacuum was in 7 range |  | Need to ensure a checklist is attached to the top plate for vacuum operations.
27 | MISTAKE | 2010-10-18 |  | R100_06 | R100-CAV-INSP-ADPT | R1 | 14 | 3 | 7 | Failed to connect thermocouples to waveguide flange
28 | MISTAKE | 2010-10-18 |  | 11 | R100-CAV-INSP-ADPT | R1 | 16 | 3 | 7 | Failed to attach thermocouples to waveguide flange.
29 | MISTAKE | 2010-10-18 |  | 16 | R100-CAV-INSP-ADPT | R1 | 19 | 3 | 7 | Failed to attach thermocouples to waveguide flange.
30 | DISCREPANCY | 2010-10-26 | castagno | R100_02 | [5 empty] | Today this cavity had the FPC flange relapped after helium ves. welding and CMM inspection but prior to HPR and CR assy. The FPC flange was lapped with the new polishing cloth. I noticed some areas of the flange appeared to have material embedded into the surface of the flange or small pitting. No scratches were observed. |  | Embedded particles - a possible anomaly of the new polishing cloth? Not sure if this is a good thing or a bad thing. |  | Proceed with HPR and clean room assy. | castagno
31 | LEAK | 2010-11-01 | kdavis | R100_02 | R100-CAV-VTRF | R1 | 21 |  |  | Cavity vacuum was unstable during cooldown, noted a correlation between dewar VAT valve closed and cavity vacuum deterioration from 1e-7 range to 1e-4 range. No He breakdown noted during RF test. |  | cryogenically test test-stand inserts to check for leaks. |  | TS 28-008 and TS28-009 where both tested, TS28-009 was found to have a small leak in the pump-out line. |  | 2010-11-19
32 | DEVIATION | 2010-11-04 | dreyfuss | R100_04 | R100-CAV-ASSY2 | R1 | 17 |  |  | test stand did not have pump down system on it use ILC slow pump down system pump down rate not the same as normal system so pump down happened alot quicker then normal. |  |  |  | went with it will find out on test weather it makes a difference or not | dreyfuss | 2010-11-04
33 | HPR | 2010-11-03 | stipton | R100_04 | R100-CAV-HPR | R1 | 34 |  |  | the wand head ont he HPR was changed to the R&D wand head |  | the wand head ont he HPR was changed to the R&D wand head due to the normal wand head not aligning properly and having to be remachined |  |  | stipton | 2010-11-03
34 | HPR | 2010-11-04 | stipton | R100_08 | R100-CAV-HPR | R1 | 35 |  |  | HPR wand head change prior to run |  | the HPR wand head had been changed out prior to the run for a R&D cavity. It was and then changed back and a clean up run was ran before the final R100-008 rinse was started. This D3 is a waste of time due to the fact that the HPR is made to have the wand head changed and there is always a clean up run done after a wand head change, therefore nothing is deviated from the normal run for the R100-008 cavity. |  |  | stipton
35 | HPR | 2010-11-06 | jdaven | R100_01 | R100-CAV-HPR | R1 | 37 | 1 | 3 | One of the support "feet" bolts was not engaged at all with the surface of the FPC strike plate. |  |  |  | The bolt in question was tightened to provide support to the FPC strike plate. HPR setup continued as normal. | jdaven | 2010-11-06
36 | DISCREPANCY | 2010-11-08 | dreyfuss | R100_04 | R100-CAV-ASSY2 | R1 | 17 |  |  | cavity was slow pumped by ILC manifold and was pumped down faster then the normal method. |  |  |  | no resolve | dreyfuss | 2010-11-08
37 | DEVIATION | 2010-11-09 | forehand | R100_007 | R100-CAV-INST-HELV | R1 | 8 | 2 | 11 | Skipped RF measurement before proceeding with HV welding |  | No action necessary |  | No action necessary | forehand | 2010-11-09
38 | DEVIATION | 2010-11-09 | forehand | R100_006 | R100-CAV-INST-HELV | R1 | 9 | 2 | 11 | Skipped RF measurement before proceeding with HV welding |  | No actions necessary |  | No actions necessary | forehand | 2010-11-09
39 | HPR | 2010-11-10 | tmharris | R100_04 | R100-CAV-HPR | R1 | 40 | 1 | 1 | re HPR due to poor performance |  | the cavity was taken out of the VTA and brought into the cleanroom to be reHPRd to see if it would perform better. |  | Cavity was EP'd, 5um, then sent for 2 HPR passes. Re-test to come. |  | 2010-11-19
40 | DISCREPANCY | 2010-11-11 | castagno | R100_04 | R100-CAV-ASSY2 | R1 | 21 |  |  | During slow pump down, an unusual odor was noticed immediately after the block valve was opened on the foreline circuit.
41 | HPR | 2010-11-12 | jdaven | R100_05 | R100-CAV-HPR | R1 | 45 | 1 | 1 | HPR wand head change out. Peter's 3.5 Cell had been rinsed prior, so the smaller head was in place. |  |  |  | Removed the smaller wand head and replaced with the regular wand head. | jdaven | 2010-11-12
42 | BCP | 2010-11-12 | jdaven | R100_05 | R100-CAV-HPR | R1 | 39 | 1 | 5 | Peter Kneisel moved his 3.5 cell cavity from the small Class 10 area into the larger class 10 area. He did so to make room for his unscheduled drying of a half cell cavity in the small Class 10 area. |  | Communicate |  | R100-05 could not progress directly into the Class 10 due to the presence of unscheduled work. Cavity remained covered and sat in Class 10 room.
43 | LEAK | 2010-11-16 | castagno | R100_05 | R100-CAV-ASSY2 | R1 | 23 |  |  | After pumping overnight, pressure was in the mid five range. Leak testing shows that HOM B is leaking. |  | The cavity was bled up and removed from test stand. Lower beam line and HOM B were removed and inspected. Nothing unusual found. Leak tested The HOM and found a leak in the braze between the copper body and SS flange. |  | Re-HPR the cavity. Remove the leaking HOM from service (47726-016) and replace with an acceptable substitute (47726-47) after cavity has dried. Assemble remainder of cavity and progress with leak test and VTA testing. | castagno | 2010-11-17
44 | RF_TEST_FAILED | 2010-11-16 | kdavis | R100_04 | R100-CAV-VTRF | R1 | 28 |  |  | R100_04 has an open in the FP cable, probably close to the dewar top plate. All data associated with this RF test have unknown error bars due to field probe Qext unknown. I scaled Eacc numbers to previous test results.
45 | RF_TEST_FAILED | 2010-11-17 | kdavis | R100_08 | R100-CAV-VTRF | R1 | 29 |  |  | previous test on R100_08 in D3 showed input coupler heating at 4 MV/m. This retest in D4 was good, suspect bad level sensor in D3. |  | Check level sensor on Dewar 3 |  | Level sensor channel damaged, heater crushed. Repair carried out, re-qual of dewar to come. |  | 2010-11-19
46 | DEVIATION | 2010-11-18 | dreyfuss | R100-05 | R100-CAV-ASSY2 | R1 | 24 |  |  | cavity was moved to landing was put back on turbo but not on scroll pump cavity did pump on turbo all night |  | opened scroll pump block valve and pressure came down |  | cavity appears to be fine and leak checked good | dreyfuss | 2010-11-18
47 | HPR | 2010-11-18 | jdaven | R100_04 | R100-CAV-HPR | R1 | 47 | 1 | 1 | Assembly group was 2 hours late in providing the Chemistry group with a cage for the cavity. |  | Communicate |  | Chemistry group had to come back unnecessarily to remove cavity from HPR.
48 | ELECTRICAL | 2010-11-21 | castagno | R100_04 | R100-CAV-BAKE | R1 | 9 |  |  | Halfway through ramp up to 120 C, GFCI outlet tripped that heater power supply cord was plugged into. | 76341BEC-0207-ECDB-EB91E55C90D0271C | Heater power supply cord was rerouted to a non-GFCI outlet. Montitored the remainder of temp ramp up to 120 C. |  | Continue with cavity processing. | castagno | 2010-11-22
49 | EP | 2010-11-17 | jdaven | R100_04 | R100-CAV-EPOL | R1 | 9 | 1 | 2 | R100-04 is welded into the helium vessel yet another 5um EP has been requested. |  | Extra steps taken in an effort to reduce possible safety hazards and equipment damage. |  | 1. Engine hoist used to install/remove cavity from EP carriage due to additional weight of the helium vessel. 2. Hose clamps were used in an effort to create a "spider support web" for the helium vessel between the end groups and the carriage cage. 3. Aluminum blocks were wedged between the helium vessel and the carriage cage at two points to support weight. | jdaven | 2010-11-17
50 |  | 2010-12-02 | kdavis | R100_4 | R100-CAV-VTRF | R1 | 34 | 4 | 21 | Q0 vs Temperature measurements were made at 17.5 MV/m due to field probe Qext error at time of test.
51 | OTHER | 2011-01-14 |  | 0TEST | R100-CAV-D3 | R1 | 1 | [9 empty] | 8623C638-F056-3A05-2924E101F0708872

## R100-D3
Status Report for R100-D3 on 10-Jun-20
TRAV_SEQ_NUM | D3CLASSIFICATION | D3OPENEDTIME | CATEGORIZETECHNICIAN | SERIALID | TRAVELERID | REVNO | SEQNO | PAGENO | STEPNO | D3DESCRIPTION | DESCRIPTIONFILES | D3ACTIONS | ACTIONFILES | D3RESOLVE | RESOLVETECHNICIAN | D3_RESOLVED_TIME | RESOLVEFILES
1 | RF_TEST_ABORTED | 2010-09-08 | [7 empty] | Regarding the VTA RF tests on R100_8: at 2.07K, recording the initial parameters for VTA testing, found cavity vacuum at 1.6e-5mbar. Contact technical support. Closing the main vacuum valve, the pressure drops to 1 e-9 mbar (vacuum in the pumping system). Opening the valve, the pressure went up to 1.3 e-5 mbar. Warming up the cavity, the leak closed bellow 4K, actually at 2.2K. Seems to have a super fluid vacuum leak. |  | Identify the type of leak. VTA tests on R100_8, will be done at 1.6 e-5 mbar. |  |  | FOREHAND | 2010-09-08
2 | BCP | 2010-12-01 | dreyfuss | R100-001 string | R100-CST-ASSY-ASO1 | R1 | 2 |  |  | R100-006, 005, and 004 helium vessels where not wiped done before waveguide was assemblied onto cavity |  | wiped down waveguide after assembly |  | put cavities on string | dreyfuss | 2010-12-01
3 | BCP | 2010-12-01 | dreyfuss | R100-001 string | R100-CST-ASSY-ASO1 | R1 | 2 |  |  | supply and return valves needed modification to the studs and backing plates and stiffing rings |  | Shortened four studs and cut half the stiffing ring and widened backing plates and remmed out holes |  | Attached to cavity | dreyfuss | 2010-12-01
4 | BCP | 2010-12-07 | macha | R100_String | R100-CST-ASSY-ASO1 | R1 | 2 |  |  | Four leaks detected during final leak test. | C1FAD1C9-01A8-2185-59FAE5FC6C8CB7B4 | Slightly increased torque to 43'lbs on leaking flanges. The group decided to up the torque of the leaking flanges on the long sides only. See updated file. |  | Group decided to move forward with assembling the string into a cryomodule. | macha | 2010-12-07
5 |  | 2010-12-07 | macha | R100_STRING | R100-CST-ASSY-ASO1 | R1 | 2 |  |  | Waveguide and window serial number recording. Drop-down in the traveler are not correct. | C2684657-B0B7-7C61-F0B8C92F2E7A245F | See attachment
6 | BCP | 2011-01-13 |  | 2KCIRCUIT | R100-CM-ASSY-COLD | R1 | 2 |  |  | During initial leak check found a leak in supply line bellows located between wave guides 4 and 5 1e-8 leak rate found. Only leaked when bagged. Leak was identified and found in the factory seam of the bellows. |  | Bellows assembly was removed from circuit and new bellows welded in its' place. |  | Did initial leak check prior to bagging to check weld. Ok. bagged bellows and leak checked Data submitted on circuit will verify repair is leak tight | dcotton | 2011-01-13
7 | OTHER | 2011-01-14 |  | 0TEST | R100-D3 | R1 | 1
8 | OTHER | 2011-03-03 | lpage | R100 | R100-CM-ASSY-SFR | R1 | 2 |  |  | During the Renascence build (years ago), a section of the thermal shield (above cavity position no.-6) was removed to provide access to repair a 2K leak on Renascence. This section was not installed when the rest of the shield was re-assembled as part of R100 assembly. This was discovered while attaching the G-10 supports in preparation of hanging the shield in the space frame. |  | Permanently tack welded original patch section of 50k shield onto R100 that was previous removed from Renascence to repair a leak. | 7C62A0B4-D4A1-ADB7-08543DBC877EEE87 | After the welding was complete leak test performed on 50k circuit to verify integrity. No leaks detected. | worland | 2011-03-02 | 7C533770-B82C-ACB9-42CB7600266A8DC2