| Traveler Title | SNSPPU Cavity Receiving Inspection |  |  |  |
| --- | --- | --- | --- | --- |
| Traveler Abstract | Incoming Inspection of SNSPPU Production Cavities. |  |  |  |
| Traveler ID | SNSPPU-CAV-INSP-CAV |  |  |  |
| Traveler Revision | R2 |  |  |  |
| Traveler Author | Pashupati Dhakal |  |  |  |
| Traveler Date | 6-Oct-20 |  |  |  |
| NCR Informative Emails | kwilson, edaly, dhakal, macha |  |  |  |
| NCR Dispositioners | kwilson, edaly, dhakal, macha |  |  |  |
| D3 Emails | kwilson, edaly, dhakal, macha |  |  |  |
| Approval Names | Pashupati Dhakal | Kurt Macha | G. Dekerlegand | Ed Daly |
| Approval Signatures |  |  |  |  |
| Approval Dates |  |  |  |  |
| Approval Title | Author | Reviewer | Reviewer | Project Manager |

| References | List and Hyperlink all documents related to this traveler. This includes, but is not limited to: safety (THAs, SOPs, etc), drawings, procedures, and facility related documents. |  |  |  |
| --- | --- | --- | --- | --- |
| 104211800-M8U-8200-A001 Cavity Assembly | 104211800-M8U-8200-A002 Probe End Sub-Assy | 104211800-M8U-8200-A002 FPC End Sub-Assy | Cavity in Shipping Crate | 104211800-M8U-8200-A017 FPC Flange Detail. |
| 104211800-M8U-8200-A016 FPC End Flange Detail | 104211800-M8U-8200-A011 Probe End Flange Detail |  |  |  |

| Revision Note |  |
| --- | --- |
| R1 | Initial release of this Traveler. |
| R2 | First Revision after the first cavity Inspection |

| Step No. | Instructions | Data Input |
| --- | --- | --- |
| 1 | Enter the cavity serial number ID that is going to be inspected. The SN should be ingraved on end dish. It not contact cavity SOTR. | [[CAVSN]] <<CAVSN>> [[Inspector1]] <<SRF>> |
| 3 | Cell Serialization ( FPC-FP, Cell 1-6) Each cell has two SN for each half record both. The numers are alpha numeric. | [[Cell 1]] <<FLOAT>> [[Cell 2]] <<FLOAT>> [[Cell 3]] <<FLOAT>> [[Cell 4]] <<FLOAT>> [[Cell 5]] <<FLOAT>> [[Cell 6]] <<FLOAT>> |
| 4 | Ensure the right Angle Valve is in the closed position. | [[RightAngleValve Clsoed]] <<YESNO>> [[ValveComments]] <<Comments>> |
| 5 | Torque Value Check. A286 8/32 – 40 in-lb. A286 5/16 – 30 ft-lb Joints are numbered 1 to 7. NOTE - Starting location is at valve end of cavity | [[Are torque values correct]] << YES NO>> [[Torque]] <Comments>> |
|  | 1. Valve screws ~30 ft-lb | [[DIMTOL25]] <<YESNO>> [[TorqComment1]] <<COMMENT>> |
|  | 2. Beamline flange FPC End ~ 30 ft-lb | [[DIMTOL26]] <<YESNO>> [[TorqComment2]] <<COMMENT>> |
|  | 3. FPC Flange ~ 30 ft-lb | [[DIMTOL27]] <<YESNO>> [[TorqComment3]] <<COMMENT>> |
|  | 4. FPC coupler ~ 30 ft-lb | [[DIMTOL28]] <<YESNO>> [[TorqComment4]] <<COMMENT>> |
|  | 5. Burst disk joint ~ 40 in-lb | [[DIMTOL29]] <<YESNO>> [[TorqComment5]] <<COMMENT>> |
|  | 6. Beamline flange Field Probe end ~30 ft-lb | [[DIMTOL30]] <<YESNO>> [[TorqComment6]] <<COMMENT>> |
|  | 7. Field probe ~ 30 ft-lb | [[DIMTOL31]] <<YESNO>> [[TorqComment7]] <<COMMENT>> |
| 6 | CMM Inspections Perform dimensional inspections on CMM. Fill out NCR for any out of tolerance measurements. 104211800-M8U-8200-A001 Cavity Assembly 104211800-M8U-8200-A002 Probe End Sub-Assy 104211800-M8U-8200-A003 FPC End Sub-Assy 104211800-M8U-8200-A017 FPC Flange Detail 104211800-M8U-8200-A011 Probe End Flange Detail 104211800-M8U-8200-A016 FPC End Flange Detail | [[Inspector3]] <<SRF>> [[CmmComments]] <<COMMENT>> [[AdditionalFile2]] <<FILEUPLOAD>> |

| Drawing number | Description | Drawing Value (inch) | Measured Value | Within Tolerance |
| --- | --- | --- | --- | --- |
| 104211800-M8U-8200-A001 | Distance between end dishes. Measure from machined step to machined step. | 35.423 | [[DIM1]] <<FLOAT>> | Reference Dimension N/A |
| 104211800-M8U-8200-A003 | Coupler flange perpendicularity. | .005 | [[DIM2]] <<FLOAT>> | [[DIMTOL2]] <<YESNO>> |
| 104211800-M8U-8200-A017 | Coupler flange perpendicularity OD | 4.889 | [[DIM3]] <<FLOAT>> | [[DIMTOL3]] <<YESNO>> |
| 104211800-M8U-8200-A016 | Beamline Flange OD- FPC side | 4.889 | [[DIM4]] <<FLOAT>> | [[DIMTOL4]] <<YESNO>> |
| 104211800-M8U-8200-A011 | Beamline Flange OD – FP side | 4.889 | [[DIM5]] <<FLOAT>> | [[DIMTOL5]] <<YESNO>> |
| 104211800-M8U-8200-A002 | End Flange to Dish ( FP side) | 6.397 | [[DIM6]] <<FLOAT>> | Reference Dimension N/A |
| 104211800-M8U-8200-A003 | End Flange to Dish ( FPC Side) | 9.01 | [[DIM7]] <<FLOAT>> | Reference Dimension N/A |
| 104211800-M8U-8200-A003 | Perpendicularity of FPC end dish to beam center | .005 | [[DIM8]] <<FLOAT>> | N/A |
| 104211800-M8U-8200-A002 | Perpendicularity of FP end dish to beam center | .005 | [[DIM9]] <<FLOAT>> | N/A |
| 104211800-M8U-8200-A002 | Tuner 120 Deg. C-Bore location | 2.559 5.118 | [[DIM10]] <<FLOAT>> [[DIM11]] <<FLOAT>> | [[DIMTOL10]] <<YESNO>> [[DIMTOL11]] <<YESNO>> |
| 104211800-M8U-8200-A002 | Tuner tapped hole locations. Verify 10-32 Thread inserts are installed and test thread. | .590 1.180 2.559 5.118 | [[DIM12]] <<FLOAT>> [[DIM13]] <<FLOAT>> [[DIM14]] <<FLOAT>> [[DIM15]] <<FLOAT>> | [[DIMTOL12]] <<YESNO>> [[DIMTOL13]] <<YESNO>> [[DIMTOL14]] <<YESNO>> [[DIMTOL15]] <<YESNO>> [[ThreadCheck1]] <<YESNO>> |
| 104211800-M8U-8200-A002 104211800-M8U-8200-A001 | Lollipop tapped hole C-Bore face. Side of flange at FP side. Flange rotation, Lollipop tapped hole 0 Deg. To FPC flange. Verify 5/8-11 threads with thread gauge. | 2.351 0 Deg. | [[DIM16]] <<FLOAT>> [[DIM17]] <<FLOAT>> | [[DIMTOL16]] <<YESNO>> [[DIMTOL17]] <<YESNO>> [[ThreadCheck2]] <<YESNO>> |
| 104211800-M8U-8200-A002 | Alignment tool tapped hole C-Bore face. Side of flange at FP side. Hole location 90 deg from FPC face Verify 5/8-11 threads with thread gauge. | 2.351 2.351 90 Deg | [[DIM18]] <<FLOAT>> [[DIM19]] <<FLOAT>> [[DIM20]] <<FLOAT>> | [[DIMTOL18]] <<YESNO>> [[DIMTOL19]] <<YESNO>> [[DIMTOL20]] <<YESNO>> [[ThreadCheck3]] <<YESNO>> |
| 104211800-M8U-8200-A003 104211800-M8U-8200-A001 | Lollipop tapped hole C-Bore face. Side of flange at FPC side. Flange rotation, Lollipop tapped hole 0 Deg. To FPC flange. Verify 5/8-11 threads with thread gauge. | 2.351 0 Deg. | [[DIM21]] <<FLOAT>> [[DIM22]] <<FLOAT>> | [[DIMTOL21]] <<YESNO>> [[DIMTOL22]] <<YESNO>> [[ThreadCheck4]] <<YESNO>> |
| 104211800-M8U-8200-A002 104211800-M8U-8200-A001 | Alignment tool tapped hole C-Bore face. Side of flange at FPC side. Hole location 90 deg from FPC face Verify 5/8-11 threads with thread gauge. Are the threads okay? | 2.351 2.351 90 Deg | [[DIM23]] <<FLOAT>> [[DIM24]] <<FLOAT>> | [[DIMTOL23]] <<YESNO>> [[DIMTOL24]] <<YESNO>> [[ThreadCheck5]] <<YESNO>> |