| Traveler Title | SNSPPU Cavity Receiving Inspection |  |  |  |
| --- | --- | --- | --- | --- |
| Traveler Abstract | Incoming Inspection of SNSPPU Production Cavities. |  |  |  |
| Traveler ID | SNSPPU-INSP-CAV |  |  |  |
| Traveler Revision | R2 |  |  |  |
| Traveler Author | Pashupati Dhakal |  |  |  |
| Traveler Date | 9-Oct-20 |  |  |  |
| NCR Informative Emails | Edaly |  |  |  |
| NCR Dispositioners | kwilson,dhakal,macha |  |  |  |
| D3 Emails | kwilson,edaly,dhakal,macha |  |  |  |
| Approval Names | Pashupati Dhakal | Kurt Macha | Earl Mosby | Ed Daly |
| Approval Signatures |  |  |  |  |
| Approval Dates |  |  |  |  |
| Approval Title | Author | Reviewer | Reviewer | Project Manager |

| References | List and Hyperlink all documents related to this traveler. This includes, but is not limited to: safety (THAs, SOPs, etc), drawings, procedures, and facility related documents. |  |  |  |
| --- | --- | --- | --- | --- |
| 104211800-M8U-8200-A001 Cavity Assembly | 104211800-M8U-8200-A002 Probe End Sub-Assy | 104211800-M8U-8200-A002 FPC End Sub-Assy | Cavity in Shipping Crate | 104211800-M8U-8200-A017 FPC Flange Detail |
| 104211800-M8U-8200-A016 FPC End Flange Detail | 104211800-M8U-8200-A011 Probe End Flange Detail |  |  |  |

| Revision Note |  |
| --- | --- |
| R1 | Initial release of this Traveler. |
| R2 | Frist revision. |

| Step No. | Instructions | Data Input |
| --- | --- | --- |
| 1 | Un-crate the cavity from the shipping box and enter the cavity serial number ID that is going to be inspected. | [[CAVSN]] <<CAVSN>> [[Inspector1]] <<SRF>> |
| 2 | Inspect for any visual damage. Record the component serial numbers, if any. Visual inspection should be looking for any obvious damage to the cavity as well as damage to the input and FPC connectors. Pay special attention the the feedthrough pins. Inspect the busrt disk to ensure it is not dented or broken. Field Probe SN Input Probe SN End Group FPC side SN End Group FP side SN Flange FPC side SN Flange FPC side SN Burst Disc SN Cavity in Shipping Crate | [[FPFTSN]] <<SN>> [[INFTSN]] <<SN>> [[EndGroupFPC]] <<SN>> [[EndGroupFP]] <<SN>> [[Flange 1]] <<SN>> [[Flange 2]] <<SN>> [[BDSN]] <<SN>> [[PassVisualInspectionOK]] <<YESNO>> [[VisualInspectComments]] <<COMMENT>> [[Inspector2]] <<SRF>> [[AdditionalFile1]] <<FILEUPLOAD>> |
| 3 | Ensure the right Angle Valve is in the closed position. | [[RightAngleValveClosed]] <<YESNO>> [[ValveComments]] <<COMMENT>> |
| 4 | Are cell Serialized? Record ( FPC to FP, Cell 1-6). Each cell has two SN for each half record both. The numers are alpha numeric. | [[Cell1]] <<SN>> [[Cell2]] <<SN>> [[Cell3]] <<SN>> [[Cell4]] <<SN>> [[Cell5]] <<SN>> [[Cell6]] <<SN>> |
| 5 | Check Torque Value on Screws Joints are numbered 1 to 7. NOTE - Starting location is at valve end of cavity A286 5/16 ~ 30 ft-lb (Locations 2, 3 and 6) A286 8-32 ~ 40 in-lb (Locations 5 and 7) Valve flange snd Input flance (2 ¾ ) > 11 ft-lb (Locations 1 and 4) | [[Location 1]] <<YESNO>> [[TorqComment1]] <<COMMENT>> [[Location 2]] <<YESNO>> [[TorqComment2]] <<COMMENT>> [[Location 3]] <<YESNO>> [[TorqComment3]] <<COMMENT>> [[Location 4]] <<YESNO>> [[TorqComment4]] <<COMMENT>> [[Location 5]] <<YESNO>> [[TorqComment5]] <<COMMENT>> [[Location 6]] <<YESNO>> [[TorqComment6]] <<COMMENT>> [[Location 7]] <<YESNO>> [[TorqComment7]] <<COMMENT>> [[Are torque values correct]] << YES NO>> [[Torque]] <Comments>> |
| 6 | CMM Inspections Perform dimensional inspections on CMM. Fill out NCR for any out of tolerance measurements. 104211800-M8U-8200-A001 Cavity Assembly 104211800-M8U-8200-A002 Probe End Sub-Assy 104211800-M8U-8200-A003 FPC End Sub-Assy 104211800-M8U-8200-A017 FPC Flange Detail 104211800-M8U-8200-A011 Probe End Flange Detail 104211800-M8U-8200-A016 FPC End Flange Detail | [[Inspector3]] <<SRF>> [[CmmComments]] <<COMMENT>> [[AdditionalFile2]] <<FILEUPLOAD>> |

| Drawing number | Description | Drawing Value (inch) | Measured Value | Within Tolerance |
| --- | --- | --- | --- | --- |
| 104211800-M8U-8200-A001 | Distance between end dishes. Measure from machined step to machined step. | 35.423 | [[DIM1]] <<FLOAT>> | Reference Dimension N/A |
| 104211800-M8U-8200-A003 | Coupler flange perpendicularity. | .005 | [[DIM2]] <<FLOAT>> | [[DIMTOL2]] <<YESNO>> |
| 104211800-M8U-8200-A017 | Coupler flange perpendicularity OD | 4.889 | [[DIM3]] <<FLOAT>> | [[DIMTOL3]] <<YESNO>> |
| 104211800-M8U-8200-A016 | Beamline Flange OD- FPC side | 4.889 | [[DIM4]] <<FLOAT>> | [[DIMTOL4]] <<YESNO>> |
| 104211800-M8U-8200-A011 | Beamline Flange OD – FP side | 4.889 | [[DIM5]] <<FLOAT>> | [[DIMTOL5]] <<YESNO>> |
| 104211800-M8U-8200-A002 | End Flange to Dish ( FP side) | 6.397 | [[DIM6]] <<FLOAT>> | Reference Dimension N/A |
| 104211800-M8U-8200-A003 | End Flange to Dish ( FPC Side) | 9.01 | [[DIM7]] <<FLOAT>> | Reference Dimension N/A |
| 104211800-M8U-8200-A003 | Perpendicularity of FPC end dish to beam center | .005 | [[DIM8]] <<FLOAT>> | N/A |
| 104211800-M8U-8200-A002 | Perpendicularity of FP end dish to beam center | .005 | [[DIM9]] <<FLOAT>> | N/A |
| 104211800-M8U-8200-A002 | Tuner 120 Deg. C-Bore location | 2.559 5.118 | [[DIM10]] <<FLOAT>> [[DIM11]] <<FLOAT>> | [[DIMTOL10]] <<YESNO>> [[DIMTOL11]] <<YESNO>> |
| 104211800-M8U-8200-A002 | Tuner tapped hole locations. Verify 10-32 Thread inserts are installed and test thread. | .590 1.180 2.559 5.118 | [[DIM12]] <<FLOAT>> [[DIM13]] <<FLOAT>> [[DIM14]] <<FLOAT>> [[DIM15]] <<FLOAT>> | [[DIMTOL12]] <<YESNO>> [[DIMTOL13]] <<YESNO>> [[DIMTOL14]] <<YESNO>> [[DIMTOL15]] <<YESNO>> [[ThreadCheck1]] <<YESNO>> |
| 104211800-M8U-8200-A002 104211800-M8U-8200-A001 | Lollipop tapped hole C-Bore face. Side of flange at FP side. Flange rotation, Lollipop tapped hole 0 Deg. To FPC flange. Verify 5/8-11 threads with thread gauge. | 2.351 0 Deg. | [[DIM16]] <<FLOAT>> [[DIM17]] <<FLOAT>> | [[DIMTOL16]] <<YESNO>> [[DIMTOL17]] <<YESNO>> [[ThreadCheck2]] <<YESNO>> |
| 104211800-M8U-8200-A002 | Alignment tool tapped hole C-Bore face. Side of flange at FP side. Hole location 90 deg from FPC face Verify 5/8-11 threads with thread gauge. | 2.351 2.351 90 Deg | [[DIM18]] <<FLOAT>> [[DIM19]] <<FLOAT>> [[DIM20]] <<FLOAT>> | [[DIMTOL18]] <<YESNO>> [[DIMTOL19]] <<YESNO>> [[DIMTOL20]] <<YESNO>> [[ThreadCheck3]] <<YESNO>> |
| 104211800-M8U-8200-A003 104211800-M8U-8200-A001 | Lollipop tapped hole C-Bore face. Side of flange at FPC side. Flange rotation, Lollipop tapped hole 0 Deg. To FPC flange. Verify 5/8-11 threads with thread gauge. | 2.351 0 Deg. | [[DIM21]] <<FLOAT>> [[DIM22]] <<FLOAT>> | [[DIMTOL21]] <<YESNO>> [[DIMTOL22]] <<YESNO>> [[ThreadCheck4]] <<YESNO>> |
| 104211800-M8U-8200-A002 104211800-M8U-8200-A001 | Alignment tool tapped hole C-Bore face. Side of flange at FPC side. Hole location 90 deg from FPC face Verify 5/8-11 threads with thread gauge. Are the threads okay? | 2.351 2.351 90 Deg | [[DIM23]] <<FLOAT>> [[DIM24]] <<FLOAT>> | [[DIMTOL23]] <<YESNO>> [[DIMTOL24]] <<YESNO>> [[ThreadCheck5]] <<YESNO>> |