|  |  |
| --- | --- |
| Traveler Title | L2 HE Cavity High Pressure Rinse |
| Traveler Abstract | This traveler documents an L2 HE cavity high pressure rinse. |
| Traveler ID | L2HE-CHEM-CAV-HPR |
| Traveler Revision  | R1 |
| Traveler Author | Alex Wildeson |
| Traveler Date | 28-Apr-2021 |
| NCR Informative Emails | ashleya |
| NCR Dispositioners | kwilson, ganey  |
| Approval Names | A. Wildeson | A. Mitchell | T. Ganey | K. Wilson |
| Approval Signatures |  |  |  |  |
| Approval Date |  |  |  |  |
| Approval Title | Author | Reviewer | Reviewer | Project Manager |

|  |  |
| --- | --- |
| References | List and Hyperlink all documents related to this traveler. This includes, but is not limited to: safety (THAs, SOPs, etc), drawings, procedures, and facility related documents. |
|  |  |  |  |  |
|  |  |  |  |  |

|  |  |
| --- | --- |
| Revision Note |  |
| R1 | Initial release of this Traveler. |

|  |  |  |
| --- | --- | --- |
| **Step No.** | **Instructions** | **Data Input** |
| 1 | Record Cavity ID, Rinse Operator, Technician (if additional person assisting), Date and Time | [[CAVSN]] <<CAVSN>>[[Operator1]] <<SRFCVP>>[[Operator2]] <<SRFCVP>>[[DateAndTime]] <<TIMESTAMP>> |
| 2 | Select whether this is the first or second (final) HPRRecord HPR data:* Record cavity cage SN:
* Record HPR head type (“Production Nozzle” is with six fan-jet nozzles, 15deg jet angle)
* Record number of passes
* Record HPR start time
* Record HPR Pump Pressure
* Record HPR tool inlet UPW resistivity value
* Record HPR End time
* Record any comments for the HPR run

Record cavity has been properly tipped to drain water after HPR | [[HPRNumber]] {{First,Second}} <<SELECT>>[[Cage]] <<SN>>[[HPRHead]] <<COMMENT>>[[NumberPasses]] <<INTEGER>>[[HPRStartTime]] <<TIMESTAMP>>[[PumpPressure]] <<FLOAT>>psi[[WaterInletResistivity]] <<FLOAT>>Mohm cm[[HPREndTime]] <<TIMESTAMP>>[[HPRComments]] <<COMMENT>>[[CavityTipped]] <<YESNO>> |