## Sheet1
NCR SEQ NUM | NCR OPEN DATE | NCR CLOSE DATE | TRAVELER ID | PART DESCRIPTION | SN | DESCRIBE COMMENT | INITIAL DISPOSITION | FINAL DISPOSITION | FINAl COMMENT
1 | 2021-03-09 |  | L2HE-INSP-FPCWG |  | 4 | Waveguide coupler flange (Dwg # SA-375-600-90) has multiple scratches present on flange. Please see attached. Multiple scratched areas checked w/profilometer. The worst area found was Rz (peak to valley) measured 294 u inch and RA (surface finish) measured 49 u inch. | Return To Vendor |  | Complete remainder of inspection traveler before returning to vendor
2 | 2021-05-04 |  | L2HE-VTA-CAV-VTRF |  | 331 | Q0(20.8) is 2.4e10, slightly lower than spec which is 2.5e10
3 | 2021-05-05 |  | L2HE-VTA-CAV-VTRF | HE_Cavity | CAVR009 | Cavity did not pass due to FE up to 1R/hr.
4 | 2021-05-18 | 2021-08-10 | L2HE-VTA-CAV-VTRF |  | HE-40 | Cavity frequency was low at 1299.975 MHz |  |  | re-test
5 | 2021-05-18 |  | L2HE-VTA-CAV-VTRF |  | HE-40 | Operator error except for two decay measurements the test was done at the 8pi/9 mode. For those measurements, Qo=2.9e10, Qfpc=1.94e10,Qfp=3.4e11 E = 2.9 MV/m |  |  | re-rinse/re-test did not meet requirements
7 | 2021-05-19 |  | L2HE-INSP-FPFT | PICKUP FEEDTHROUGH | 2061 | Probe has stain
8 | 2021-05-19 |  | L2HE-INSP-FPFT |  | 2062 | Scratches down outside diameter.
9 | 2021-05-25 |  | L2HE-INSP-HMFT |  | 2134 | HEH2134 has multiple dings down Antenna. | Use As Is |  | Surface finish poor but field in this area is not very strong either. Visual findings can be used to relate to cavity performance later on.
10 | 2021-05-25 | 2021-07-28 | L2HE-INSP-HMFT |  | 2143 | HEH2143 probe antenna has metallic foreign material on face. | Hold/Modify |  | See if chemistry will be able to remove the foreign material or residue.
11 | 2021-05-26 |  | L2HE-VTA-CAV-VTRF |  | HE-R009 | Qext2 is 2.9200e+11, out of spec which is 0.75-2.5 e12.
12 | 2021-05-26 | 2021-08-10 | L2HE-VTA-CAV-VTRF |  | HE-R009 | QextHOMb is 1.11e+11, out of spec which is >=2.7e11. Also, PHOMB at 21MVm was 3.52 W, out of spec which is ? 1.7 W |  |  | retune HOM
13 | 2021-05-27 |  | L2HE-INSP-FPFT |  | 2078 | Probe has multiple dings down along stem. | Use As Is
14 | 2021-05-27 |  | L2HE-INSP-FPFT |  | 2082 | Surface seal area has small imperfection. | Hold/Modify |  | Send to machine shop to remove the irregularity on seal path.
15 | 2021-05-27 |  | L2HE-INSP-FPFT |  | 2083 | Probe has gouge on O.D
16 | 2021-05-27 |  | L2HE-INSP-FPFT |  | 2085 | Probe O.D all the way around has rough finish. | Use As Is
17 | 2021-06-02 |  | L2HE-INSP-FPFT | PICKUP FEEDTHROUGH | 2143 | Probe has line of rough imperfections.
18 | 2021-06-02 |  | L2HE-INSP-FPFT |  | 2151 | Seal surface area has rough finish.
19 | 2021-06-03 |  | L2HE-INSP-FPFT |  | 2157 | Probe has rough finish | Use As Is
20 | 2021-06-03 |  | L2HE-INSP-FPFT |  | 2158 | Probe has rough finish | Use As Is
21 | 2021-06-03 |  | L2HE-INSP-HMFT |  | 2235 | Seal surface area has rough finish. Probe 6mm radius has rough finish.
22 | 2021-06-03 |  | L2HE-INSP-HMFT | HOM FEEDTHROUGH | 2238 | Seal surface area is not smooth. Probe radius has rough finish.
23 | 2021-06-04 |  | L2HE-INSP-CAV | LCLS-II PRODUCTION CAVITY ASSEMBLY | 284 | Visual inspection of cavity bellows on SN CAV-284 found (3) minor dents in between bellow convolutions. See pictures in attach box.
24 | 2021-06-08 |  | L2HE-INSP-HMFT |  | 2049 | Flange seal surface to probe tip distance measured 29.50 mm, slightly out of spec which is 29.51-29.61 mm. | Hold/Re-measure |  | Hold for now, further evaluation will be done.
26 | 2021-06-15 |  | L2HE-INSP-FPFT |  | 2048 | Probe is loose.
27 | 2021-06-18 |  | L2HE-INSP-FPFT | PICKUP FEEDTHROUGH | 2138 | Seal surface finish is not smooth
28 | 2021-06-24 | 2021-06-25 | L2HE-INSP-CAV | LCLS-II PRODUCTION CAVITY ASSEMBLY | 145 | Visual inspection of bellow convolutions found (2) small dents in bottom of convolution. | Use As Is | Use As Is
29 | 2021-06-24 |  | L2HE-INSP-CAV | LCLS-II PRODUCTION CAVITY ASSEMBLY | 118 | Visual inspection of bellow convolution only,found (1) small dent on bottom of bellow convolution. | Use As Is
30 | 2021-07-06 |  | L2HE-CLNRM-FPCC-INSP | UNKN | CPI-FPCC-HE-001 | found small ding in bellows of CPI-001 coupler
31 | 2021-07-07 | 2021-07-07 | L2HE-INSP-CAV | LCLS-II PRODUCTION CAVITY ASSEMBLY | 342 | Testing Only! | Hold/Modify | Conforming | TESTING ONLY
32 | 2021-07-12 | 2021-08-10 | L2HE-CMA-UCM-INSP | LCLS-II UPPER COLD MASS ASSEMBLY | RFCHE-UC-002 | Drawing part # F10074420 Vertical support lines twisted. Please see attached. |  | Use As Is
33 | 2021-07-13 | 2021-08-10 | L2HE-CMA-UCM-INSP | LCLS-II UPPER COLD MASS ASSEMBLY | RFCHE-UC-002 | Copper clamp indium foil applied to wrong half. | Hold/Modify | Conforming | 1. Check/disassemble all 16 copper clamps. 2. Apply indium foil to the long halves on the clamps. No need to remove the indium on the short half. 3. Re-assemble clamps, torque three times with 8-hr intervals.
34 | 2021-07-13 | 2021-08-10 | L2HE-CMA-UCM-INSP | LCLS-II UPPER COLD MASS ASSEMBLY | RFCHE-UC-002 | Cavity hangar has sharp edges. | Hold/Modify | Conforming | Deburr sharp edges
35 | 2021-07-13 | 2021-08-10 | L2HE-CMA-UCM-INSP | LCLS-II UPPER COLD MASS ASSEMBLY | RFCHE-UC-002 | During the inspection, we came across bolts that had suspicious head markings.We had M16 and 1/4-20 hardware that had "vvv" on the head instead of "THE". The proper authorities were notified. |  | Use As Is
36 | 2021-07-13 | 2021-08-11 | L2HE-CWI-VV | VESSEL LCLS-II AFTER MACHING WELDMENT | RFCHE-VV-002 | The fillet weld located on weld#C1 have been reduced in size via some machining operation. The welds are not complaint with the drawing, specifications or the welding code. It is my intention to reject the fillet weld portion of this flange weld joint that also includes a CJP groove weld. However, I found the groove weld portion of this same flange weld joint to be in compliance. |  | Use As Is
37 | 2021-07-13 |  | L2HE-VTA-CAV-VTRF | LCLS-II PRODUCTION CAVITY ASSEMBLY | 331 | Cavity quench at 21.8 MV/m | Hold/Re-measure |  | We will look into re-rinse/re-test for this cavity
38 | 2021-07-31 | 2021-07-31 | L2HE-VTA-CAV-VTRF | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-331 | FE onset 5.5 MV/m, FE loading and 1R/hr at 14 MV/m | Hold/Re-measure | Reject | re-rinse https://bit.ly/3jbEoiG
39 | 2021-06-22 | 2021-08-10 | L2HE-CMA-FPCW-INSP | FPC WARM PART | CPI-FPCW-HE-001 | Found scratch on copper plating of DN40 flange |  | Return To Vendor
40 | 2021-06-23 |  | L2HE-CMA-FPCW-INSP | FPC WARM PART | CPI-FPCW-HE-002 | Found scratch on copper plating of DN40 flange
41 | 2021-07-06 |  | L2HE-CMA-FPCW-INSP | FPC WARM PART | CPI-FPCW-HE-007 | Found scratches on copper plating of DN40 flange
42 | 2021-07-19 |  | L2HE-CMA-FPCW-INSP | FPC WARM PART | CPI-FPCW-HE-005 | Found scratch on copper plating of DN40 flange
43 | 2021-07-19 |  | L2HE-CMA-FPCW-INSP | FPC WARM PART | CPI-FPCW-HE-006 | Push rod bellows distorted and slightly twisted. Push rod was difficult to remove and reinstall. The M6 antenna vented bolt was loose
44 | 2021-07-20 |  | L2HE-CMA-FPCW-INSP | FPC WARM PART | CPI-FPCW-HE-009 | Face if DN100 flange scratched, small scratch on outer conductor copper plating.
45 | 2021-07-20 |  | L2HE-CMA-FPCW-INSP | FPC WARM PART | CPI-FPCW-HE-010 | Push rod bellows very bent and twisted, some convolutions were pinched. Push rod was not engaged. One of the RF contact springs was undone.
46 | 2021-07-20 |  | L2HE-CMA-FPCW-INSP | FPC WARM PART | CPI-FPCW-HE-010 | RF contact area of inner conductor copper was pitted and copper plating had defects
47 | 2021-08-10 |  | L2HE-INSP-IMAG | MAGNETIC SHIELD CAVITY ASSEMBLY | RFCHE-MS-009 | Part I.D.# F10048728 Rev.A Shield, LCLS-11 2nd layer center section A. has minor ding about 5/16" diameter and .010 deep. Part I.D.# F10048731 Shield, LCLS-11 2nd layer coupler side A. Has (2) sets of imperfections (1)is about 6" long the 2nd is about 4" long .
48 | 2021-08-10 |  | L2HE-INSP-IMAG | MAGNETIC SHIELD CAVITY ASSEMBLY | RFCHE-MS-009 | Part I.D.# F10048728 Rev.A Shield, LCLS-11 2nd layer center section A. Has light ding 3/8" diameter and .010 deep. Part I.D.# F10048731 Rev.A Shield,LCLS-11 2nd layer coupler side A. Has 6" long imperfection. marked in red.
49 | 2021-08-10 |  | L2HE-INSP-HMFT | HOM FEEDTHROUGH | 2161 | Flange seal surface to probe tip is 29.5 mm, slightly out of spec which is 29.51-29.61 mm

## Sheet2
DateTime Text | Date Text | Date
2021-03-16 14:24:27.0 | 2021-03-16 | 2021-03-16
Formulas:
Date = DATEVALUE([Date Text])