|  |
| --- |
| **Ionized Nitrogen Parts Cleaning** |
| **Document Number:** | SRF-MSPR-CLNRM-CST-ION | **Effective Date:** | DD Mmm YYYY |
| **Revision Number:** | R1 | **Periodic Review Date:** | DD Mmm YYYY |
| **Document Owner:** | D. Forehand | **Department Owner:** | SRF Operations |

# Purpose

The purpose of this document is to <enter text>.

# Scope

This procedure describes the preferred ionized Nitrogen cleaning of parts for the SNSPPU project.

# Terms and Definitions

The following terms have specific meanings within this procedure.

|  |  |
| --- | --- |
| **Term** | **Definition** |
| <Term 1> | <Definition> |
| <Term 2> | <Definition> |
|  |  |

# Roles and Responsibilities

The following roles have responsibilities described in this document.

|  |  |
| --- | --- |
| **Role** | **Responsibility** |
| <Job Title> | <Very short summary of activities this job title performs in this procedure.> |
|  |  |
|  |  |
|  |  |

# Procedure

* All associated hardware and flanges should already be UHV cleaned, bagged and ready in the production pass-through.

The parts cart shall be cleaned off with alcohol and several clean-room wipes spread out for the parts to rest on. Remove the associated hardware from the pass-through and layout on a table to ensure all the parts are there.

* The nitrogen cleaning table shall be sprayed with ionized N2 to verify cleanliness before any of the parts are blown off table to spec 1 on the particle counter. If the table particle counts are above 1 on any scale the table should be blown until the counts drop below one, or in extreme conditions the table will need to be wiped down with DI water and be allowed to dry.
* Appropriate sized dust covers for all flanges will be sprayed until they meet or exceed spec 2 on the particle counter and placed on cart.
* The tips of an appropriate number of spring clamps shall be sprayed until they reach or exceed spec 2 on the particle counter and placed on cart.
* Dust covers can now be placed on assembly IAW the appropriate assembly procedure.
* As the flanges and gaskets are removed from the bags, they shall be inspected visually for any defects that might cause problems with sealing during the assembly. If any scratches or defects are found, the part shall not be used for the assembly. If the issue cannot be resolved immediately, the assembly will be delayed until replacement or repaired parts are available.
* As parts are placed on the cart, ensure that they are laid out in a manner that allows the assembly tech to pick up the needed pieces without reaching over any parts that will be used later in the assembly.
* If the assembly will require any number of sub-assemblies, perform the following steps.

If no sub-assemblies are needed move on to the next step.

* + Any gaskets needed for sub-assemblies can be sprayed with ionized N2 until they are able to meet or exceed spec 1 on the particle counter. Place the gaskets on the parts cart.
	+ At this time all the sub-assembly flanges shall be sprayed with the ionized nitrogen until all counts reach spec 2 on the particle counter. They shall then be placed on a cloth wiper on the parts cart.
	+ These parts can now be put together and made into the appropriate sub-assemblies.
	+ The sub-assemblies can now be sprayed again until they reach or exceed spec 1i on the particle counter.
* All flanges and associated gaskets shall be sprayed with ionized N2 until they reach or exceed spec 1 on the particle counter. Parts can be carefully placed on cart.
* All tools to be used during assembly shall also be sprayed until they reach or exceed spec
* 2 and placed on the cart.
* After all the proceeding steps are satisfied, the cart can be moved to the assembly area.

1 Particle counts are to be zero on all scales except 0.3µm, which can be zero or 1 in five seconds.

2 Particle counts can be 1 count per second or less on the 1µm scale.

# References

|  |  |
| --- | --- |
| **Document No.** | **Title** |
| SRF-01-ML-001 | SRF Quality Manual |
| <SRF-FM-###> | <Document Title> |
|  |  |

# Release and Revision History

|  |  |  |
| --- | --- | --- |
| **Rev #** | **Major Changes** | **Effective Date:** |
| 1 | Initial version | DD Mmm YYY |
|  |  |  |
|  |  |  |
|  |  |  |

# Approvals

|  |  |  |  |
| --- | --- | --- | --- |
| **Approved by:** | **Name:** | **Signature:** | **Date:** |
| Document Owner | D. Forehand |  | DD Mmm YYY |
| Subject Matter Expert | K. Davis |  | DD Mmm YYY |
| Work Center Lead  | <First Last Name> |  | DD Mmm YYY |
| Project Leader  | <First Last Name> |  | DD Mmm YYY |