## Sheet1
NCR SEQ NUM | NCR OPEN DATE | NCR CLOSE DATE | TRAVELER ID | PART DESCRIPTION | SN | DESCRIBE COMMENT | INITIAL DISPOSITION | INITIAL DISPOSITION COMMENT | FINAL DISPOSITION | FINAL COMMENT
1 | 2021-03-09 | 2021-03-16 | L2HE-INSP-FPCWG |  | 004 | Waveguide coupler flange (Dwg # SA-375-600-90) has multiple scratches present on flange. Please see attached. Multiple scratched areas checked w/profilometer. The worst area found was Rz (peak to valley) measured 294 u inch and RA (surface finish) measured 49 u inch. | Return To Vendor | Complete remainder of inspection traveler before returning to vendor
2 | 2021-05-04 | 2021-10-06 | L2HE-VTA-CAV-VTRF | Nine Cell | 0331 | Q0(20.8) is 2.4e10, slightly lower than spec which is 2.5e10 | Use As Is
3 | 2021-05-05 | 2021-10-06 | L2HE-VTA-CAV-VTRF | UNKN | R009 | Cavity did not pass due to FE up to 1R/hr. | Use As Is
4 | 2021-05-18 | 2021-08-10 | L2HE-VTA-CAV-VTRF | UNKN | 0040 | Cavity frequency was low at 1299.975 MHz | Hold/Re-measure | re-test | Reject
5 | 2021-05-18 | 2021-08-10 | L2HE-VTA-CAV-VTRF | UNKN | 0040 | Operator error except for two decay measurements the test was done at the 8pi/9 mode. For those measurements, Qo=2.9e10, Qfpc=1.94e10,Qfp=3.4e11 E = 2.9 MV/m | Reject | re-rinse/re-test did not meet requirements
7 | 2021-05-19 | 2021-08-17 | L2HE-INSP-FPFT | PICKUP FEEDTHROUGH | 2061 | Probe has stain | Hold/Modify | Send to chemistry, see if the stain can be removed. | Use As Is | Passed visual inspection after stain removal. Okay to use.
8 | 2021-05-19 | 2021-08-17 | L2HE-INSP-FPFT |  | 2062 | Scratches down outside diameter. | Use As Is | The scratches are not significant. Pictures are captured and okay to proceed to use.
9 | 2021-05-25 | 2021-06-25 | L2HE-INSP-HMFT |  | 2134 | HEH2134 has multiple dings down Antenna. | Use As Is | Surface finish poor but field in this area is not very strong either. Visual findings can be used to relate to cavity performance later on.
10 | 2021-05-25 | 2021-07-28 | L2HE-INSP-HMFT |  | 2143 | HEH2143 probe antenna has metallic foreign material on face. | Hold/Modify | See if chemistry will be able to remove the foreign material or residue. | Use As Is
11 | 2021-05-26 | 2021-08-10 | L2HE-VTA-CAV-VTRF | PRODUCTION CAVITY ASSEMBLY | R009 | Qext2 is 2.9200e+11, out of spec which is 0.75-2.5 e12. | Use As Is
12 | 2021-05-26 | 2021-08-10 | L2HE-VTA-CAV-VTRF |  | HE-R009 | QextHOMb is 1.11e+11, out of spec which is >=2.7e11. Also, PHOMB at 21MVm was 3.52 W, out of spec which is ? 1.7 W | Hold/Re-measure | retune HOM | Use As Is
13 | 2021-05-27 | 2021-06-25 | L2HE-INSP-FPFT |  | 2078 | Probe has multiple dings down along stem. | Use As Is
14 | 2021-05-27 | 2021-08-17 | L2HE-INSP-FPFT |  | 2082 | Surface seal area has small imperfection. | Hold/Modify | Send to machine shop to remove the irregularity on seal path. | Use As Is | Passed visual inspection after sealing area was polished. Okay to proceed to use.
15 | 2021-05-27 | 2021-08-17 | L2HE-INSP-FPFT |  | 2083 | Probe has gouge on O.D | Use As Is | The gouge does not appear significant. Pictures captured as reference, and okay to proceed to use.
16 | 2021-05-27 | 2021-06-25 | L2HE-INSP-FPFT |  | 2085 | Probe O.D all the way around has rough finish. | Use As Is
17 | 2021-06-02 | 2021-09-09 | L2HE-INSP-FPFT | PICKUP FEEDTHROUGH | 2143 | Probe has line of rough imperfections. | Use As Is | Imperfection is not significant. Probe pictures are captured and available for reference if there is future issue.
18 | 2021-06-02 | 2021-09-09 | L2HE-INSP-FPFT |  | 2151 | Seal surface area has rough finish. | Use As Is | Seal surface is measured with profilometer, average Ra is 12 uin, within drawing requirement which is 64 uin (1.6 um).
19 | 2021-06-03 | 2021-06-25 | L2HE-INSP-FPFT |  | 2157 | Probe has rough finish | Use As Is
20 | 2021-06-03 | 2021-06-25 | L2HE-INSP-FPFT |  | 2158 | Probe has rough finish | Use As Is
21 | 2021-06-03 | 2021-09-09 | L2HE-INSP-HMFT |  | 2235 | Seal surface area has rough finish. Probe 6mm radius has rough finish. | Use As Is | Seal surface measured with profilometer, average Ra is 22 uin, within drawing requirement which is 32 uin (0.8 um).
22 | 2021-06-03 | 2021-09-09 | L2HE-INSP-HMFT | HOM FEEDTHROUGH | 2238 | Seal surface area is not smooth. Probe radius has rough finish. | Use As Is | Seal surface measured with profilometer, average Ra 21 uin, within drawing requirement which is 32 uin (0.8 um). Probe roughness is okay to use as is. Probe pictures are captured as reference if there is issue later on.
23 | 2021-06-04 |  | L2HE-INSP-CAV | PRODUCTION CAVITY ASSEMBLY | 0284 | Visual inspection of cavity bellows on SN CAV-284 found (3) minor dents in between bellow convolutions. See pictures in attach box.
24 | 2021-06-08 | 2021-09-24 | L2HE-INSP-HMFT |  | 2049 | Flange seal surface to probe tip distance measured 29.50 mm, slightly out of spec which is 29.51-29.61 mm. | Hold/Re-measure | Hold for now, further evaluation will be done. | Use As Is | Per communication with SLAC, the antenna being 0.01 mm short is not likely to cause a problem in this case.
26 | 2021-06-15 | 2021-09-07 | L2HE-INSP-FPFT |  | 2048 | Probe is loose. | Reject | Set aside and do not use for production.
27 | 2021-06-18 | 2021-09-09 | L2HE-INSP-FPFT | PICKUP FEEDTHROUGH | 2138 | Seal surface finish is not smooth | Use As Is | Seal surface measured with profilometer, average Ra 12 uin, within drawing requirement which is 64 uin (1.6 um).
28 | 2021-06-24 | 2021-06-25 | L2HE-INSP-CAV | PRODUCTION CAVITY ASSEMBLY | 0145 | Visual inspection of bellow convolutions found (2) small dents in bottom of convolution. | Use As Is |  | Use As Is
29 | 2021-06-24 | 2021-06-25 | L2HE-INSP-CAV | PRODUCTION CAVITY ASSEMBLY | 0118 | Visual inspection of bellow convolution only,found (1) small dent on bottom of bellow convolution. | Use As Is
30 | 2021-07-06 | 2021-08-11 | L2HE-CLNRM-FPCC-INSP | UNKN | CPI-FPCC-HE-001 | found small ding in bellows of CPI-001 coupler | Return To Vendor
31 | 2021-07-07 | 2021-07-07 | L2HE-INSP-CAV | PRODUCTION CAVITY ASSEMBLY | 0342 | Testing Only! | Hold/Modify | TESTING ONLY | Conforming
32 | 2021-07-12 | 2021-08-10 | L2HE-CMA-UCM-INSP | UPPER COLD MASS ASSEMBLY | RFCHE-UC-002 | Drawing part # F10074420 Vertical support lines twisted. Please see attached. |  |  | Use As Is | Checked with Fermi. Rods can be used as is. See attached transcript
33 | 2021-07-13 | 2021-08-10 | L2HE-CMA-UCM-INSP | UPPER COLD MASS ASSEMBLY | RFCHE-UC-002 | Copper clamp indium foil applied to wrong half. | Hold/Modify | 1. Check/disassemble all 16 copper clamps. 2. Apply indium foil to the long halves on the clamps. No need to remove the indium on the short half. 3. Re-assemble clamps, torque three times with 8-hr intervals. | Conforming
34 | 2021-07-13 | 2021-08-10 | L2HE-CMA-UCM-INSP | UPPER COLD MASS ASSEMBLY | RFCHE-UC-002 | Cavity hangar has sharp edges. | Hold/Modify | Deburr sharp edges | Conforming
35 | 2021-07-13 | 2021-08-10 | L2HE-CMA-UCM-INSP | UPPER COLD MASS ASSEMBLY | RFCHE-UC-002 | During the inspection, we came across bolts that had suspicious head markings.We had M16 and 1/4-20 hardware that had "vvv" on the head instead of "THE". The proper authorities were notified. |  |  | Use As Is | Vendor supplied material certs per JLAB request, see attached. Certs look good. JLAB also run metal analysis on a sample screw and the meterial is found to be 316 SS. JLAB consulted Fermi on this issue. Fermi's comments are "Graded fasteners are required to have specific markings, much like lifting hardware. Non-graded fasteners which are required to follow some standard (e.g. ASTM) will also have markings, and it’s common industry practice to have markings. That being said, not all fasteners are required to have markings", "we didn’t specify any graded fasteners...nor invoke any fastener codes for the coldmass upper assembly".
36 | 2021-07-13 | 2021-08-11 | L2HE-CWI-VV | VESSEL AFTER MACHING WELDMENT | RFCHE-VV-002 | The fillet weld located on weld#C1 have been reduced in size via some machining operation. The welds are not complaint with the drawing, specifications or the welding code. It is my intention to reject the fillet weld portion of this flange weld joint that also includes a CJP groove weld. However, I found the groove weld portion of this same flange weld joint to be in compliance. |  |  | Use As Is | After consulting Fermi on 8/11, decision has been made to accept VV002 welds as is.
37 | 2021-07-13 | 2021-10-06 | L2HE-VTA-CAV-VTRF | PRODUCTION CAVITY ASSEMBLY | 0331 | Cavity quench at 21.8 MV/m | Hold/Re-measure | We will look into re-rinse/re-test for this cavity | Use As Is
38 | 2021-07-31 | 2021-07-31 | L2HE-VTA-CAV-VTRF | PRODUCTION CAVITY ASSEMBLY | HE-331 | FE onset 5.5 MV/m, FE loading and 1R/hr at 14 MV/m | Hold/Re-measure | re-rinse https://bit.ly/3jbEoiG | Reject | Wait for HPR evaluation Re-test
39 | 2021-08-02 | 2021-08-10 | L2HE-CMA-FPCW-INSP | FPC WARM PART | CPI-FPCW-HE-001 | Found scratch on copper plating of DN40 flange |  |  | Return To Vendor
40 | 2021-08-02 | 2021-08-13 | L2HE-CMA-FPCW-INSP | FPC WARM PART | CPI-FPCW-HE-002 | Found scratch on copper plating of DN40 flange |  |  | Return To Vendor
41 | 2021-08-02 | 2021-08-13 | L2HE-CMA-FPCW-INSP | FPC WARM PART | CPI-FPCW-HE-007 | Found scratches on copper plating of DN40 flange |  |  | Return To Vendor
42 | 2021-08-02 | 2021-08-13 | L2HE-CMA-FPCW-INSP | FPC WARM PART | CPI-FPCW-HE-005 | Found scratch on copper plating of DN40 flange |  |  | Return To Vendor
43 | 2021-08-02 | 2021-08-13 | L2HE-CMA-FPCW-INSP | FPC WARM PART | CPI-FPCW-HE-006 | Push rod bellows distorted and slightly twisted. Push rod was difficult to remove and reinstall. The M6 antenna vented bolt was loose |  |  | Return To Vendor
44 | 2021-08-02 | 2021-08-13 | L2HE-CMA-FPCW-INSP | FPC WARM PART | CPI-FPCW-HE-009 | Face if DN100 flange scratched, small scratch on outer conductor copper plating. |  |  | Use As Is
45 | 2021-08-02 | 2021-08-13 | L2HE-CMA-FPCW-INSP | FPC WARM PART | CPI-FPCW-HE-010 | Push rod bellows very bent and twisted, some convolutions were pinched. Push rod was not engaged. |  |  | Return To Vendor
46 | 2021-08-02 | 2021-08-13 | L2HE-CMA-FPCW-INSP | FPC WARM PART | CPI-FPCW-HE-010 | RF contact area of inner conductor copper was pitted and copper plating had defects |  |  | Return To Vendor
47 | 2021-08-10 | 2021-09-15 | L2HE-INSP-IMAG | MAGNETIC SHIELD CAVITY ASSEMBLY | RFCHE-MS-028 | Part I.D.# F10048728 Rev.A Shield, LCLS-11 2nd layer center section A. has minor ding about 5/16 diameter and .010 deep. Part I.D.# F10048731 Shield, LCLS-11 2nd layer coupler side A. Has (2) sets of imperfections (1)is about 6" long the 2nd is about 4" long . | Hold/Re-measure | SN 28 is selected to carry out on-axis field mapping to see if the shielding performance is compromised by the imperfections | Use As Is | The shield's performance is not compromised hence it can be used as is
48 | 2021-08-10 | 2021-09-15 | L2HE-INSP-IMAG | MAGNETIC SHIELD CAVITY ASSEMBLY | RFCHE-MS-026 | Part I.D.# F10048728 Rev.A Shield, LCLS-11 2nd layer center section A. Has light ding 3/8 diameter and .010 deep. Part I.D.# F10048731 Rev.A Shield,LCLS-11 2nd layer coupler side A. Has 6" long imperfection. marked in red." | Hold/Re-measure | SN 26 is selected to carry out on-axis field mapping to see if the shielding performance is compromised by the imperfections | Use As Is | Shelding performance is not compromised so the shield can be used as is
49 | 2021-08-10 | 2021-09-24 | L2HE-INSP-HMFT | HOM FEEDTHROUGH | 2161 | Flange seal surface to probe tip is 29.5 mm, slightly out of spec which is 29.51-29.61 mm | Hold/Re-measure | Hold for now. Further evaluation is needed. | Use As Is | Per communication with SLAC, the antennas being 0.01 mm short is not likely to cause a problem in this case.
50 | 2021-08-12 | 2021-08-31 | L2HE-INSP-IMAG | MAGNETIC SHIELD CAVITY ASSEMBLY | RFCHE-MS-046 | Damage noted during visual inspection. Part ID#s F10048802 & F10048803 has a cut through the MAG Cavity turner end. See attached photos. Normal shield shown. |  |  | Conforming | This shield, SN 46, is determined to be a cavity 6 magnetic shield that suppose to have the cut
51 | 2021-08-13 | 2021-08-13 | L2HE-CMA-FPCW-INSP | FPC WARM PART | CPI-FPCW-HE-010 | One of the RF spring had come undone from itself |  |  | Return To Vendor
52 | 2021-08-13 | 2021-08-31 | L2HE-CMA-FPCW-INSP | FPC WARM PART | CPI-FPCW-HE-005 | Kapton sleeve missing covering the ceramic. | Hold/Modify | Install spare kapton sleeve prior to storage | Return To Vendor
53 | 2021-08-13 | 2021-08-31 | L2HE-CMA-FPCW-INSP | FPC WARM PART | CPI-FPCW-HE-006 | Kapton sleeve missing covering the ceramic. | Hold/Modify | Install spare kapton sleeve prior to storage | Return To Vendor
54 | 2021-08-13 | 2021-08-31 | L2HE-CMA-FPCW-INSP | FPC WARM PART | CPI-FPCW-HE-009 | Kapton sleeve missing covering ceramic. | Hold/Modify | Install spare kapton sleeve prior to storage | Return To Vendor
55 | 2021-08-13 | 2021-08-31 | L2HE-CMA-FPCW-INSP | FPC WARM PART | CPI-FPCW-HE-010 | Kapton sleeve missing covering the ceramic. | Hold/Modify | Install spare kapton sleeve prior to storage | Return To Vendor
56 | 2021-08-19 | 2021-08-20 | L2HE-VTA-CAV-VTRF | PRODUCTION CAVITY ASSEMBLY | HE-040 | The lock frequency was below specification. It was 1299.966 MHz. |  |  | Reject
57 | 2021-08-27 | 2021-10-19 | L2HE-CMA-VV-INSP | VESSEL AFTER MACHING WELDMENT | RFCHE-VV-004 | Paint present inside of groove of flange and on face of flanges. Drawing number from picture identifies flange. Please see attached. | Hold/Modify | Having discussed with cryomodule assembly team lead, we will review the visual inspection findings and make corrections as needed. Vendor and Fermi are noted of the issues for improvements on future vessels. | Conforming | Repair work has been done. We will keep track of this VV to see if leak would develop at final assembly level.
58 | 2021-08-27 | 2021-10-19 | L2HE-CMA-VV-INSP | VESSEL AFTER MACHING WELDMENT | RFCHE-VV-004 | Scratches and pitting present on multiple flanges. Please see attached pictures. Drawing number from pictures identify flange with issue. | Hold/Modify | Having discussed with cryomodule assembly team lead, we will review the visual inspection findings and make corrections as needed. Vendor and Fermi are noted of the issues for improvements on future vessels. | Conforming | Repair work has been done. We will keep track of this VV to see if leak would develop at final assembly level.
59 | 2021-09-02 | 2021-09-03 | L2HE-VTA-CAV-VTRF | PRODUCTION CAVITY ASSEMBLY | HE-295 | Cavity lock frequency 50 kHz below target range |  |  | Use As Is
60 | 2021-09-02 | 2021-09-03 | L2HE-VTA-CAV-VTRF | PRODUCTION CAVITY ASSEMBLY | HE-295 | Qext 1 slightly above allowable range, Qext2 slightly below allowable range (per Traveler criteria) |  |  | Use As Is
61 | 2021-09-08 | 2021-09-17 | L2HE-INSP-BLBU | WELDMENT BELLOWS - US END | 007 | Dim2 & Dim9 are OOT. See attached IR. Dim9 is multiplied (x2) to get the 96.0 dimension. |  |  | Use As Is | OOT values are for bellows only, are relatively small, and will not effect the flange features.
62 | 2021-09-09 | 2021-09-17 | L2HE-INSP-BLBU | WELDMENT BELLOWS - US END | 001 | Dim2, Dim3, & Dim9 are OOT. Dim9 is multiplied (x2) to get the 96.0 dimension. |  |  | Use As Is | OOT values are for bellows only, are relatively small, and will not effect the flange features.
63 | 2021-09-09 | 2021-09-17 | L2HE-INSP-BLBU | WELDMENT BELLOWS - US END | 004 | Dim2 is OOT. See attached report. |  |  | Use As Is | OOT values are for bellows only, are relatively small, and will not effect the flange features.
64 | 2021-09-09 | 2021-09-17 | L2HE-INSP-BLBU | WELDMENT BELLOWS - US END | 010 | Dim2 & Dim3 are OOT. See attached IR. |  |  | Use As Is | OOT values are for bellows only, are relatively small, and will not effect the flange features.
65 | 2021-09-14 | 2021-09-16 | L2HE-VTA-CAV-VTRF |  | 0128 | Cavity HE-CAV128 Qext2 was low during the decay measurement. |  |  | Use As Is
66 | 2021-09-15 | 2021-09-21 | L2HE-CWI-VV | VESSEL AFTER MACHING WELDMENT | RFCHE-VV-004 | The internal CJP groove weld on Weld Joint #C1 is acceptable (*see weld-map). However, the external flange intermittent fillet welds were reduced in leg size. That reduction in fillet weld size is below an acceptable size in accordance with the drawings and specifications. |  |  | Use As Is | Having consulted Fermi Lab SOTR, this weld NCR is deemed non-critical
67 | 2021-09-16 | 2021-09-20 | L2HE-CMA-UCM-INSP | UPPER COLD MASS ASSEMBLY | RFCHE-UC-004 | F10074420 Line D vertical supports twisted/bent. Please see attached picture. |  |  | Use As Is | On UC002, similar issues are found on Line D supports. Fermi and the vendor are both noted. Fermi advises to use as is. The same approach is adopted for UC004
68 | 2021-09-16 | 2021-09-21 | L2HE-CMA-UCM-INSP | UPPER COLD MASS ASSEMBLY | RFCHE-UC-004 | Cold mass assembly is missing two F10043365 5k clamps. | Hold/Re-measure | Suspect that cryomodule assembly group took a couple 5K clamps off to work on re-application of indium foils, let's double check | Conforming
69 | 2021-09-21 | 2021-11-03 | L2HE-CMA-UCM-INSP | UPPER COLD MASS ASSEMBLY | RFCHE-UC-004 | Indium foil applied to wrong half. | Hold/Modify | 1. Check/disassemble all 16 copper clamps. 2. Apply indium foil to the long halves on the clamps. No need to remove the indium on the short half. 3. Re-assemble clamps, torque three times with 8-hr intervals. | Conforming | ready for use
70 | 2021-09-27 | 2021-10-19 | L2HE-CMA-VV-INSP | VESSEL AFTER MACHING WELDMENT | RFCHE-VV-002 | Visual inspection performed on 17-Sep-2021. All photos on the server: M:\cmm\Pictures\Inspection Pictures\L2HE\Vacuum vessel\VV002\Receiving Inspection | Hold/Modify | Having discussed with cryomodule assembly team lead, we will review the visual inspection findings and make corrections as needed. Vendor and Fermi are noted of the issues for improvements on future vessels. | Conforming | Repair work has been done. We will keep track of this VV to see if leak would develop at final assembly level.
71 | 2021-10-17 | 2021-10-18 | L2HE-VTA-CAV-VTRF |  | HE-331 | Pi mode frequency 1300.1297 MHz is lower than min frequency specifications (1300.150MHz). |  |  | Use As Is | The acceptance criteria has been updated. This will be reflected in the future traveler.
72 | 2021-10-20 |  | L2HE-INSP-CFL | CURRENT LEADS CF FLANGE ASSEMBLY | 0043 | All CL Flange lead wires are 12 inches, not 15 inches. According to FNAL Drawing F10022642-Rev G, under Table 1, Item 10, part number FC0062680 should be 15 inches long. Every CL Flange showed up with 12 inch long lead wires instead of the proper 15 inch long lead wires.
73 | 2021-10-25 |  | L2HE-INSP-BLBU | WELDMENT BELLOWS - US END | 001 | Slotted flange seal area has several small dings and light scratches on the bottom. | Hold/Modify | Deliver to machine shop for polishing of seal surfaces.
74 | 2021-10-25 |  | L2HE-INSP-BLBU | WELDMENT BELLOWS - US END | 002 | Slotted flange has light scratches and small dings on bottom of seal surface. | Hold/Modify | Deliver to machine shop for polishing of seal surfaces.
75 | 2021-10-25 |  | L2HE-INSP-BLBU | WELDMENT BELLOWS - US END | 003 | Slotted flange has several light scratches and (2) small dings. | Hold/Modify | Deliver to machine shop for polishing of seal surfaces.
76 | 2021-10-25 |  | L2HE-INSP-BLBU | WELDMENT BELLOWS - US END | 004 | Slotted flange has several light scratches on bottom of seal surface area. | Hold/Modify | Deliver to machine shop for polishing of seal surfaces.
77 | 2021-11-02 |  | L2HE-INSP-BLBS | WELDMENT BELLOWS SHORT | 002 | Please see attached for dimensional non-conforming.
78 | 2021-11-03 |  | L2HE-INSP-BLBP | WELDMENT BELLOWS PRCM | 001 | Please see attached for dimensions out of tolerance.