## Sheet1
TRAV_SEQ_NUM | CAVSN | SPECIALHANDLING | HELIUMVESSEL | TUNERATTACHED | HOMSCONNECTED | LABVIEWFILE | TESTDATE | DEWAR | VTATSSN | TESTOPERATOR1 | TESTOPERATOR2 | CAVITYVACUUM | CAVITYVACUUMOK | CAVITYVACUUMCOMMENT | DEWARLHELEVELCM | DEWARTEMPK | DEWARPRESSURETORR | POWERMETERZEROED | CABLECALIBRATIONOK | CABLECALIBRATIONCOMMENT | FREQ_7_7PI | FREQ_6_7PI | FREQ_5_7PI | FREQ_4_7PI | FREQ_3_7PI | FREQ_2_7PI | FREQ_1_7PI | TESTOPERATORHOMS | TESTOPERARHOMS_OTHER | HOMSFILE | HOMSCOMMENT | MEETS SPECS | LOCKFREQ | DEWARPRESS | LOCKFREQMEETSPEC | CAVITYCOUPLING | TDS_TXT | EEACC | Q0 | QEXTIN | QERXTFP | QEXTFPERROR | RAD | QEXTHOMA | QEXTHOMB | EMAXMVM | Q0ATEMAX | Q0AT20MVM | INIT_RADAT20MVM | INIT_FEONSETMVM | RADMAX | QOAT4MVM | PERFLIMITAT2_07K | PERFLIMITAT2_07K_OTHER | RFPROCESSING | RFPROCESSINGCOMMENT | KLOREN | KLORENMEETSSPEC | KLORENCOMMENT | PP_RADAT20MVM | PP_FEONSETMVM | PP_RADMAX | EACCUNSCALEDQUENCH_6_7PI | EACCUNSCALEDQUENCH_5_7PI | EACCUNSCALEDQUENCH_4_7PI | EACCUNSCALEDQUENCH_3_7PI | EACCUNSCALEDQUENCH_2_7PI | EACCUNSCALEDQUENCH_1_7PI | QUENCHSTUDYCOMMENT | Q0AT20MVM_2_07K | Q0AT20MVM_2_02K | Q0AT20MVM_1_99K | Q0AT20MVM_1_93K | CAVITYPERFORMANCE | RF_TESTRAWDATA | RF_TESTPROCESSED | Q0VSEACC | RADVSEACC | HOMAANDHOMBVSEACC | FVSEACC2 | UPLOADFILES | INFORMHOMSCHECKERS | INFORMHOMSCHECKERSHPNAME | INFORMAHOMSCHECKERSHPDATE | CAVITYMEETSSPECIFICATIONS | DEWAR_PRESSURETORR | CAVITYVACUUMTORR | DEWARHELEVEL | RADAT20MVM | FEONSETMVM | CAVITYID_IF_NOT_C100 | HOM_CONNECTED | MHA_7_7PI | MHA_6_7PI | MHA_5_7PI | MHA_4_7PI | MHA_3_7PI | MHA_2_7PI | MHA_1_7PI | TEST_OPERATORHOMS | HOM_F1_1886MHZ | HOM_QL1 | HOM_QL1 | HOM_QL1 | HOM_QL1 | HOM_QL1 | HOM_QL1 | HOM_QL1 | HOM_QL1 | HOM_QL1 | [25 empty] | HOM_FILE | LOCKFREQMHZ | DEWARPRESSTORR | TDSTXT | R_AT20MVM | RADMAX_MRH | Q020MVM | PULSEPROCESSING | Q0_AT_2_07K | Q0_AT_2_02K | Q0_AT_1_99K | Q0_AT_1_93K
1 | C100-1 |  | No | No |  | C100_E.txt | 2010-06-02 18:34 | 8 | TS28-007 | GRENOBLE | REECE |  |  |  | 170 | 2.07 | 29 | Yes | Yes | [15 empty] | No | Overcoupled |  | 2.74 | 15000000000 | 3500000000 | 990000000000 |  | 0 | 2200000000000 | 1000000000000 | 25.4 | 11800000000 | 11600000000 |  |  |  | 13300000000 | RF power | [21 empty] | 04D618DB-FE5A-7E50-56CCEBB47324F826 | [11 empty] | 0 |  |  | 20 |  | 2 | 1496.8493 | 1495.6338 | 1492.3839 | 1487.6461 | 1482.6795 | 1477.9729 | 1474.8173 | Mircea |  | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0E1E3201-9AF3-31E5-0529FC86AEE8CE76 | 1496.8441 | 2.07 | FF525A1A-DB57-7FC6-6AD44A1A54494698 | 0.01 | 0.01 | 12300000000 | No | 12300000000 | 14000000000 | 14700000000 | 117600000000
4 | C100-RI-002 |  | No | No | 2 | C100_RI_002.txt | 2010-09-29 | 4 | TS28-003 | Grenoble | KDavis |  |  |  |  | 2.07 |  | Yes | Yes |  | 1496.656 | 1495.44 | 1492.319 | 1487.892 | 1482.942 | 1478.447 |  | Marhauser | Grenoble | 70EA2594-E9EE-ACF1-AAB1CC394C8F3CBD | F.M.: First measurement before chemistry upon arrival (no issues), but tilted TM111 pi/7-mode pair measured on bench, luckily to HOM-damper side, therefore lower Qs than usual (actually not good if modes are tilted). Second measurement by Christiana. One high Q mode probably faulty. Else slightly different Qs compared to first measurement, but all below BBU threshold. | Yes | 1496.6558 | 29 | Yes | Overcoupled | 69589BB3-D12E-14AD-7EEC60F98B9E3CDD | 5.11 | 19200000000 | 5840000000 | 1240000000000 | 15 | 0.009 | 2400000000000000 | 38700000000000 | 27.1 | 9870000000 | 12600000000 |  |  | 677 | 16600000000 | FE | Admin limit of 27 MV/m when Rad > 100mR/hr | Yes | FE onset at 14.7 MV/m following processing | [13 empty] | 12600000000 | 13400000000 | 14400000000 | 16900000000 |  | 695E28C4-A5B5-FE50-CC96A0763E3C2D3B | 695E8D86-A6A2-4261-79B84FCBA7F4140B | 695EDB90-A722-83DC-E14CC24FCC818D35 | 695F460B-0B84-4614-33F649E7CBB03F75 | 695F8B15-9E33-ECA0-B0F814E2D0D33688 | 695FCC95-B225-C329-EAADB223CC8C3E64 | 8460C398-9352-AF04-E73A23D4D4014F01 | 1 | HAIPENG | 2012-02-22 15:25 | Yes | 29 | 4e-09 | 157 | 5.97 | 22.5
5 | C100-RI-003 |  | No | No | 2 | C100_RI_003.txt | 2010-10-09 | 8 | TS28-001 | STIRBET | KDAVIS |  |  |  |  | 2.07 |  | Yes | Yes |  | 1496.6296 | 1495.4529 | 1492.3135 | 1487.8042 | 1482.6832 | 1478.2783 | 1475.1683 | Grenoble |  | 9BEFC258-EACF-9598-C0842D58380C137F | HOMs meet spec., but cable on HOM A was loose as detected during tests. After warm-up it was found that the loose connection was at the tophat feedthrough. The HOM measurements need to be repeated once the HV is welded to the cavity. | Yes | 1496.6287 | 29 | Yes | Overcoupled | 9BF0C7CB-EA3A-64A3-7111F1CBE0E01F0D | 5.1 | 12900000000 | 10100000000 | 1480000000000 | 11.6 | 0.003 | 300000000000000 | 25000000000000 | 27 | 9230000000 | 10400000000 |  |  | 795 | 13100000000 | FE | [16 empty] | 10400000000 | 11900000000 | 12100000000 | 13200000000 |  | 9BF78B78-DDD0-12DE-8279C21856F17619 | 9BF7D19A-9AA7-C888-2B28BB8D69BE8AF0 | 9BF829E8-B182-5DFA-4BA65DE905F0DB01 | 9BF87DC8-EAC7-CAB2-D9D2A307712D8BC7 | 9BF8E109-B8CF-C97D-FA6B7B000F20EBE3 | 9BF936CB-B768-A626-7629289D5055203E | 9BF972CF-0809-85B9-AF7FC23340FF55AA | 1 | HAIPENG | 2012-02-24 12:38 | Yes | 29 | 7.5006e-10 | 191.23 | 8.76 | 15.5
6 | C100-RI-004 |  | No | No | 0 | C100_RI_004.txt | 2010-10-14 18:25 | 8 | TS28-008 | KDAVIS | KDAVIS |  |  |  |  | 2.07 |  | Yes | Yes |  | 1496.6547 | 1495.4727 | 1492.3159 | 1487.8395 | 1482.8617 | 1478.3112 | 1475.1513 | Mircea | Marhauser | C48AA6D7-E2C5-34B8-F56781F09EF0E572 | no issues | Yes | 1496.653 | 29 | Yes | Overcoupled |  | 3.82 | 11900000000 | 10200000000 | 924000000000 | 11.7 | 0.003 |  |  | 32.5 | 8010000000 | 9500000000 |  |  | 53.7 | 11900000000 | Cable | Subsequent testing was prevented by persistent helium breakdown due to cavity superfluid leak | No | [14 empty] | 9550000000 | 10200000000 | 10500000000 | 11500000000 |  | CB88B87A-A264-9903-C6FE536FAAB867C2 | CB92FF44-E2CA-F5C7-FADB75F9B8B90B4F | CB93A701-C0E1-DA40-0780F331F355BAEC | CB93FD1D-012F-C54C-D2B0574F77C27F33 |  | CB94AEA6-B112-E623-6E59D6342146A99B | 1C4E1E71-BBEE-2E3D-2905030A0F131961 | 1 | MARHAUSE | 2010-10-20 21:48 | Yes | 29 | 6.5e-06 | 187 | 0.00194 | 20.5
7 | C100-RI-005 |  | No | No | 2 | C100_RI_005.txt | 2010-10-18 | 8 | TS28-007 | KDAVIS | STIRBET |  |  |  |  | 2.07 |  | Yes | Yes |  | 1496.6643 | 1495.5314 | 1492.355 | 1487.9174 | 1482.9054 | 1478.461 | 1475.2634 | Mircea | Marhauser | EF19DFE6-9F40-10F6-B67553FC5CAC2F4A | problems with cable on HOM B. Either pin retraction or HOM feedthrough problem. Needs to be inspected upon warm-up. Cavity HOMs to be remeasured with HV. | Yes | 1496.6643 | 29 | Yes | Overcoupled | E9C9D9F0-0DDF-5D18-5E7C2CF5A9F4AEA1 | 3.81 | 13300000000 | 6240000000 | 1030000000000 | 12.6 | 0.006 |  |  | 27 | 9660000000 | 10800000000 |  |  | 220 | 13400000000 | Other | admin limit 100mR/hr at 27 MV/m or greater field | Yes | post processing FE onset was at 16 MV/m | [13 empty] | 10800000000 | 11600000000 | 12100000000 | 13200000000 |  | E9E38FD2-B1D2-A4F3-F554C53707BEE83F | E9E52F7C-9B74-508F-EBE9D962FF24A1E3 | E9E9191B-C814-6084-59D722FD58E040C5 | E9E96E1F-F189-2A16-1B2D9A4EF6F03008 | E9E9CD0E-F486-994C-1D22FB05F7795810 | E9EA21AB-E8B3-9C38-61442E65BCE9C2C9 | E9EAA78A-0906-C4C7-D8E5ABE10E41EBBF | 1 | MARHAUSE | 2010-10-26 15:06 | Yes | 29 | 3.8e-07 | 195 | 1.14 | 21.5
8 | C100-RI-006 |  | No | No | 2 |  |  | 8 |  | KDAVIS | KDAVIS | [16 empty] | Grenoble | Marhauser | D481330E-0EBB-E1EB-38CC071593EC9885 | no issues. The high power tests of this cavity at this date (2010-10-21) had to be postponed due to a superfluid leak (2e-4mbar) (F.M.) | Yes |  |  | Yes | Overcoupled | [45 empty] | 1 | HAIPENG | 2012-02-24 13:31 | Yes
9 | C100-RI-006 |  | No | No | 2 | C100_RI_006A.txt | 2010-11-03 | 7 | TS28-011 | GRENOBLE | KDAVIS |  |  |  |  | 2.07 |  | Yes | Yes |  | 1496.6838 | 1495.499 | 1492.3583 | 1487.82 | 1482.726 | 1478.1389 | 1475.001 | Mircea | Grenoble | 177DAA6F-EB9A-A9F8-D3AA5790AA208136 | no issues | Yes | 1496.6819 | 29 | Yes | Undercoupled |  | 5.13 | 13800000000 | 14700000000 | 1320000000000 | 11.5 | 0.003 | 90000000000000 | 9000000000000 | 41.6 | 9240000000 | 10800000000 |  |  | 9.02 | 14200000000 | Quench (non FE) |  | No | [14 empty] | 9990000000 | 11100000000 | 11600000000 | 13100000000 |  | 17776C65-C412-9273-8245FE710BADEB2A | 1777F889-FE29-62CD-2676C45C75AC4E40 | 17785BAE-E6AC-5482-0466BEB9D0A4ECAD | 1778AE2B-F1DC-8238-DF4D63613C07C593 | 17790D8A-98DB-2B08-772BB1CE5902B83E | 17795D22-F82E-A3CC-D52E904983B93C5B | 1779D03A-0231-1EE0-44C46DF6DE04F711 | 1 | MARHAUSE | 2010-11-04 12:04 | Yes | 29 | 4.9e-09 | 203 | 0.00314 | 30
10 | C100-RI-008 |  | No | No | 2 | C100_RI_008.txt | 2010-11-21 | 4 |  | KDAVIS | Marhauser | [9 empty] | 1496.6149 | 1495.4128 | 1492.264 | 1487.7193 | 1482.7782 | 1478.1758 | 1475.0678 | Marhauser |  | 7523E0CC-CB08-499D-41275F14E8F4DD28 | no issues | Yes |  |  |  | Overcoupled | [45 empty] | 1 | HAIPENG | 2012-02-24 13:35 | Yes
11 | C100-RI-008 |  | No | No | 2 | C100_RI_008.txt | 2010-11-21 | 4 | TS28-004 | KDAVIS | KDAVIS |  |  |  |  | 2.07 |  | Yes | Yes |  | 1496.6178 | 1495.4155 | 1492.2667 | 1487.722 | 1582.7809 | 1478.1784 | 1475.0706 | Marhauser |  | 7524CC93-C4AC-639A-1D7DC0E8ACAB537A | no issues. | Yes | 1496.6155 | 29 | Yes | Overcoupled |  | 5.45 | 15400000000 | 8800000000 | 882000000000 | 12 | 0.003 | 60000000000000 | 2000000000000 | 12 | 8860000000 |  |  |  | 644 | 15700000000 | Other | FE loaded at 12 MV/m | Yes | Radiation spike at 17 MV/m caused reduced performance. Subsequent RF processing to recover performance was ineffective. | [18 empty] | 7527FF45-DA82-543C-79080E6BAF78CE24 | 752858FA-0374-11AA-5AD620B46012D560 | 752893B8-D479-3E04-493215A8BE1CE042 | 7528CA9B-9107-A897-822260B040C74B9A | 752905D0-E2BF-BE57-EF8551BD62BFCC02 | 752940C7-EC7F-7051-DFE226B0CF197711 |  | 1 | MARHAUSE | 2010-11-22 15:02 | No | 29 | 5e-09 | 168 |  | 6
12 | C100-RI-007 |  | No | No | 2 | C100-RI-007.txt | 2010-11-30 18:34 | 8 | TS28-008 | KDAVIS | KDAVIS |  |  |  |  | 2.07 |  | Yes | Yes |  | 1496.6703 | 1495.4895 | 1492.3373 | 1487.7752 | 1482.6851 | 1478.1015 | 1474.9384 | Grenoble |  | A767D0A3-0B40-EE75-A32A9A1A7BF772D8 | no issues (F.M) | Yes | 1496.6684 | 29 | Yes | Overcoupled | C27AFFF7-ABB9-2D26-3AF04B82C3F0B3D5 | 4.86 | 12600000000 | 10200000000 | 505000000000 | 11.7 | 0.003 |  |  | 26.46 | 7000000000 | 10000000000 |  |  | 909 |  | Other | Test aborted due to cavity vacuum leak. Helium discharging at 1W Pinc by end of test. | No | [19 empty] | C2814FF5-F7CD-1068-60467ED19585AB32 | C28294C7-DC3D-558B-5E4DE07AB5E18E45 | C283653A-CEA7-D5B9-2E88F55413173370 | C283BA6E-E20A-60B9-5E85EAB7D84EC045 |  | C2849DF5-D90A-2610-C0C672B0EC2D3ED2 | 1C3405B1-0356-38E1-78CF74B5200552D6 | 1 | MARHAUSE | 2010-12-07 17:50 | Yes | 29 | 1.1e-05 | 183 | 0.125 | 18.6
13 | C100-RI-005 |  | Yes | No | 0 | C100_RI_005A.txt | 2010-12-02 10:06 | 7 | TS28-001 | GRENOBLE | GENG |  |  |  |  | 2.07 |  | Yes | Yes |  | 1496.4188 | 1495.2805 | 1492.0923 | 1487.6467 | 1482.6557 | 1478.2034 | 1474.9862 | Grenoble |  | B201DD24-ADEA-C6AC-20D6078D2BF1EBDA | no issues | Yes | 1496.4192 | 29 | Yes | Overcoupled |  | 5.4 | 13900000000 | 11100000000 | 892000000000000 | 26 | 0.003 |  |  | 27.2 | 9400000000 | 10100000000 |  |  | 935 |  | Cable | FP cable short at tophat feedthrough | No | [19 empty] | B2090025-F00C-7080-748A0846A939A75F | B209B54D-A140-4DFA-CA16126FCB4FF110 | B20AA782-98A7-7508-67BF63FFD313656C | B20B6E23-D3C4-B808-F05FE5220C775BF7 |  | B20C72EA-A57C-BEBD-9E3534CDD1CFBA51 | B20CC0AA-9582-F162-736149E4D90A70B8 | 1 | MARHAUSE | 2010-12-06 08:27 | No | 29 |  |  | 0.003 | 20.2
14 | C100-RI-004 | previous test was FE loaded at 15 MV/m; cavity was high-pressure rinsed subsequently | Yes | No | 2 | C100-RI-004B.txt | 2010-12-05 09:43 | 4 | TS28-004 | GRENOBLE | KDAVIS |  |  |  |  | 2.07 |  | Yes | Yes |  | 1496.4406 | 1495.2664 | 1492.0907 | 1487.6134 | 1482.6286 | 1478.0713 | 1474.8919 | Other | na |  | previously performed |  | 1496.4389 | 29 | Yes | Overcoupled | BC2A9378-0033-FD3C-886F9FCE6B901987 | 5.73 | 15600000000 | 11500000000 | 892000000000 | 11.7 | 0.001 | 5000000000000 | 20000000000000 | 27.1 | 8400000000 | 10200000000 |  |  | 8.73 | 16600000000 | Other | admin limited to 27 MV/m | No | [14 empty] | 10300000000 | 11600000000 | 12200000000 |  |  | BC34F53D-DE6A-8858-27BDAF6AF43A1C28 | BC3547BF-F9E6-00B2-081B06B1911EFF47 | BC35AC74-BB02-EB39-AEF19B91A7445A1C | BC3612F1-D882-AEE7-FB457C3E0B2B0FC9 | BC366704-F248-5C4E-4A72CEDA131466B3 | BC36FA4F-C022-4B66-ACCDBE6E22743638 | BC379825-99C2-7D9C-4E7026F151339283 | 1 | MARHAUSE | 2010-12-06 11:02 | Yes | 29 | 1.1e-08 | 162 | 0.0904 | 18.5
15 | C100-RI-003 |  | Yes | No | 2 | C100_RI_003A.txt | 2010-12-04 10:06 | 8 | TS28-007 | KDAVIS | GRENOBLE |  |  |  |  | 2.07 |  | Yes | Yes |  | 1496.4033 | 1495.2348 | 1492.0784 | 1487.563 | 1482.5846 | 1478.0254 | 1474.895 | Grenoble |  | BCA7C768-FA99-1FBC-49E3806522552FB4 |  | Yes | 1496.402 | 29 | Yes | Undercoupled | BC3D7BF0-E2C8-E59F-FFC0AB8926B3907B | 4.33 | 11000000000 | 13800000000 | 963000000000 | 11.6 | 0.001 | 4000000000000 | 20000000000000 | 27 | 7480000000 | 8470000000 |  |  | 810 | 11200000000 | Other | admin limit 27 MV/m | No | [14 empty] | 8520000000 | 9200000000 | 9520000000 | 10200000000 |  | BC48176B-FBF6-A7E1-0FDC7EA9419EB87A | BC488D98-A65F-BC09-109A4A61DAE56A81 | BC48E4B3-E122-4823-AE9F4E707AC0C350 | BC4939B7-CAD2-DFDF-70F7E45317ED567A | BC49D899-C462-A94E-D917224613F57A80 | BC4A2847-FAE4-F25C-8DA1FBBC5C473775 | BC4A901B-F470-0EF7-F9D49BA14E196555 | 1 | MARHAUSE | 2010-12-06 11:13 | Yes | 29 | 3e-08 |  | 2.46 | 18.7
17 | C100-RI-005 |  | Yes | No | 2 | C100_RI_005D | 2010-12-06 | 7 | TS28-001 | STIRBET | KDAVIS |  |  |  |  | 2.07 |  | Yes | Yes |  | 1496.4224 | 1495.2844 | 1492.0962 | 1487.6505 | 1482.6588 | 1478.2069 | 1474.9898 | Other |  |  | HOM survey done on 2nd Dec 2010 (Christiana). |  | 1496.4203 | 29 | Yes | Overcoupled | C1731C2D-D6D5-CB1A-9E557F4612EE5F7F | 5.45 | 15000000000 | 11000000000 | 621000000000 | 11.7 | 0.003 | 70000000000000 | 200000000000000 | 27.2 | 11500000000 | 12100000000 |  |  | 579 | 15600000000 | Operator | Administrative limit 27 MV/m | No | [14 empty] | 12200000000 | 13500000000 | 14500000000 | 16600000000 |  | C17A7922-C0CD-BEC5-DC371852B569A358 | C17B0D9B-A5B9-C6C5-4C8ABF12C80FFC88 | C17B569A-949D-398C-F65EA1BEC42EB093 | C17BE4DA-0C13-712F-B6CFA748DAC7EB36 | C17C3CB7-C58D-C2AD-BB7FC1C306CE19D8 | C17C7ED4-9DA6-1598-89B50B86D391C2A9 | C17CD467-C432-EEEB-DABCE6BD647FF577 | 1 | HAIPENG | 2012-02-24 13:24 | Yes | 29 |  | 199.63 | 1.21 | 18
18 | C100-RI-006 |  | Yes | No | 2 | C100_006B.txt | 2010-11-24 15:31 | 3 | TS28-003 | GRENOBLE | KDAVIS |  |  |  |  | 2.07 |  | Yes | Yes |  | 1496.4774 | 1495.3035 | 1492.1354 | 1487.5984 | 1482.5022 | 1477.9189 | 1474.7444 | Other |  |  | HOM survey completed satisfactorily on 3 Nov 2011, traveler seq num 9. | Yes | 1496.475 | 29 | Yes | Overcoupled | 75683D96-BA5A-D75E-EAD974CF55403680 | 4.94 | 15200000000 | 10800000000 | 10900000000000 | 11.7 | 0.002 | 13000000000000 | 1200000000000000 | 30.4 | 11600000000 | 13700000000 |  |  | 108 | 15300000000 | Other | 100mR/hr radiation admin limit | No | [19 empty] | 75842EF7-D54F-1836-9827AB54F6171208 | 7584C07A-CA79-1731-CAFD4061E3CFA7E2 | 75850F4D-FB3B-3700-C9ED8DB44CAEE61A | 7585529A-B1B8-869B-BCBD189B8074BB4F | 7585BE4D-FBD4-1104-4CEF6C7565392281 | 7585FA41-B973-8959-08C591BE37CC444E | 1C0683E1-C83C-CC10-EF3F1C2051B898B5 | 1 | MARHAUSE | 2011-01-11 10:21 | Yes | 29 | 5e-09 | 160 | 0.106 | 13
19 | C100-RI-007 |  | Yes | No | 2 | C100_RI_007B.txt | 2010-12-11 15:18 | 8 | TS28-001 | KDAVIS | KDAVIS |  |  |  |  | 2.07 |  | Yes | Yes |  | 1496.3986 | 1495.2224 | 1492.0651 | 1487.4628 | 1482.3891 | 1477.7808 | 1474.6181 | KDavis |  | DFF8C88D-AE9C-0926-3FE5FF60D30995FB | Previously performed |  | 1496.3975 | 29 | No | Overcoupled | DC41FD6E-A829-319C-D307B4834BF7CAC4 | 5.2 | 14300000000 | 10800000000 | 647000000000 | 11.7 | 0.001 | 20000000000000 | 5000000000000 | 26.9 | 11400000000 | 11800000000 |  |  | 2.8 | 15000000000 | Other | Admin limit 27 MV/m gradient. | Yes | Processed multipacting barrier between 18 to 22 MV/m | [13 empty] | 11800000000 | 12800000000 | 13700000000 | 15400000000 |  | DC4C3A54-9F74-0709-58711FA77ACAF9EA | DC4CAE77-0350-8738-EBE4E60C2ED90C28 | DC4D1704-E6EE-5BB9-10D3F34A23C1A4B1 | DC4D6FEF-AE2C-8922-7FF3F7E54A2B82B9 | DC4DD124-F7CF-67B1-9FAE664C36628F27 | DC4E21EF-C20D-CF98-1259128376A2856E | DC4E856D-AE21-D179-17C95441525AF3B4 | 1 | MARHAUSE | 2010-12-14 15:10 | Yes | 29 | 6e-09 | 187 | 0.048 | 18.3
20 | C100-RI-012 |  | Yes | No | 2 | C100_RI_012.txt | 2010-12-10 19:21 | 7 | TS28-007 | KDAVIS | KDAVIS |  |  |  |  | 2.07 |  | Yes | Yes |  | 1496.252 | 1495.093 | 1491.948 | 1487.348 | 1482.407 | 1477.856 | 1474.65 | Mircea |  | DFFD1F25-9F1A-6FCC-D5E71DA115FEDC0E |  |  | 1496.2506 | 29 | No | Undercoupled | DC79EE88-90B9-E989-0F2BCFC7BC2A5831 | 5.09 | 12900000000 | 13900000000 | 596000000000 | 11.7 | 0.006 | 3000000000000 | 2000000000000 | 27 | 9350000000 | 10100000000 |  |  | 723 | 12800000000 | Other | Admin limited to 27 MV/m due to radiation > 100mR/hr | Yes | RF power processing in 3/7 and 2/7-pi modes. Little radiation in those modes. No improvement in radiation in pi mode finally. | [13 empty] | 10100000000 | 11200000000 | 11900000000 | 13200000000 |  | DC81590B-B304-87B1-1D7EF57DAFDD4B3D | DC81BF32-DFD4-E449-466B6C28F0651A15 | DC8222EC-E5E8-2D9B-635A4D9CB25E4013 | DC827ACE-B6CE-ECD6-AB66273B2F161068 | DC82C794-D6EA-9991-C61B6AF11F8222D0 | DC8312E8-AF4B-609C-C8425631A8B36D0E | DC836178-BAB2-EAE3-08BE82C90E28D4D5 | 1 | MARHAUSE | 2010-12-14 15:13 | Yes | 29 | 1.3e-08 | 192 | 3.96 | 18.7
21 | C100-RI-010 |  | Yes | No | 2 | C100_RI_010.txt | 2010-12-09 09:21 | 4 | TS28-003 | KDAVIS | KDAVIS |  |  |  |  | 2.07 |  | Yes | Yes |  | 1496.3714 | 1495.1814 | 1492.0551 | 1487.4291 | 1482.4085 | 1477.7758 | 1474.6868 | Mircea |  | E007141E-F930-A15B-6412ED29D92D59CE |  |  | 1496.3718 | 29 | No | Overcoupled | E544FB26-C70E-6ADA-9B51443EC8F7FB80 | 5.28 | 14600000000 | 11100000000 | 763000000000 | 11.7 | 0.001 | 13000000000000 | 750000000000000 | 27.03 | 10100000000 | 10800000000 |  |  | 0.6 | 13700000000 | Other | admin limited (gradient) to 27 MV/m | Yes | a few radiation bursts and quenches before reaching 27 MV/m. Then got a clean QvsE curve. | [13 empty] | 10900000000 | 12100000000 | 12600000000 | 13500000000 |  | E54F6DCA-9D71-2649-FEE479F01A1BD2CF | E5500203-9EFE-0BCE-B2A7BF723F88BD51 | E55053A1-9DBA-2DF3-BF5B96610D9A27ED | E550B15E-9B32-17ED-F57ECA5E4D4BE80A | E5512682-0F47-09AB-BD8D7E6E55223E15 | E5517691-E028-0011-FDC6C3DFA9BB041B | E551DAA8-9035-9817-5949182CFA3B3A1A | 1 | MARHAUSE | 2010-12-14 15:27 | Yes | 29 | 6.8e-09 | 169 | 0.00907 | 18
22 | C100-RI-011 |  | Yes |  | 2 |  |  | 8 |  | KDAVIS | KDAVIS | [16 empty] | Mircea |  | E00953E5-F256-E84D-3808DAE1F8BA5B78 | [5 empty] | Overcoupled | [45 empty] | 1 | HAIPENG | 2012-02-24 13:59 | Yes
23 | C100-RI-011 |  | Yes | No | 2 | C100_RI_011a.txt | 2010-12-13 15:06 | 4 | TS28-004 | KDAVIS | KDAVIS |  |  |  |  | 2.07 |  | Yes | Yes |  | 1496.438 | 1495.289 | 1492.146 | 1487.558 | 1482.593 | 1478.073 | 1474.963 | Mircea |  | E512B9C0-D67F-350F-CC0C8D0C9EE5C4A1 | HOM Survey performed previously | Yes | 1496.4364 | 29 | Yes | Overcoupled | E502186B-BD04-05ED-7FD2AC5E5E4D5EB8 | 6.06 | 19400000000 | 12100000000 | 783000000000 | 11.9 | 0.001 | 600000000000000 | 8000000000000 | 26.96 | 11700000000 | 12900000000 |  |  | 0.185 | 20600000000 | Other | admin limited to 27 MV/m (gradient) | No | [14 empty] | 13100000000 | 15400000000 | 16600000000 | 19400000000 |  | E5088FCD-FF6F-AE87-280C7D2D0B538688 | E508EBA1-CC76-F6F6-6B745F8603FCA6CA | E5094984-D4E3-E70C-B4DCFF54C12C98DD | E509A88A-D218-D5E9-A0CF197F195AAF04 | E509F681-0D43-20B8-800E3E315E6BDF50 | E50A6C64-08B7-0BAC-DEB8A25E8DD26C89 | E50AC670-B9C2-DCC6-F8E457C126EEAFFE | 1 | MARHAUSE | 2010-12-14 15:15 | Yes | 29 | 1.1e-08 | 163 | 0.0031 | 24
24 | C100-RI-003 |  | Yes | No | 2 | C100_RI_003A.txt | 2010-12-14 17:05 | 8 | TS28-007 | KDAVIS | GRENOBLE |  |  |  |  | 2.07 |  | Yes | Yes |  | 1496.403 | 1495.235 | 1492.0784 | 1487.563 | 1482.585 | 1478.025 | 1474.895 | Grenoble |  | E598E028-A491-5F60-67A5395585015804 |  | Yes | 1496.402 | 29 | Yes | Overcoupled | E55D7E9E-C324-BB7A-F66D76716A3A2AE9 | 4.33 | 11000000000 | 13800000000 | 963000000000 | 11.6 | 0.001 | 4000000000000 | 20000000000000 | 27 | 7480000000 | 8470000000 |  |  | 833 | 11200000000 | Other | admin limited to 27 MV/m due to radiation >100mR/hr | Yes | radiation bursts and lost lock several times as power increased up to 27 MV/m | [13 empty] | 8520000000 | 9200000000 | 9520000000 | 10200000000 |  | E5670C68-BC32-5BD5-B94DC0D3D8DBF3EF | E56781D0-FED7-0A8C-8C7F27D447A3E743 | E567D9AE-DEAC-4C61-24448CADFF3FB182 | E5683C83-D563-07BA-9178F281CFFFA89B | E568931E-B132-525F-2960015D612F6477 | E568E8D4-EFAF-CC10-91C684A78A649221 | E5693FC0-F275-BEDD-64360F50C7E48462 | 1 | MARHAUSE | 2010-12-14 15:33 | Yes | 29 | 3e-08 | 163 | 2.46 | 18.7
25 | C100-RI-005 | Field Cable short in or near the top hat. Test performed for field emission measurements only. | Yes | No | 2 | C100_RI_005A.txt | 2010-12-02 15:52 | 7 | TS28-001 | Grenoble | KDavis |  |  |  |  | 2.07 |  | Yes | Yes |  | 1496.4188 | 1495.2805 | 1492.0923 | 1487.6467 | 1482.6557 | 1478.2034 | 1474.9862 | Grenoble |  | 4C7168D0-ABF0-2369-9B7C4D15C8D0325B |  | Yes | 1496.4192 | 29 | Yes | Overcoupled |  | 5.4 | 13900000000 | 11100000000 | 892000000000000 | 26 | 0.003 |  |  | 27.2 | 9390000000 | 10100000000 |  |  | 935 |  | Cable | field probe cable short in top hat. Test performed for measuring field emission only. | No | [19 empty] | E6B7F634-ADC8-4C17-A95F8ABB69ECE2CB | E6B88320-EE10-5E87-89708BCB186A7101 | E6B9DF64-F689-2643-4A7424F75655BD7B | E6BA0B2D-975D-54AB-4FBC3443C109EC5F |  | E6BA3EA4-FC41-5799-2BCE825B9FD9AE92 |  | 1 | MARHAUSE | 2010-12-14 16:14 | No | 29 |  |  | 0.00296 | 20
26 | C100-RI-008 | HPR since previous RF test | Yes | No | 2 | C100_RI_008a | 2011-01-18 18:02 | 8 | TS28-007 | KDAVIS | KDAVIS |  |  |  |  | 2.07 |  | Yes | Yes |  | 1496.607 | 1495.404 | 1492.258 | 1487.711 | 1482.767 | 1478.162 | 1475.058 | KDavis |  |  | HOM Survey previously completed satisfactorily traveller seq no 11 | Yes | 1496.66 | 29 | Yes | Overcoupled | 9EAB312D-F169-A7D1-EA334FB949E8E101 | 4.02 | 10000000000 | 15000000000 | 820000000000 | 11.8 | 0.003 | 6000000000000 | 2000000000000000 | 8 | 4510000000 |  |  |  | 647 | 10000000000 | Other | FE loaded at 7 MV/m | Yes | attempted to process through radiation barrier unsuccessfully | [18 empty] | 9EB49203-C267-B160-5D57BB31D18A7223 | 9EB4F22B-9FF8-6432-B0C5CFAF02EF363E | 9EB54AEB-F2A6-D3F7-29AD8EF03F360775 | 9EB59FCD-EA64-BB1E-6B36F47FEE0BE963 | 9EB5F9FF-D4A6-6252-551B510B81EFE285 | 9EB65FE9-0B61-1CE4-71EAA83B1B4E3205 | 9EB6B208-CD9C-C70D-DA6E31CD0B921343 | 1 | HAIPENG | 2012-02-24 13:36 | Yes | 29 | 8.3e-09 | 188 |  | 5
28 | [20 empty] | 123 | 654 | 879 |  |  |  |  | KDavis | KDavis | B9B16E5A-B33A-84B9-991D46255E60DB32 | [5 empty] | Overcoupled | B9B1A294-AD71-9731-71A60AACFFE3F30F | [36 empty] | teteste | [10 empty] | No
29
30 | C100-RI-014 |  | Yes | No | 2 | C100_RI_014 | 2011-01-22 18:03 | 8 | TS28-001 | KDAVIS | KDAVIS | 4e-08 |  |  | 177 | 2.07 |  | Yes | Yes |  | 1496.334 | 1495.169 | 1491.998 | 1487.504 | 1482.491 | 1477.993 | 1474.868 | KDavis | KDavis | B97139FB-0408-187E-30E06F6C46F8F3E6 | test |  | 1496.3347 | 29 | No | Overcoupled | B972246A-D5F6-9CEE-3E11D4546F9A32F4 | 4.65 | 12300000000 | 15600000000 | 613000000000 | 11.7 | 0.002 | 4000000000000 | 5000000000000 | 27.1 | 8680000000 | 9330000000 | 0.00693 | 19.2 | 191 | 11600000000 | Other | admin limited to 27 MV/m | Yes | several quenches between 25 and 27 MV/m | -5.2 |  |  | 1 | 15 | 172 | [7 empty] | 9440000000 | 10200000000 | 10700000000 | 11800000000 |  | B97F94E0-B89C-ACF0-16C12F1710214485 | B980A407-B649-A6C6-7757B6943599D836 | B980FBB6-C95C-D2B4-2F05D0F26E529329 | B98160B1-9449-266E-C2386BBD8B0A5FA5 | B981B12A-A2B0-44B3-616AEA7F57F1D57C | B9820061-B199-CFA3-92623061253C4A20 | B9826B45-CD5B-75EB-BE1E6754F596B72C | 1 | MARHAUSE | 2011-01-24 14:57 | Yes | 29
31 | C100-RI-013 | the four line valve was closed for about one hour before testing | Yes | No | 2 | C100_RI_013.txt | 2011-01-28 09:51 | 8 | TS28-007 | GRENOBLE | KDAVIS | 2.1e-07 |  |  | 181.94 | 2.07 |  | Yes | Yes |  | 1496.4415 | 1495.2442 | 1492.0858 | 1487.55 | 1482.6767 | 1478.1472 | 1474.9444 | Grenoble | Marhauser | D275D610-9FD2-6174-F985F285973D01B3 | 5 peaks noted between TM110 and TM111 screen shots may be found in HOM measurements data file. | Yes | 1496.4407 | 29 | Yes | Overcoupled | CD7E31DB-D4B6-739A-194DCE5D8D684804 | 4.6 | 11500000000 | 9720000000 | 1110000000000 | 11.7 | 0.001 | 1170000000000 | 576000000000 | 14.67 | 8210000000 |  |  | 8.5 | 4010 | 12000000000 | FE | Cavity is FE loaded at 15 MV/m | Yes | attempted to process with limited results. |  |  |  |  | 8.87 | [12 empty] | cavity was FE loaded below 15MV/m | CE01F99E-EE99-34BF-A93FF7237716D641 | CE047320-F4E9-336B-E918F2E64E0B4928 | CE06130D-D73F-5F4D-D98810ADB93302EE | CE0A6A4F-9082-25EA-3CB2E15052C0155A | CE0AAD60-E3AC-95F7-3F66CE0CB28F2017 | CE0B02BE-B5F2-BED1-C0CDB59CB7E139EF | CE0B9B35-01CF-26BF-58DEBC41ABB2785A | 1 | MARHAUSE | 2011-02-10 09:45 | No | 29
32 | C100-RI-015 |  | Yes | No | 2 | C100_RI_015.txt | 2011-01-26 | 7 | TS28-008 | KDAVIS | KDAVIS | 9.2e-09 |  |  | 195 | 2.07 |  | Yes | Yes |  | 1496.4567 | 1495.3002 | 1492.1276 | 1487.5561 | 1482.5284 | 1477.9675 | 1474.8721 | KDavis | KDavis | CE1E79A7-9F42-6F6D-957019C99AF72E83 | no issues | Yes | 1496.4575 | 29 | Yes | Overcoupled | CE1FE11E-F8E6-B1D5-490881136F9682C8 | 5.6 | 16000000000 | 12400000000 | 1180000000000 | 11.6 | 0.001 | 14000000000000 | 23000000000000 | 27 | 11200000000 | 11800000000 | 2.19 | 18 | 16.9 | 14800000000 | Other | admin limited to 27 MV/m | Yes | quench processing from 24 MV/m to 27 MV/m, radiation finally decreased to background | -4.6 |  |  | 0.006 | 26.5 | 0.0289 | [7 empty] | 11800000000 | 12800000000 | 13500000000 | 15000000000 | quench processing from 24MV/m to 27MV/m cleaned up FE | CE297F9F-9797-DD44-A8063E88301437D5 | CE29F9DF-9A27-38C3-FDD0DA4E655A3B8D | CE2A9F85-BBF2-9073-47F9F729AD1394A2 | CE2BB473-FDC3-E304-A37B67DE53749774 | CE2C09FE-0BAE-FF78-E72B52DA47EE626F | CE2CF477-E904-6613-00A48611E88916CA | CE2D609B-A8E3-6DBB-8EC9B9824A622E40 | 1 | MARHAUSE | 2011-01-28 16:56 | Yes | 29
33
34
35 | C100-RI-001 |  | Yes | No | 2 | C100_RI_001a | 2011-02-01 | 8 | TS28-001 | STIRBET | KDAVIS | 1e-09 |  |  | 189.1 | 2.07 |  | Yes | Yes | Cable for incident RF power SN 011 from July 2009 | 1496.4419 | 1495.2463 | 1492.1197 | 1487.6679 | 1482.6345 | 1478.2126 | 1475.1409 | Mircea | KDavis |  | HOM survey on C100_RI_001 done 24th Sept 2010 M:\asd\asddata\VTA_RF_DATA\1497_DATA\HOM data files |  | 1496.4402 | 29 | Yes | Undercoupled | ED2FD04C-B2C4-6D72-6B4E866AD4B0FA0D | 4.86 | 10200000000 | 12100000000 | 541000000000 | 11.6 | 0.036 | 19000000000000 | 2370000000000 | 27 | 8280000000 | 8550000000 | 0.00105 | 27 | 0.0191 | 10500000000 | Operator | Admin limit 27 MV/m | No | C100_RI_001 had no RF processing. However, the cavity have been retested twice on 2nd Feb (see data in C100_RI_001b)to check for issues with Lorentz detuning. Data for Post Processing are from file C100_RI_001b. Initial test wa withour radiation. On subsequent RF tests, radiation onset at 18 MV/m and KLoren from C100_RI_001b. | -3.5517 |  |  | 0.201 | 18 | 26.2 | [7 empty] | 8710000000 | 9.39e+19 | 9790000000 | 10600000000 | Cavity C100_RI_001 have been tested twice: on 1st Feb 2011 , no radiation onset up to 27MV/m but with small Lorentz coefficient (and fanny shape). Retested again next day (C100_RI_001b) with radiation onset this time at 18 MV/m, Lorentz coefficient still small (- 3.5517), but with linear shape. Probably RF cable issues. Consider cavity qualified for string assembly. The HOM data has been collected on the first test. Attached are all files for the two RF tests. | ED72A890-E5BB-8A5B-7CDE49DF67ECFF24 | ED73BD9C-A0C9-E7E0-4F14B190144AC39B | ED7471A1-D220-38B7-9E7F6F425F7DD9B7 | ED77BA98-0929-3251-45275017BB31152B | ED787414-AD34-E730-4BE98DB179758359 | ED79816B-0788-E0E8-0ADDDA229893474D | ED7A107E-D01F-C130-351CCFB9569D23C6 | 1 | MARHAUSE | 2011-02-08 08:57 | Yes | 29
36 | C100-RI-002 |  | Yes | No | 2 | C100_RI_002.txt | 2011-02-03 15:00 | 4 | TS28-004 | KDAVIS | WESTBY | 8e-09 |  |  | 168 | 2.07 |  | Yes | Yes | Directional coupler under shield lid. Circulator and 500 Watt load under shield lid. HOM power meter under shield lid | 1.4964 | 1.4952 | 1.492 | 1.4876 | 1.4827 | 1.4782 | 1.4751 | KDavis | KDavis |  | Previously performed satisfactorily; see C100-CAV-VTRF R2 Sequence #4. | Yes | 1496.3874 | 29 | No | Overcoupled |  | 5.84 | 19700000000 | 9930000000 | 882000000000 | 12.5 | 1500 | 1050000000000 | 5670000000000 | 27 | 11900000000 | 13500000000 | 0.00464 | 0 | 0.0282 | 21000000000 | Other | Admin limited to 27 MV/m gradient. | No |  | -4.8 |  |  | 0.00464 | 0 | 0.0282 | [7 empty] | 12800000000 | 14700000000 | 15700000000 | 18500000000 | Admin limit (gradient) 27 MV/m. | F145C689-FBB8-44C2-B38F17DD4C9BFF55 | F14653F8-A0BE-1335-A34C66EAEA92DB09 | F146B391-A075-E237-A60936B3E1CE8743 | F1470A28-AA61-FFA0-9C11A9B6B89E87F0 | F1476370-B954-3CAC-E2324683412DD7B2 | F1486DFE-AB2F-6AD5-6DBB4247397EED4E | F14927B8-0A08-6BD0-9A9C75A2A0AC8772 | 1 | MARHAUSE | 2011-02-08 08:58 | Yes | 29
37 | C100-RI-008 | Previous test FE loaded, HPR before this test | Yes | No | 2 | C100_RI_008b.txt | 2011-02-05 18:42 | 8 | TS28-007 | KDAVIS | KDAVIS | 5.5e-09 |  |  | 187 | 2.07 |  | Yes | Yes | HOM PM under the shield lid. Pinc/Pref directional coupler, circulator and 500W load under the shield lid. | 1496.609 | 1495.406 | 1492.26 | 1487.713 | 1482.768 | 1478.163 | 1475.059 | KDavis | KDavis |  | Previously performed, see traveler seq. num. 11 | Yes | 1496.6066 | 29 | Yes | Undercoupled | 01A7F786-C0AA-211D-AC3C2F3DAA89B328 | 4.36 | 10600000000 | 12600000000 | 1090000000000 | 11.6 | 0.001 | 30000000000000 | 70000000000000 | 14.4 | 5200000000 |  | 147 | 10 | 333 | 10800000000 | Other | FE loaded at 14.4 MV/m post-processing | Yes | Performed RF processing, quench processing in 6/7, 4/7, 3/7, and 2/7 pi modes. | -6.1 |  |  |  | 7.5 | 1460 | 18 |  | 23.5 | 15.5 | 8.5 |  | Q-switching in pi mode at 21 MV/m. Overall results point to heating in cell 2 or 6. |  |  |  |  | Cavity performance deteriorated following RF processing. Final power rise was FE loaded at 14.4 MV/m | 01B68703-F2CD-DC5C-C577B985D30ECCE2 | 01B7090F-BBC6-9607-C9A3D97077D4DB3D | 01B7DBEC-BF95-DD90-DDEA1FA3D460421D | 01B80C54-D7C9-1F09-04E1BD5FC271FC36 | 01B88D6D-AE88-7BAC-E10EC02F07ABB3BA | 01B8E9F0-A76C-A311-28997AEA5EA05CB4 | 01BA8DFA-F31C-1D78-96A4A88AC793F9C7 | 1 | MARHAUSE | 2011-02-08 09:02 | Yes | 29
38 | C100-RI-013 |  | Yes | No | 2 | C100_RI_013A.txt | 2011-02-08 | 8 | TS28-009 | TBASS | STIRBET | 5e-09 |  |  | 187.04 | 2.07 |  | Yes | Yes | Serial number for FPC cable 024 09/10 25693 | 1496.4388 | 1495.2424 | 1492.0831 | 1487.547 | 1482.6733 | 1478.1435 | 1474.9697 | Grenoble | Grenoble | E4E2433D-F471-C4B2-9C968C12970BE240 | HOM Survey done on 27 or 28 Jan 2011-Christiana |  | 1496.4365 | 29 | Yes | Overcoupled | 0C552612-C5B0-FB7C-A661B4D736F9810B | 5.38 | 12800000000 | 8320000000 | 1250000000000 | 11.8 | 0.006 | 4280000000000 | 2450000000000 | 27.1 | 8420000000 | 9270000000 | 0.00338 | 23.5 | 30 | 13400000000 | Operator | Admin Limit @ 27MV/m | No | No RF Processing done, cavity performed very well | -5.2874 | [12 empty] | 9190000000 | 10100000000 | 10600000000 | 11600000000 | Radiation spike at 20MV/m, radiation onset at 23.5MV/m (30 mrem/h), no RF processing necessary | 0C6CEBB8-0D2C-6017-AD8B3FF756D89B8A | 0C735D08-9229-9192-D412A76F7AD9F1E4 | 0C73A2AF-9810-8612-E214F55B16E89BC3 | 0C73E25E-C190-DC85-16B6ADF019E80223 | 0C7417CD-D54A-97AF-947EDD3D7AD30AB1 | 0C745185-B572-D36B-0EFE7A665CB22F9B | 0C749CAD-FAA7-CDDC-B25A02337BC966A7 | 1 | HAIPENG | 2012-02-24 13:07 | Yes | 29
39 | C100-RI-017 | Unable to do high power RF test because cavity vacuum below lambda was too high. | Yes | No | 2 |  | 2011-02-11 15:06 | 8 | TS28-008 | GRENOBLE | WESTBY | 0.0005 | [8 empty] | 1496.3832 | 1495.2087 | 1492.1189 | 1487.5189 | 1482.4856 | 1477.9165 | 1474.8689 | Grenoble | Westby | 1654F6B9-B78B-CE7C-7F9A6BE7DEE98156 |  | Yes |  |  |  | Overcoupled | [16 empty] | RF power | No high power RF test were performed. | [27 empty] | 1 | HAIPENG | 2012-02-24 14:12 | Yes
40 | C100-RI-018 |  | Yes | No | 2 | C100_RI_018.txt | 2011-02-15 | 8 | TS28-001 | TBASS | STIRBET | 1e-09 |  |  | 193.4 | 2.07 |  | Yes | Yes | HOMb cable calibration indicates a problem with the cable inside the dewar: HOMb Cable Calibration correction factor = 99.3, Cable Return Loss Popen = 24.2 mW and PHOMb = 22.4 Similar data regarding HOMb observed during initial HOM survey | 1496.3829 | 1495.2195 | 1492.0831 | 1487.5463 | 1482.5847 | 1478.0595 | 1474.9855 | Mircea | Bass | 2EFB7256-97FF-5AEB-99676442E300E52D | Initial HOM survey done with 20000 points indicates an issue with HOMB cable |  | 1496.3837 | 29 | No | Undercoupled | 2EC3B33D-9A24-FEC2-A86B8D69C63D6D8D | 5.26 | 11800000000 | 14000000000 | 747000000000 | 11.6 | 0.01 | 4570000000000 | 6.27e+17 | 27.0956 | 8800000000 | 9300000000 | 0.0037766 | 50 | 0.013433 | 12100000000 | Operator | Admin Limited at 27 MV/m, without any radiation (during test, background radiation) | No |  | -5.2086 | [12 empty] | 9370000000 | 10200000000 | 10700000000 | 11700000000 | Cavity admin limited at 27 MV/m, no radiation during testing above background. Cable on HOMb with issues.During RF during and after RF power qualification of C100_RI_018, cavity vacuum was better than 10-9 mbar. 30 min after opening dewar lid, cavity vacuum went suddenly to 10-7 mbar. No power glitch. Turbo running. | 2ECD36F5-A3E2-0FF4-4FFD8CF1C74B0303 | 2ECDAC1F-E367-9882-4F09E1A5DD4FCA06 | 2ECDDFD5-E19D-3407-6AB163379D03F2D7 | 2ECE0BFF-B43B-2ABB-CA17D4F6658DA775 | 2ECE29D4-02BB-4874-2E0518C7481A327A | 2ECE485D-E195-79CB-68C7C2F9B541FC17 | 2ECE7D2C-EADB-FC63-C3EA04EE99409D64 | 1 | HAIPENG | 2012-02-24 14:18 | Yes | 29
41 | C100-RI-019 |  | Yes | No | 2 | C100_RI_019.txt | 2011-02-17 16:23 | 8 | TS28-004 | KDAVIS | GRENOBLE | 1e-09 |  |  | 175 | 2.07 |  | Yes | Yes |  | 1496.291 | 495.118 | 1492.012 | 1487.422 | 1482.481 | 1477.952 | 1474.85 | Grenoble | Westby |  |  |  | 1496.2914 | 29 | No | Overcoupled | 6D7BB3C1-FABD-2A78-EBF443E0EE60DD06 | 6.04 | 13400000000 | 11200000000 | 1070000000000 | 11.6 | 0.004 | 20000000000000 | 7000000000000 | 27 | 9180000000 | 9990000000 | 0.01 | 7.5 | 2.66 | 13900000000 | Other | Admin limited (gradient) at 27 MV/m | Yes | RF power processing beginning at 7.5 MV/m on initial power rise. No radiation >1e-2 measured during post-processing data run. | -3.8 |  |  | 0.00457 | 0 | 0.0183 | [7 empty] | 8930000000 | 9910000000 | 10400000000 | 11600000000 | Initial two data sets yielded Kl=-6.8. Recalibrated cables and final data set yeilded Kl=-3.8. | 6D90DA82-0453-A93C-D47A555700F2458E | 6D9118A7-AAF8-E528-5FD3C6E4FC9E8B77 | 6D91796C-DF72-AADE-684B31BFFA0110A6 | 6D91B879-A07B-FA75-961A30DA4C373F35 | 6D91ED66-053D-B944-463460539A4BE040 | 6D922EA6-B937-ACA4-F8553B4E75E4D38E | 6D92B644-CB20-A622-A42E8D15418D1AF3 | 1 | MARHAUSE | 2011-03-01 08:22 | Yes | 29
42 | C100-RI-017 |  | Yes | No | 2 | C100_RI_017.txt | 2011-02-24 | 4 | TS28-003 | TBASS | WESTBY | 1.32e-08 |  |  | 177.95 | 2.07 |  | Yes | No | FP Cable Return Loss measurement was Popen=24.2mW and Pfp=3.46mW (typical values Pfp=13mW). Initial cavity testing indicated problems with the FP cable, limiting the cavity test. The HOMb and the FP cables were switched under the dewar lid. Retesting of the cavity was successful using HOM-B signal for field measurements. Note: this explains 5e12 Q2 for final testing. | 1496.3813 | 1495.2078 | 1492.1178 | 1487.5175 | 1482.4831 | 1477.9141 | 1474.8663 | Grenoble | Westby |  | HOM Survey performed on 11-Feb-11, Traveler Seq Num 39 | Yes | 1496.379 | 29 | No | Overcoupled | 5E448164-9F9A-4985-4BA8D45900E924E7 | 5.71 | 13400000000 | 11000000000 | 4660000000000 | 11.6 | 0.001 | 14300000000000 | 875000000000 | 27 | 9700000000 | 10500000000 | 0.00618 | 22 | 8.76 | 14300000000 | Other | Admin Limited (gradient) at 27MV/m | Yes | Lost lock at 21 MV/m with radiation spike of 6E-1 mrem/h. Cavity repeatedly lost lock above 22MV/m with radiation up to 10mrem/h. Re-established lock and continued increase gradient up to 27MV/m. After reaching 27MV/m, performed a clean test from 2-27MV/m without radiation. Characteristics of processing suggest multipacting. | -5.4 |  |  | 0.00464 | 25 | 0.0197 | [7 empty] | 10700000000 | 11900000000 | 12500000000 | 14100000000 | Note that HOM-B was used to measure field. See previous comment in cable cal section. | 5E571C26-EFE1-D2BA-6EE54F9CEBCD6C9B | 5E57CB8D-AB16-4C2A-FC06422227054306 | 5E57F196-0665-3AAA-2ED2B71EE614F763 | 5E581BE7-B65C-42C8-D01793FBAEFAEDF9 | 5E583956-BD76-2300-7396BC00FA77FE63 | 5E58580D-EFE3-54E7-23AE1EAAC3114873 | 5E587D25-D90D-1080-94BDED104E26A893 | 1 | HAIPENG | 2011-02-25 14:37 | Yes | 29
43 | C100-RI-020 | Super fluid leak HOM measurement only |  |  | 2 |  | 2011-02-25 18:48 | 8 | TS28-008 | KDAVIS | GRENOBLE | 0.00023 |  |  | 186.5 | 29 | [11 empty] | Grenoble | Grenoble | 5F3EBE60-CF4C-52ED-65CF2982AC1AC5A8 | unable to upload file using specified file name name too long and has spaces |  |  |  |  | Overcoupled | [45 empty] | 1 | HAIPENG | 2012-02-24 14:20 | Yes | 2.07
44 | C100-RI-021 |  | Yes | No | 2 | C100_RI_021.txt | 2011-03-01 15:29 | 4 |  | KDAVIS | TBASS | 5.06e-09 |  |  | 173 | 2.07 |  |  |  |  | 1496.2902 | 1495.1495 | 1491.9962 | 1487.4167 | 1482.5065 | 1478.0017 | 1474.8622 | Grenoble | Westby | 7320AA33-F7FB-3652-86A69CD596C32977 |  |  | 1496.2905 | 29 | No | Overcoupled | 736BD90B-FBD5-065C-9E20E63419F7F154 | 6.04 | 12900000000 | 8960000000 | 992000000000 | 11.7 | 0.28 | 17200000000000 | 11500000000000 | 7 | 8980000000 |  |  | 5 | 9.79 |  | FE |  | No | [19 empty] | 736E9529-91E0-DA4E-9B1240F855BF22AE | 73740CAD-AD3A-2BA5-CEBAE74E9067B301 | [5 empty] | 1 | MARHAUSE | 2011-03-04 12:53 | Yes | 29.3
45 | C100-RI-022 |  | Yes | No | 2 |  |  | 8 |  | WESTBY | TBASS | [9 empty] | 1496.4822 | 1495.3067 | 1492.1839 | 1487.6013 | 1482.585 | 1478.0522 | 1475.0527 | Westby | Bass | 7D8AB165-A6C5-830D-8A350A911C79FDF0 |  | Yes |  |  |  | Overcoupled | [45 empty] | 1 | HAIPENG | 2012-02-24 14:29 | Yes
46 | C100-RI-020 |  | Yes | No | 2 | C100_RI_020.txt | 2011-03-06 11:18 | 7 | TS28-010 | KDAVIS | KDAVIS | 9e-09 |  |  | 207 | 2.07 |  | Yes | Yes |  | 1496.404 | 1495.25 | 1492.084 | 1487.518 | 1482.602 | 1478.068 | 1474.953 | KDavis | KDavis | 90B37CE3-0F10-C24B-886FED0F75450E19 | HOM survey previously completed. See traveler seq. no. 43 from 2/25/2011. |  | 1496.402 | 29 | Yes | Overcoupled | 90B4020E-C905-1E71-5DE7E589167A8956 | 5.18 | 16300000000 | 11600000000 | 857000000000 | 11.8 | 0.001 | 5000000000000 | 1300000000000000 | 27.2 | 10100000000 | 10800000000 | 0.00991 | 19.5 | 133 | 15300000000 | Other | Admin limited to 100mR/hr radiation at 27 MV/m | Yes | quenching began at 23 MV/m on initial power rise. RF/quench processed cavity to 27 MV/m. | -5.3 |  |  | 0.511 | 16 | 133 | [7 empty] | 10900000000 | 12100000000 | 12800000000 | 14100000000 | 133 mR/hr at 27.2 MV/m. Admin limited | 917F1631-B4F6-DE96-A3F9E6D9B8AE300B | 917F5651-B846-E84F-3AA20FD32E0BC464 | 917F978F-E19D-7D9C-2FA2DFB015CB8E76 | 917FCED3-CE63-D025-7730B2D36D425216 | 91800544-9927-7212-882539D1E2BAF454 | 918055BA-F1C3-6726-A7917BAF69BE197C | 91809D39-CD6E-E83D-FCAB7E7C8A4A18F9 | 1 | MARHAUSE | 2011-03-07 15:38 | Yes | 29
47 | C100-RI-021 |  | Yes | No | 2 | C100_RI_021A.txt | 2011-03-10 15:23 | 4 | TS28-004 | TBASS | WESTBY | 1.29e-08 |  |  | 175.8 | 2.07 |  | Yes | Yes |  | 1496.3077 | [6 empty] | Grenoble | Westby |  | HOM Survey performed on 1Mar2011, Traveler Seq Num 44, HOM Survey meets specifications |  | 1496.3052 | 29 | No | Overcoupled | A1889BFC-F8F9-5268-26AADC7B7AF789C5 | 6.24 | 16300000000 | 9300000000 | 993000000000 | 12 | 0.00152 | 21300000000000 | 12600000000000 | 28.6 | 8850000000 | 9930000000 | 1.59 | 16.97 | 347 | 15600000000 | FE |  | Yes | Attempted to clean cavity by processing in 6/7, 4/7 and 3/7 pi modes. | -4.8951 |  |  | 1.59 | 16.97 | 347 | [12 empty] | A187A440-FBE4-7786-884B917E6D30455B | A1863520-D172-8A0C-8A3A1093874CBB30 | A1852CE9-D78F-378B-2837CE21ECFCF2E2 | A1855521-0832-2042-BA076E463B51914B | A18576C9-DB84-E6B9-A5B9B72E946B46E3 | A1859AD1-BA14-A382-ACA535C403068D26 | A185C328-D464-275A-55F990521D35B323 | 1 | MARHAUSE | 2011-03-11 08:14 | Yes | 29
48 | C100-RI-022 | test stand leak suspected cold cavity isolation valve closed during cooldown and high power test | Yes | No | 2 | C100_RI_022.txt | 2011-03-10 10:10 | 7 | TS28-008 | GRENOBLE | KDAVIS | 1.5e-08 |  |  | 190 | 2.07 |  | Yes | Yes |  | 1496.4822 | 1495.3067 | 1492.1839 | 1487.6013 | 1482.585 | 1478.0522 | 1475.0527 | Bass | Westby |  |  | Yes | 1496.4874 | 29 | Yes | Overcoupled |  | 4.93 | 17000000000 | 7720000000 | 469000000000 | 12.8 | 0.003 | 4530000000000 | 1210000000000 | 28.01 | 11200000000 | 12400000000 | 0.00617 | 24.4 | 138 | 17600000000 | FE | very little radiation in pre RF processed data | Yes | quench processing at 23.7 to 25 MV/m radiation and quenching at 30.1MV/m radiation 138mR/hr | 4.48 |  |  | 0.336 | 21.93 | 92.8 | [7 empty] | 12600000000 | 13900000000 | 14800000000 | 16600000000 | 28MV/m radiation limited at 100mR/hr | A5B6D02A-BF6B-8860-E4292B5B61FD43A5 | A5B79945-F53F-FFDC-A49D909392D78077 | A5B802ED-E653-DA86-A669287B01F052CB | A5B8A9FF-F479-9839-F2757C8FB51EA646 | A5B90A3D-E699-9FE5-D89B1FDDA140E274 | A5B94FD4-BE5A-E3E1-8A7B7966C32E88DB | A5B99C3D-ACC9-D4B3-416552CCA8FA55D2 | 1 | MARHAUSE | 2011-03-11 11:25 | Yes | 29
49 | C100-RI-023 |  | Yes | No | 2 | C100_RI_023.txt | 2011-03-13 12:45 | 7 | TS28-010 | GRENOBLE | KDAVIS | 5e-09 |  |  | 155 | 2.07 |  | Yes | Yes | initial fwd pwr into detuned cavity had an error will repeat with high power amp repeat with high power amp ci 30726, cr 22603, ct 64.8 | 1496.4507 | 1495.3102 | 1492.1429 | 1487.5078 | 1482.4749 | 1477.8785 | 1474.7994 | Marhauser | KDavis | C47D718F-DF0D-4B42-113B7A9F0003AB4E |  | Yes | 1496.4478 | 29 | Yes | Overcoupled | B6027101-BABE-7DAA-B57C0FD380B0CCD6 | 4.92 | 15700000000 | 9330000000 | 782000000000 | 12 | 0.001 | 6000000000000 | 1000000000000000 | 27.01 | 9040000000 | 10500000000 | 6.33 | 14.6 | 350 | 15700000000 | Other | Administrative limit | Yes |  | -5.24 |  |  | 5.18 | 14.6 | 350 | [7 empty] | 10300000000 | 11700000000 | 12500000000 | 13800000000 |  | B14A7BF4-F09C-DEB1-A98969FC5D9F804C | B60DE01C-B919-F60E-BD190C6C696E2BD0 | B612BDAF-D669-FDD8-84AF361E3B1F2732 | B612F760-DC85-391D-B1EAB5D76AF028EC | B6132F9D-B531-7998-542356B615FE49F5 | B6145AF5-C8DE-3733-613A1CAE7C8251ED | B6148AEE-EF05-14AD-7B761221476DBF33 | 1 | MARHAUSE | 2011-03-14 10:54 |  | 29
50 | C100-RI-025 |  | Yes | No | 2 |  |  | 8 |  | STIRBET | KDAVIS | [16 empty] | Marhauser | KDavis | CA724CF6-C71B-0BBD-099CBE74C3F2A7A1 |  | Yes |  |  |  | Overcoupled | [45 empty] | 1 | MARHAUSE | 2011-03-18 15:31 | Yes
51 | C100-RI-025 |  | Yes | No | 2 | C100_RI_025A.txt | 2011-03-20 | 8 | TS28-010 | STIRBET | KDAVIS | 5e-09 |  |  | 183 | 2.07 |  | Yes | Yes | Test done with directional coupler under lid, HOM power meter under lid. | 1496.4016 | 1495.2603 | 1492.1124 | 1487.5174 | 1482.435 | 1477.861 | 1474.761 | Marhauser | Marhauser | DD8EF021-C5C4-112C-EA9AADBBBBC9BB6C | HOM survey done on 11th March 2011 |  | 1496.4212 | 29 | Yes | Overcoupled | DA2D449B-CC3D-2C08-6CE1EEE1EDEF083F | 5.73 | 12900000000 | 11300000000 | 1350000000000 | 11.6 | 0.001 | 5680000000000 | 53200000000000 | 30 | 7290000000 | 8130000000 | 0.0378 | 18.5 | 72.5 | 10600000000 | Operator | Radiation limitation 100 mR/h at 30 MV/m |  | Quench processing in the first run. In second run no more quenches up to 30 MV/m. | -5.1983 |  |  | 88.8 | 19.5 | 26.2 | [7 empty] | 8490000000 | 9170000000 | 9540000000 | 10300000000 | In first run, radiation onset with a spike at 18.5 MV/m. Starting with 21.5 MV/m frequent quenches, processed away. At 22.5MV/m Qo drop.Able to reach 30 MV/m with radiation about 100mR/h. The second "clean run" radiation onset at 19.5MV/m, no more quenches up to 30 MV/m. Significant lower Qo for the second run. | DA3B4A93-AB27-5E21-EB7937A8F0E28C1B | DA3BB934-BE23-A4A8-D06CAC60153A089C | DA3C0BC5-C5D1-ED9A-8E44FA93972D1B7D | DA3C4E5D-9F1C-844D-7A292554DA6F878A | DA3C9E12-D520-4B0B-71465AADD123930F | DA3CE5DB-BE78-3043-4E17599CF7732CFE | DA3D4EA4-F6A2-597E-CC325D7D64075397 | 1 | MARHAUSE | 2011-03-22 08:34 | Yes | 29
52 | C100-RI-028 |  | Yes | No | 2 | C100_RI_028.txt | 2011-03-26 19:00 | 8 | TS28-009 | KDAVIS | KDAVIS | 5.5e-09 |  |  | 183 | 2.07 |  | Yes | Yes |  | 1496.584 | 1495.44 | 1492.295 | 1487.682 | 1482.798 | 1478.163 | 1475.069 | Marhauser | Marhauser | F84E51C5-BE24-64E9-CFC8E70C859A72DB |  | Yes | 1496.5823 | 29 | Yes | Undercoupled | FCC335E3-A0F0-FEC4-1CFB1C3645FCAD1E | 4.28 | 10600000000 | 13600000000 | 1140000000000 | 11.6 | 0.009 | 30000000000000 | 10000000000000 | 27 | 7770000000 | 8630000000 | 0.000825 | 22 | 789 | 10500000000 | Other | Admin limited by radiation > 100mR at 27MV/m | No |  | -5.3 |  |  | 11.3 | 14.6 | 710 | [7 empty] | 6800000000 | 7330000000 | 7610000000 | 8520000000 | Input cable in dewar displayed strong dependency between input power and cable loss. | FCD25DC7-E200-18A1-E78C22AA489AA6EC | FCD2988A-94E0-2493-FA66F2A606A2695F | FCD2D849-F07D-C969-2D7F6316184E001E | FCD30C3F-CDB8-B909-59CD9127B58098B4 | FCD3462B-A835-F881-745CD26ECCA2B11A | FCD37C5B-E568-1F69-7517F25B1B794D21 | FCD3EBE5-D162-7E30-0216CCFE27A7FB49 | 1 | MARHAUSE | 2011-03-27 13:13 | Yes | 29
53 | C100-RI-029 |  | Yes |  | 2 | C100_RI_029.txt | 2011-03-31 | 4 | TS28-004 | KDAVIS | TBASS | 1e-09 |  |  | 170 | 2.07 |  | Yes | Yes |  | 1496.5407 | 1795.3888 | 1492.2387 | 1487.6557 | 1482.7975 | 1478.2652 | 1475.1927 | Bass | Marhauser | 0D08490F-CB45-3DDB-0BFA559CA56C6368 |  | Yes | 1496.5386 | 29 | Yes | Overcoupled | 73CB6C0F-C9A3-86D0-059DB01C3FE590FB | 6.05 | 13800000000 | 10200000000 | 1140000000000 | 11.7 | 0.003 | 16800000000000 | 9240000000000 | 26.7 | 10900000000 | 11500000000 | 0.00458 |  | 0.0153 | 14500000000 | Quench (non FE) |  | No |  | -3.76 | [5 empty] | 30 | 33.3 | 14.7 | 19.9 | 12.7 | 5.4 | Cell(s) 3/5 quenching at about 22.6 MV/m (corrected) |  |  |  |  | Cavity was tested twice on 3/31/2011. 1st test was quench limited at 21 MV/m, with lorentz coeff -6.1. Cable calibrations were repeated. 2nd test was quench limited at 27MV/m with lorentz coeff -3.8. On 4/2/2011, passband modes were studied to find quenching from cell 3/5. No temperature measurements were performed. | 73D30F35-A180-C135-DA9E1F28A858CCF8 | 73D397D1-F3C3-322D-7E60F16807DAF2FA | 73D3BE8A-0F15-989E-15207D0671F14F14 | 73D3D501-D100-F4F1-40226BF67FE15C56 | 73D3F289-E8C4-64E2-143E9D66A068E335 | 73D40DA1-F1F7-F223-0532B971D430F6E4 | 73DF451F-F76B-5477-A0B8DA7E3B9EFB3A | 1 | MARHAUSE | 2011-03-31 13:49 | Yes | 29
54 | C100-RI-032 |  | No | No | 2 | C100_RI_032.txt | 2011-04-05 | 7 | TS28-010 | TBASS | KDAVIS | 1e-09 |  |  | 208.6 | 2.07 |  | Yes | Yes |  | 1496.8782 | 1495.7354 | 1492.6028 | 1487.9882 | 1483.0395 | 1478.4232 | 1475.3612 | Bass | Marhauser | 2BF8EAB1-F20D-76A0-CF6309A6F16ABFAE |  | Yes | 1496.8754 | 29 | No | Undercoupled | 2BFC30AB-F6FD-FBF6-00C37536EAEF9D2D | [36 empty] | Could not perform high RF power test due to low cavity Q. | 2BFD8452-9BBB-051E-A9329FD053CB37B1 | [6 empty] | 1 | MARHAUSE | 2011-04-06 14:08 | Yes | 29
55 | C100-RI-030 |  | Yes | No | 2 | C100_RI_030.txt | 2011-04-09 | 8 | TS28-001 | STIRBET | STIRBET | 5e-09 |  |  | 192.98 | 2.07 |  | Yes | Yes | Ct=51.2 | 1496.5151 | 1495.3227 | 1492.1665 | 1487.5781 | 1482.8164 | 1478.3231 | 1475.2195 | Mircea | Marhauser | 44E0A63E-D675-8EF5-4C41564FAD493539 | HOM data with 20000 points collected on 9th Apr 2011 |  | 1496.5126 | 29 | Yes | Overcoupled | 44ADE031-0E70-9E78-44CE03A4A762FD3F | 5.37 | 12700000000 | 10700000000 | 741000000000 | 11.6 | 0.0015 | 8560000000000 | 21100000000000 | 27.8 | 10900000000 | 11100000000 | 0.00627 | 50 | 0.0214 | 13200000000 | Quench (non FE) |  |  | Cavity quench limited at 27.8 MV/m with practically no radiation onset. | -4.6361 | [12 empty] | 11100000000 | 12000000000 | 12400000000 | 13200000000 | Cavity performed very well in VTA tests, without radiation onset. One jump in HOMb data at about 25 MV/m. | 44B68AE2-A88D-721E-52355769EB11EF28 | 44B6D8AF-D98D-AB7A-2F4C3EF0F94C8F22 | 44B74F87-DCBA-12DF-172B58D8F9121AE7 | 44B7E6B2-021A-8A59-6EFCF68D25CC4024 | 44B88489-9679-98E2-EE7EC3FD0182EF1F | 44B8D8C0-9169-8CFB-14CDEC16796AC91D | 44B93DF1-E439-2A35-08D1588AA1E04FB8 | 1 | MARHAUSE | 2011-04-11 10:02 | Yes | 29
56 | C100-RI-031 |  | Yes | No | 2 | C100_RI_031.txt | 2011-04-14 | 8 | TS28-009 | TBASS | TBASS | 1e-09 |  |  | 183.8 | 2.07 |  | Yes | Yes |  | 1496.5829 | 1495.4403 | 1492.2317 | 1487.6398 | 1482.5916 | 1478.0554 | 1474.9661 | Bass | Marhauser | 593861F8-960C-9452-BBE5DD18048183FD | no issues |  | 1496.5801 | 29 | Yes | Overcoupled | 593ED4EA-A6DD-CBCA-E5B15788B197FA64 | 5.54 | 13200000000 | 10500000000 | 1150000000000 | 11.6 | 0.001 | 21900000000000 | 4720000000000 | 25.7 | 9240000000 | 9980000000 | 0.00165 | 22 | 100 | 13600000000 | Quench (non FE) |  | Yes | processed cavity in 6/7 and 4/7 pi modes. 6/7pi mode quenched at 27.6MV/m with 2.4mR/h. 4/7pi mode quenched at 32MV/m with radiation spike of 10mR/h, further testing in this mode was limited to 29.2MV/m. | -5.13 |  |  | 2.19 | 16 | 140 | [7 empty] | 9880000000 | 10800000000 | 11400000000 | 12500000000 | initial testing was quench limited at 24 MV/m. Processed cavity in 4/7 and 6/7 pi modes. Post processing testing was quench limited at 25.7 MV/m with 140mR/h. | 594DF35B-D334-7393-3B81B11038E82FA2 | 59553CDC-A25E-2452-D1018340553BC749 | 59556823-A15F-C330-59AD84A8D8F45D23 | 5955A00B-DC3B-68D2-9184E8B12EDB9876 | 5955BC6B-F0F7-FE46-7C88C22DA1BD2EE1 | 5955E2F3-9BA1-90EA-5C7D2817277B2EF9 | 5956281B-BE77-EFD3-D3B2BFB33AE28630 | 1 | MARHAUSE | 2011-04-15 14:24 | Yes | 29
57 | C100-RI-034 |  | Yes | No | 2 | C100_RI_034.txt | 2011-04-15 20:52 | 4 | TS28-004 | TBASS | TBASS | 1e-09 |  |  | 169.45 | 2.07 |  | Yes | Yes |  | 1496.5464 | 1495.3801 | 1492.2959 | 1487.6454 | 1482.7856 | 1478.1024 | 1475.053 | Bass | Bass | 5BCF0308-AC9F-259A-477F9939B8BD58B4 |  |  | 1496.5436 | 29 | Yes | Overcoupled | 5BCFBC27-D02A-F83F-CFF5AEC5542EA78A | 6.2 | 14900000000 | 9910000000 | 684000000000 | 11.8 | 0.00099 | 7560000000000 | 19400000000000 | 24.8 | 10000000000 | 10600000000 | 0.0239 | 17.5 | 1 | 13500000000 | Quench (non FE) |  |  |  | -4.78 | [5 empty] | 25.6 | 28.4 | 20.1 | 22.6 | 18.8 |  |  | 10600000000 | 11800000000 | 12600000000 | 14100000000 | Cavity was quench limited to 24.8 MV/m. Ran cavity through passbands to find quench limit in each mode. Was unable to find quench limit for 1/7pi mode. | 5BD69DA4-A651-1ED9-85B230A7C34AD758 | 5BD6D000-FD21-385F-930EC6FACA9CB5D0 | 5BD6FF42-E9FB-7A4D-A06CD40D8A8F77E1 | 5BD741C7-C710-9252-07472A27C6F809E5 | 5BD77598-FDBD-0CA6-7469DB6C2E856AAB | 5BD7A648-CC09-9735-BED86EE01FD978E1 | 5BD7D468-005B-0D18-1D0F6430441503A2 | 1 | HAIPENG | 2012-02-27 10:48 | Yes | 29
58 | C100-RI-032 | Cavity has dent due to being dropped. 1st time in dewar, cavity had cold leak. HOM FT broke after warmup, while cavity was in the dewar. Cavity HOM FT was replaced for 2nd VTA test, low Qo prevented RF testing. Cavity was HPR before this test. | No | No | 2 | C100_RI_032A.txt | 2011-04-20 17:02 | 8 | TS28-010 | TBASS | TBASS | 5e-09 |  |  | 187.4 | 2.07 |  | Yes | Yes | [8 empty] | Bass | Marhauser | 7A00F27C-0958-4709-B5D6B4337EBECB18 | HOM Survey completed 5 April 2011, VTRF traveler #54 Fitted HOMs, no issues, FM |  | 1496.875 | 29 | No | Overcoupled | 74BC5DFA-C935-1B07-DA400DE2195BA43A | 5.14 | 14900000000 | 9670000000 | 1590000000000 | 11.8 | 0.0018 | 83600000000000 | 20000000000000 | 36 | 8650000000 | 11100000000 | 0.0061 | 25 | 46.2 | 15300000000 | Quench (non FE) |  | No |  | -5.26 | [12 empty] | 11000000000 | 12100000000 | 12700000000 |  | Cavity was quench limited at 36MV/m, radiation ~50mR/h. | 74C05C80-E2F3-6B48-913623C29D9EFE77 | 74C0D538-B000-FE0F-EE715FEB724DBBBB | 74C0F0C2-08EE-24FC-41780F245F5C79CC | 74C10D80-F427-B063-C79D1873D4622A27 | 74C12C29-C187-4E6E-441523BBF9F7D1C2 | 74C1458C-C0A0-6B92-64719A20809FFA9B | 74C16E98-E4B0-FD9D-A1AA49B5A5C733B0 | 1 | MARHAUSE | 2011-04-21 17:40 | Yes | 29
59 | C100-RI-027 | Cavity tested prior to 120C bake to troubleshoot cold leak problems | Yes | No | 2 | C100_RI_027.txt | 2011-04-02 | 8 | TS28-009 | KDAVIS | KDAVIS | 6e-08 |  |  | 160 | 2.07 |  |  |  |  | 1496.586 | 1495.428 | 1492.278 | 1487.654 | 1482.706 | 1478.176 | 1475.056 | KDavis | KDavis |  | No HOM survey performed on this cooldown; HOM survey to be completed on subsequent VTA test |  | 1496.5835 | 29 | Yes | Overcoupled | A15E9745-C484-E514-56BF038A3A8BE735 | 4.06 | 11100000000 | 6210000000 | 1150000000000 | 12.1 | 0.008 | 2000000000000 | 3000000000000 | 27 | 8500000000 | 9850000000 | 54.4 | 13 | 3430 | 10700000000 | FE |  |  |  | -5.09 | [16 empty] | Cavity Admin (Rad) Limited @ 27 MV/m (3R/h); no temperature measurements were made | A16D2095-F9F7-9EF6-97672C8516489603 | A16D68D9-DB47-88DC-42990821126A87B2 | A1702AB9-A7E4-4BAA-644DEC50FA2BE072 | A1704B87-0CFC-4E11-205BAF140706249B | A1706367-D7A3-68DC-C1459CA8DDAB5649 | A1707B4F-AF1F-7B3A-37375AB37096F5DE | 1B9910A5-0513-0FB8-B3B25C2D82C6A3AF | 1 | HAIPENG | 2011-04-29 15:15 | Yes | 29
60 | C100-RI-033 |  | Yes | No | 2 | C100_RI_033.txt | 2011-05-03 | 7 | TS28-010 | TBASS | KDAVIS | 5e-09 |  |  | 211.8 | 2.07 |  | Yes | Yes |  | 1496.5821 | 1495.4389 | 1492.2604 | 1487.6465 | 1482.7735 | 1478.1877 | 1475.0802 | Bass | KDavis | BB1C8427-AD69-8BE6-B934440B32B4069F |  |  | 1496.5828 | 29 | Yes | Overcoupled | BB1CFE19-0EA6-9024-6F0E9EEB22B73B08 | 4.43 | 13500000000 | 8130000000 | 1270000000000 | 12 | 0.0027 | 132000000000000 | 185000000000000 | 27.3 | 12000000000 | 11800000000 | 0.302 | 16 | 54.6 | 14100000000 | Quench (non FE) |  | No |  | -4.18 | [16 empty] | Cavity Quench limited at 27.3MV/m, ~50mR/hr. | BB200BA0-A095-F40F-6B0DBB71D05F15C4 | BB2072A8-08D1-FCB6-C6FE709AD0D652FB | BB209378-E5B0-6687-00A9E295EC7360CA | BB20AC08-F6CE-AEF1-69F66921B7627198 | BB20C32F-A573-2ED4-84A7BD05E1DDB2F2 | BB20DCA7-BB14-A751-3F63A60B8FB47466 | 1B87B356-AA0D-705D-92B9CB22FD275ADE | 1 | HAIPENG | 2011-05-04 12:13 | Yes | 29
61 | C100-RI-035 |  | Yes | No | 2 | C100_RI_035.txt | 2011-05-04 15:11 | 4 | TS28-004 | TBASS | TBASS | 5e-09 |  |  | 176.7 | 2.07 |  | Yes | Yes |  | 1496.5749 | 1495.4307 | 1492.1986 | 1487.5814 | 1482.7932 | 1478.2371 | 1475.0325 | Bass | Bass | BC6EE27D-B5DC-A59E-5D676ED000F75838 |  |  | 1496.572 | 29 | Yes | Overcoupled | BC6F6A65-901A-D294-9D0F97C2970288A3 | 5.62 | 15700000000 | 9760000000 | 1010000000000 | 11.9 | 0.004 | 74500000000000 | 5170000000000 | 32 | 8140000000 | 10300000000 | 0.224 | 18 | 106 | 15600000000 | FE |  | No |  | -4.1257 | [16 empty] | Admin Limited (Radiation) @ 32MV/m, radiation 100mR/hr | BC756F55-9722-4A76-9CDBF64374A79600 | BC75FF26-0B08-24DC-18215C9A760C92FC | BC76202D-AB91-F0CC-8E0106319EE6695A | BC76396D-E778-72A3-0C20FF74DCFC9EDC | BC764F0D-F901-10DF-8C96C59644865534 | BC766B46-EBF2-0A71-50122B24E7DE451E | 1B79BEF9-E2DB-0D8B-B2598F267D603E80 | 1 | HAIPENG | 2012-02-27 10:49 | Yes | 29
62 | C100-RI-036 | HOM FT broke during 120C bake. Top beamline flange has SS radial wedge clamp | Yes | No | 2 | C100_RI_036.txt | 2011-05-18 15:06 | 8 |  | WESTBY | TBASS | 5e-09 |  |  | 193.2 | 2.07 |  | Yes | Yes |  | 1496.6715 | 1495.5283 | 1492.3457 | 1487.7687 | 1482.8935 | 1478.3853 | 1475.2654 | Westby | Bass | 0486ACD3-0469-FA22-28B7AF0541D6A640 |  |  | 1496.6712 | 29 | Yes | Overcoupled | 08A4E673-9CBD-B67F-888D43EBD8BCE8D4 | 5.68 | 12800000000 | 10400000000 | 761000000000 | 11.6 | 0.001 | 1550000000000 | 11800000000000 | 24.4 | 6820000000 | 8470000000 | 1550 | 10 | 19600 | 13400000000 | Quench (non FE) |  | No |  | -4.8657 | [5 empty] | 26.42 | 24.7 | 26.5 | 19.4 | [7 empty] | Cavity quench limited at 24.5MV/m, field emitter @ 14MV/m, subsequently very high radiation | 08A8CE81-BEA7-05EE-AC011BA6BF974405 | 08AC476E-0848-10EF-544DA56D2B74F766 | 08AC65CE-9A24-C7AF-700A923D382BE92A | 08AC840E-C257-88D9-9BC6E60CD25D6D65 | 08ACA9EE-E5D5-55C5-D7DE182491F447DF | 08ACDF5E-C93C-50AE-9883347A1EDDAFEA | 08A91779-0643-FC3B-1A5DC828E70A2EA6 | 1 | HAIPENG | 2011-05-18 17:11 | No | 29
63 | C100-RI-027 |  | Yes | No | 2 | C100_RI_027A.txt | 2011-05-20 16:01 | 4 | TS28-004 | TBASS | WESTBY | 5e-09 |  |  | 178.23 | 2.07 |  | Yes | Yes |  | 1496.5934 | 1495.4372 | 1492.2867 | 1487.6617 | 1482.714 | 1478.1836 | 1475.0614 | Bass | Westby | 0F031A08-EE4E-ED3F-E889AE9FD73122B2 |  |  | 1496.5907 | 29 | Yes | Overcoupled |  | 5.22 | 15800000000 | 8700000000 | 1020000000000 | 12.1 | 0.002 | 25100000000000 | 40000000000000 | 28.4 | 5820000000 | 10300000000 | 3.93 | 17.5 | 584 | 16400000000 | FE |  | No |  | -4.8736 | [12 empty] | 10300000000 | 11500000000 | 13400000000 | 14700000000 | Cavity admin limited (radiation) at 28.4MV/m (929mR/h). | 1D311BB4-F39C-C3F2-91954E8CB6AF7B23 | 1D313779-A7BD-CE48-FFDE8F4CC200C964 | 0F07D869-E752-6723-67BE2E412CBA92FA | 0F07F191-DFA8-7DB3-6DEB1A84B274ED2E | 0F080B22-C5D7-6F61-02A3BDBC7949D0C6 | 0F082689-9088-7BB2-A6A7614F016CE956 | 0F085232-9D3D-7655-ECBB8572B8A7A024 | 1 | HAIPENG | 2011-05-23 10:37 | Yes | 29
64 | C100-RI-037 |  | Yes | No | 2 | C100_RI_037_052311_145837.txt | 2011-05-23 | 8 | TS28-009 | WESTBY | KDAVIS | 4.6e-07 |  |  | 193.53 | 2.07 |  | Yes | Yes |  | 1496.5884 | 1495.3818 | 1492.1618 | 1487.5425 | 1482.3618 | 1477.655 | 1474.4291 | Westby | KDavis | 23126AE8-9C9E-ED2C-FF927846BD61A8F1 |  |  | 1496.5858 | 4.15e-07 | Yes | Overcoupled | 23158B39-A302-0E76-DC4DD9262B7FA34B | 4.72 | 14100000000 | 6980000000 | 1310000000000 | 12.4 | 0.0034 | 4950000000000000 | 4790000000000 | 27.2 | 10300000000 | 10900000000 | 61.1 | 12.5 | 7570 | 14900000000 | FE | Admin Limited (radiation)to 27MV/m | No | Data collected on 2nd run | -5.0368 | [12 empty] | 8690000000 | 9600000000 | 10400000000 | 11800000000 | Cavity admin limited (radiation) 27 MV/m. Additional Note: Cavity vacuum rose from E-9 to E-7 mbar range during pumpdown, RF tested despite higher than usual cavity pressure (see attachment C100-RI-037-VTA-COOL Processed Data.xlsx")" | 2328014A-AD19-B332-9082EE3E428AE3DC | 23345E2F-0A30-A79C-736D28A03DB5306C | 232942B2-E61E-13C6-76E0CE5456775ACE | 232A8BA8-E08E-1F9B-1930753209BB2B3A | 232B7C7A-C89C-36E7-F3612AB5B0C77358 | 232D114A-CE35-D452-0B479FB1877E777A | 23301A62-FAFD-C8C6-FE0D3D651CEC7289 | 1 | HAIPENG | 2011-05-24 15:32 | Yes | 29
65 | C100-RI-016 |  | Yes | No | 2 | C100_RI-016.txt | 2011-05-24 | 7 | TS28-001 | STIRBET | WESTBY | 5e-09 |  |  | 210.6 | 2.07 |  | Yes | Yes |  | 1496.4121 | 1495.2708 | 1492.114 | 1487.7003 | 1482.7686 | 1478.1911 | 1475.7832 | Westby | Westby | E7A70EFA-C6FC-25D3-29C409E16EADC614 |  |  | 1496.409 | 29 | Yes | Overcoupled | 279D3CEA-A509-03CB-410D554614FEF6DF | 4.99 | 14000000000 | 9540000000 | 897000000000 | 11.8 | 0.0015 | 6600000000000 | 6560000000000 | 27 | 8900000000 | 10900000000 | 43.2 | 12.5 | 442 | 14400000000 | Operator | Radiation in excess of 400 mR/h | No | Kept cavity filed constant at 20MV/m form about 20 minutes. No change in radiation (15- 20 mR/h) | -3.8661 | [12 empty] | 10800000000 | 11700000000 | 12500000000 | 14000000000 | Radiation onset at quite low cavity field (12.5 MV/m), no quenches. | 27AEC653-0170-20A2-1828BB58CA1C7926 | 27AF1C48-A416-54CA-781C5F5DF5443A7D | 27AF6813-AA25-5608-98CD668AF0962ABB | 27AFB605-A41E-93FC-799465ED29F37D91 | 27B007FF-D01E-C238-FA5A0E797C3230B4 | 27B04938-07B8-4DA8-9BA823F7973F5D2F | 27B09425-0ACF-9523-321AB2FE8B307FC4 | 1 | HAIPENG | 2012-02-24 14:10 | No | 29
66 | C100-RI-008 |  | No | No | 0 | C100_RI-008C.txt | 2011-05-26 | 8 | TS28-007 | STIRBET | STIRBET | 5e-09 |  |  | 177.1 | 2.07 |  | Yes | Yes | PIN and PRFL via directional couplers assembled over circulator under Dewar lid, port 3 circulator with 500W load. | 1496.702 | 1495.5288 | 1492.3649 | 1487.75 | 1482.7363 | 1478.0553 | 1474.9356 | KDavis | KDavis |  |  |  | 1496.7016 | 29 | No | Overcoupled | 31AE8467-EAD1-3AB2-8B0CDC7D21D7A56B | 5.62 | 13500000000 | 10000000000 | 907000000000 | 11.7 | 0.00334 |  |  | 27 | 9560000000 | 10600000000 | 5.42 | 12.5 | 2080 | 14100000000 | Other | No quench events up to 27 MV/m. Cavity admin limited at this value due to high radiation: radiation onset at 12.5 - 13 MV/m, slowly increase to about 350 mR/m at 24 MV/m, then a sudden jump to higher radiation . At 27 MV/m radiation was 1.8 - 2 R/h. |  |  | -5.4606 | [12 empty] | 10600000000 | 11700000000 | 12300000000 | 13700000000 | IN the previous two VTA tests (21st Nov 2010 and 5th Feb 2011), C100_RI_008 was plagued by low cavity filed quenches, Q sitches at 18 - 21 MV/m and radiation onset at 13.5 MV/m (ro 7.5 MV/m after processing). Radiation went up to 2R/h. Scratches on several irises have been identified and removed. In this test, no more qeunces events, no more Q switches , just radiation onset at 12.5 MV/m and masx radiation 2 R/h at 27 MV/m. | 31B78967-F611-C88E-3E102723E80F05F5 | 31B7CDA0-E2CD-7969-B0F8440150300FA1 | 31B8208D-F4B7-66F9-0DA46AE248BD50EE | 31B86ABC-C399-5853-D4170D66D4932670 |  | 31B8B9E9-DF2E-887A-804D6EE2DE43614F | 31B906CF-9064-7A29-A169FCD53F8AB461 | 1 | HAIPENG | 2012-02-24 13:40 | Yes | 29
67 | C100-RI-024 |  | Yes | No | 2 | C100_RI_024.txt | 2011-06-02 11:21 | 8 | TS28-008 | GRENOBLE | TBASS | 6.9e-09 |  |  | 190.44 | 2.07 |  | Yes | Yes |  | 1496.8162 | 1495.657 | 1492.478 | 1487.9799 | 1482.9654 | 1478.6142 | 1475.668 | Westby | Bass | 6590817B-A1BE-1163-AE00130F42C6578D | After high resolution Polfitting, the TM111 passband in the first pair marginal passed QA at 2K. The bead-pull after HV welding and further cold QA HOM measurement need to be done. |  | 1496.8158 | 29 | No | Overcoupled | 659EB996-D40E-9E1B-2AFF58DB8E920D09 | 4.62 | 16200000000 | 9090000000 | 800000000000 | 12.2 | 0.003 | 66500000000000 | 9230000000000 | 27.1 | 6230000000 | 9790000000 | 57.7 | 12.07 | 3620 | 15400000000 | Other | Administrative limit at 27MV/m when FE greater than 100mR/hr. | No |  | 5.0778 | [12 empty] | 10400000000 | 9910000000 | 10500000000 | 11900000000 |  | 65A9D7BD-AFA2-0FEF-FC10E36B4A81EA69 | 65AA3E8C-BE1B-F4F1-17C73D844848529A | 65AA65CC-00E4-98E0-4908F0DC1BE39697 | 65AA93BC-F7C3-99B1-FA54F7314AA6AF53 | 65AABFFE-9F17-92DF-39C5444E62CA3107 | 65AAE04C-A4D4-1648-A1E53E6CA5EC0850 | 65AB1D65-9E8A-1AB0-5EA48909B5F7F310 | 1 | HAIPENG | 2012-02-13 16:58 | Yes | 29
68 | C100-RI-037 |  | Yes | No | 2 | C100_RI_037A.txt | 2011-06-08 | 4 | TS28-004 | STIRBET | TBASS | 2.4e-08 |  |  | 165.35 | 2.07 |  | Yes | Yes |  | 1496.5851 | 1495.3788 | 1492.1587 | 1487.5392 | 1482.3579 | 1477.6506 | 1474.425 | Westby | Bass | 749D4FC6-E637-4096-57C3ED2EA1B4FCA0 | HOM survey done in the previous VTA dest on 23rf May 2011 |  | 1496.5843 | 29 | Yes | Overcoupled | 749B1ACF-9544-A4BD-7E26F6DC98FEC868 | 5.75 | 17300000000 | 7660000000 | 686000000000 | 12.8 | 0.001 | 980000000000000 | 5230000000000 | 12.5 | 5200000000 |  |  | 7 | 1.15 |  | Operator | Radiation onset at 7 MV/m followed by Qo drop. Radiation in excess of 100 mR/m at 12.5 MV/m | No |  | -4.9411 | [16 empty] | Cavity RF tested on 23May2011 (VTRF traveler #64). Cavity was Admin Limited @ 27MV/m due to high radiation levels (7.57E+3 mR/h). Cavity did pass specifications. Cavity had HPR and is being tested again to check the HPR water quality, after clean room glycol issues.In this test, C100_RI_037 had radiation onset at 7 MV/m followed by Qo drop. At 12.5 MV/m radiatio nwas in excess of 100 mR/m. The VTA qualification test failed. | 74A80FD3-CB1A-CA46-A9AB9E09A67875F3 | 74A8563A-E2E3-5A2B-93E0A343BAE27230 | 74A8A514-A505-80E5-63EFBD3D86852DBE | 74A8F095-07FE-2655-2F9151011B3F6833 | 74A92E6E-D541-7EF2-276C470D14F16FD0 | 74A980EE-9E87-66DF-AD035788CA3D01D8 | 74A9BA23-94F0-D691-0FA3BA82CE66125A | 1 | HAIPENG | 2012-02-27 10:56 | No | 29
69 | C100-RI-029 | Cavity is being re-tested after a propylene glycol leak into the class 10 clean room. This test is to verify the integrity of our systems | Yes | No | 0 | C100_RI_29B | 2011-06-13 20:26 | 8 | TS28-009 | KDAVIS | KDAVIS | 3e-08 |  |  | 170 | 2.07 |  | Yes | Yes |  | 1496.54 | 1495.388 | 1492.237 | 1487.655 | 1482.5 | 1477.512 |  | KDavis | KDavis |  | No HOM measurements made as this is a re-test |  | 1496.54 | 29 | Yes | Undercoupled | 8141F60D-ACA5-15F7-02397C63BF55FFA9 | 4.44 | 11800000000 | 13200000000 | 1080000000000 | 11.6 | 0.001 |  |  | 20.56 | 9020000000 | 9020000000 | 264 | 13.4 | 382 | 11300000000 | Quench (non FE) |  | Yes | processed in pi mode as well as all passbands. No improvement in quench field noted. | -4.905 | [5 empty] | 27.5 | 36 | 18.5 | 14.5 | 11.5 | 1.5 | 1/7 pi mode was breaking down at all field levels, was unable to establish good lock to cavity without breakdown. Once scaled cells 2/6 and 3/5 have approximately the same maximum cell gradient of 24 MV/m | 9020000000 |  |  |  |  | 8E8F5A5C-B6A0-E578-3F17ADF328681295 | 8E8EDD9C-D7C4-0884-E72A7AA063E2DC7D | 8E8C1668-DC2C-CC24-3FDA3230B678A282 | 8E8C6608-AF59-03BB-FDA20CEEF0C48B73 |  | 8E390BFD-029C-0385-D39EDAA03E87E7C5 | 8E8DCBAB-E832-9237-6C225417A69E6598 | 1 | HAIPENG | 2011-06-14 11:08 | Yes | 29
70 | C100-RI-038 |  | Yes | No | 2 | C100_RI_038 | 2011-06-09 15:56 | 8 | TS28-001 | GRENOBLE | TBASS | 5.06e-09 |  |  | 190 | 2.07 |  | Yes | Yes |  | 1496.6279 | 1495.4028 | 1492.0953 | 1487.3185 | 1482.0099 | 1477.2954 | 1474.1273 | Bass | Westby | 8A975126-BAF1-58B9-01D03FBFEBE3244B |  |  | 1496.6276 | 29 | Yes | Overcoupled | 8A992F0D-B598-590E-342FE5AE7C181FAD | 4.82 | 15200000000 | 12100000000 | 1300000000000 | 11.7 | 0.0001 | 6570000000000 | 6430000000000 | 33.8 | 10300000000 | 12300000000 | 66.9 | 2.5 | 3270 | 15300000000 | Cable | Field emission limit of 100mR/hr at 33.5MV/m. | Yes | Processed cavity cells 1 & 7 with no improvement. Heavy field emission during initial test affected Qo of cavity. RF processed improved performance. Processing in cells 3 and 5 or 4pi/7 mode reduced field emission thus improved the performance of the cavity. | -3.72 |  |  | 0.0561 | 17.95 | 148 | [7 empty] | 7240000000 | 8230000000 | 8620000000 | 9260000000 | After RF processing in cells 3 & 5 (4pi/7 mode), the cavity met all of the specifications. The input cable heating is still a problem. | 8AB85284-EA7C-2CC9-C4B703EC44DFA8E6 | 8AB8850D-04E0-CA5C-9A4087E44EFBFD08 | 8AB8B96E-940F-4C10-BD48949A08E1249C | 8AB8F0ED-F502-503F-94ED7D759EE86A4B | 8AB917AC-CA2B-F537-FCAFD9C86D1655C5 | 8AB96504-C8ED-9950-02A4E19D8ACFB264 | 8AB9B2DE-EB1F-9B7B-71780BF335F80757 | 1 | HAIPENG | 2011-06-14 09:59 | Yes | 29
71 | C100-RI-036 |  | Yes | No | 2 | C100_RI_36A.txt | 2011-06-10 16:38 | 7 | TS28-010 | GRENOBLE | GRENOBLE | 5.06e-09 |  |  | 203 | 2.07 |  | Yes | Yes |  | 1496.671 | 1495.5288 | 1492.3463 | 1487.7681 | 1482.8926 | 1478.3762 | 1475.2626 | Bass | Westby | 8ABE4E8C-F249-C525-04E00DBB2809258F | HOMs measured during initial test of cavity on 18May2011. | Yes | 1496.6705 | 29 | No | Overcoupled | 8AC017D4-D817-9A1F-E86452614C1388DC | 4.9 | 17000000000 | 11000000000 | 1040000000000 | 11.9 | 0.001 | 1790000000000 | 14100000000000 | 33.3 | 9640000000 | 12200000000 | 0.00775 | 22.5 | 64.3 | 17600000000 | Cable |  | No |  | -5.22 | [12 empty] | 10100000000 | 11700000000 | 12900000000 | 14400000000 | Cavity 2K lock frequency was above specification. Cavity initial Qo recorded during QfieldProbe measurement was a bit higher than specification. | 8ACECA0D-AD09-986B-85999DDAF636C5F2 | 8ACF23FC-DA70-0954-382B4F24DE883625 | 8ACF5147-9A7F-0159-4FB30CCDD7FB9571 | 8ACF75AA-C224-FD55-1AB8CF7C8BB94175 | 8AD18812-B478-D7F1-676EF6E21062B483 | 8AD22DD7-B3D0-9C2C-8EDDD0A04BF14688 | 8AD255CB-0828-FCB6-DE4AA9CA6EA154F0 | 1 | HAIPENG | 2011-06-14 09:47 | Yes | 29
72 | C100-RI-023 | cavity retest to verify that cleanroom water is not contaminated | Yes | No | 0 | C100_RI_023B.txt | 2011-06-13 | 7 | TS28-008 | TBASS | TBASS | 1.95e-08 |  |  | 198.53 | 2.07 |  | Yes | Yes |  | 1496.4575 | 1495.3169 | 1492.1515 | 1487.515 | 1482.482 | 1477.8904 | 1474.8053 | KDavis | KDavis |  | HOM survey performed on previous test (13-Mar-2011, VTRF #49). Cavity met HOM Specifications. |  | 1496.4571 | 29 | Yes | Overcoupled | 8E55258F-F48F-F17A-EC77C449D9C6900C | 4.84 | 17400000000 | 8640000000 | 523000000000 | 12.5 | 0.006 |  |  | 27 | 10000000000 | 12100000000 | 0.0431 | 19 | 147 | 16700000000 | FE | Admin Limited @ 27MV/m, radiation 154 mR/h | No |  | -4.87 | [16 empty] | Admin Limited @ 27MV/m, radiation 154 mR/h | 8E58CC00-0028-BD0E-A41BAF692BCAC12C | 8E599BAD-AA68-7026-A5676843EDADB6FC | 8E59C0D3-F2ED-BF96-95E361751C9922A3 | 8E59DD6D-FD25-637D-EB00FD90665B2766 |  | 8E59FAF5-91E7-6BAE-3C4CA2CC880AA007 | 8E5A3485-081E-8ECF-8D9FB4E3B5001DF0 | 1 | HAIPENG | 2012-02-24 14:31 | Yes | 29
73 | C100-RI-037 | Cavity was HPR after last VTRF test (Traveler #68) was FE loaded at 7MV/m. | Yes | No | 2 | C100_RI_037B.txt | 2011-06-23 17:18 | 8 | TS28-001 | GRENOBLE | TBASS | 9.1e-09 |  |  | 191 | 2.07 |  | Yes | Yes |  | 1496.5881 | 1495.3807 | 1492.1623 | 1487.5424 | 1482.3611 | 1477.6506 |  | KDavis | KDavis |  | HOM Survey completed on first VTRF test (traveler #64) |  | 1496.5876 | 29 | Yes | Overcoupled | BE6285A0-AE86-1522-2786AB2CB08BE972 | 4.47 | 17400000000 | 7480000000 | 968000000000 | 13 | 0.0018 | 950000000000000 | 5170000000000 | 25.9 | 7440000000 | 11400000000 | 292 | 9.5 | 6560 | 18200000000 | Other | Admin Limited (Radiation @ 10R/h) @ 26MV/m | Yes | test operator grenoble processed the cavity in 4pi/7, 2pi/7 and 6pi/7. FE was lower after processing in the 4pi/7 and 6pi/7 modes. The 2pi/7 mode processing was not effective in lowering FE. Repeated cable calibration and measured a new Qo vs. E curve. | -4.67 |  |  | 3.9 | 15.03 | 3160 | [11 empty] | after pi mode processing the field emission at 20MV/m was reduced from 75mR/hr to 3.9mR/hr. The post processed data files are named with PPdata.pdf due to file name character limitation. | D122336B-FBD5-41B2-2AA51F4DC6C3875D | D1228A81-E5C5-9FEB-264C1AF6A2382AF2 | D122B5BA-F0C2-791D-BA4A926B93952C4C | D122D89A-E314-A851-06CA623E375AE550 | D122FBEA-9D65-106E-A12D13DBCD0F4513 | D12328FD-97ED-FC37-A3717F8BBB975F91 | D1236D72-9DF7-8E39-6287D435AABAC35B | 1 | HAIPENG | 2011-06-27 09:56 | Yes | 29
74 | C100-RI-023 | Cavity retest after high pressure rinse control failure. | Yes | No | 2 | C100_RI_023C.txt | 2011-06-30 14:56 | 8 | TS28-008 | GRENOBLE | KDAVIS | 1.09e-08 |  |  | 190.68 | 2.07 |  | Yes | Yes |  | 1496.4698 | 1495.3507 | 1492.1429 | 1487.5281 | 1482.4958 | 1477.9005 | 1474.8214 | Westby | Bass |  | HOMs completed in March 2011 |  | 1496.4691 | 29 | Yes | Overcoupled | E1EC86C0-E837-6E8D-91E1FEEB0B29BC1F | 4.55 | 16600000000 | 10000000000 | 1120000000000 | 12 | 0.003 | 7040000000000 | 1120000000000000 | 30 | 8460000000 | 11200000000 | 0.024 | 19.05 | 12.9 | 17300000000 | Quench (non FE) |  | No |  | -5.05 | [16 empty] | cavity vacuum was low 1.09e-8 mBarr | E1F1EA75-9B1F-90C1-80E18C0EBB2F5004 | E1F294B5-D03E-3527-4F2B5D34D8EBFAF6 | E1F2D743-A855-BF2B-331254A00368FD24 | E1F36A86-0012-5BF6-76B8A2529E5E9A64 | E1F3A885-02FA-7CC2-41845A822E2035BB | E1F493D8-F080-A7B0-72BD41AF6ABC2428 | E1F52E2E-F878-DEC9-0FD866BD0BA31773 | 1 | HAIPENG | 2012-02-24 14:34 | Yes | 29
75 | C100-RI-040 |  | No | No | 0 | C100_RI_040.txt | 2011-07-07 | 8 |  | STIRBET | STIRBET | 6.3e-09 |  |  | 188.75 | 2.07 |  | Yes | Yes |  | 1496.8321 | 1495.5989 | 1492.2858 | 1487.5325 | 1482.3274 | 1477.5766 | 1474.1949 | KDavis | KDavis |  | CAvity without HOMs |  | 1496.8821 | 29 | No | Undercoupled | 065720A9-CBF1-F286-7ED80F22329FE095 | 4.92 | 11500000000 | 13400000000 | 1420000000000 | 11.5 | 0.003 |  |  | 27 | 6760000000 | 8070000000 | 0.0131 | 21.5 | 1.59 | 12100000000 | Operator | Qo drop starting at low cavity field. | No |  | -5.966 |  |  | 0.0131 |  | 1.59 | [7 empty] | 8200000000 | 8800000000 | 9200000000 | 9900000000 | Qo drop starting at low cavity field. Reach 27 MV/m with very low radiation. | 06603D17-DF86-6A2E-9529FFF27F3C65D2 | 066081BA-0189-47F8-297BF58B05D27658 | 0660CF4F-B427-EC97-91292DCD134C5E44 | 066116E0-BDEE-F084-32CEA7B9BB1745B9 |  | 066184E5-02D0-040C-D0CF0CB24E102A57 | 06622464-BA9A-F712-40EFFFAD468BC6CA | 1 | HAIPENG | 2012-02-27 11:18 | No | 29
76 | C100-RI-041 | Cavity was baked in furnace prior to VTRF test | No | No | 0 | C100_RI_041.txt | 2011-07-26 13:33 | 7 | TS28-001 | TBASS | TBASS | 5e-09 |  |  | 186 | 2.07 |  | Yes | Yes |  | 1496.8544 | 1495.6383 | 1492.3782 | 1487.6792 | 1482.5824 | 1477.7046 | 1474.5606 | KDavis | KDavis |  | No HOM survey performed on this cavity test |  | 1496.8509 | 29 | No | Overcoupled | 0B559279-C73F-51A8-6995C5ACD70471E5 | 3.98 | 16900000000 | 7580000000 | 1070000000000 | 12.9 | 0.003 |  |  | 29 | 7940000000 | 10100000000 | 0.067 | 19 | 44.9 | 17000000000 | Quench (non FE) | Quench limited @ 29MV/m | No |  | -6.1 | [16 empty] | quench limited at 29MV/m | 0B7083DD-FD93-5C1E-F3D0133DE19DED6A | 0B72CAA7-E246-B75D-F82DE687C0380F0F | 0B7529DD-A826-722A-8DC1BBC55EBEFFE1 | 0B76FA48-9166-A884-4424A436DE60602D |  | 0B7AF9F6-E599-3BA0-ED722BC4F74D1D61 | 0B817F07-018D-A4C6-EFA4FE399900E799 | 1 | HAIPENG | 2012-02-27 11:27 | Yes | 29
77 | C100-RI-039 | Cavity was baked in furnace prior to VTA RF test (see cavities C100-RI-040 and 041) | No | No | 2 | C100_RI_039.txt | 2011-07-27 15:13 | 8 | TS28-009 | TBASS | TBASS | 5e-09 |  |  | 193 | 2.07 |  | Yes | Yes |  | 1496.871 | 1495.6486 | 1492.4063 | 1487.655 | 1482.3513 | 1477.6786 | 1474.3251 | KDavis | KDavis |  | HOM survey not performed on this cavity test. |  | 1496.8673 | 29 | No | Undercoupled | 42639948-FF88-0C63-AE094F368DA9A7EC | 3.37 | 10000000000 | 12300000000 | 1080000000000 | 11.7 | 0.0007 | 592000000000000 | 5780000000000 | 26.2 | 6770000000 | 7280000000 | 1.37 | 15 | 391 | 9660000000 | RF power | Cavity was tested with roll-up 250W amp. | No |  | -5.88 | [16 empty] | Cavity was RF power limited at 26.2MV/m | 427B3A6D-9C0C-6377-8A061F61B01D0DE0 | 427E72F3-0DE4-44B4-B632F02C6ED21982 | 427F62F6-D567-2BE7-F1EBD73466E6FBB1 | 42808606-BF2C-9697-AED809BB64C74B38 | 428183A7-9B4C-895B-03C2846C61EEE4C6 | 42828123-A0F2-AF20-01E406D0CD1A9B03 | 4283B084-0F95-929C-18E39C16CDF883FD | 1 | HAIPENG | 2012-08-16 11:14 | No | 29
78 | C100-RI-041 | The cavity was moved from Dewar 7 to Dewar 8. Cavity retest in Dewar 8 to compare results with test in Dewar 7. | No | No | 0 | C100_RI_041A.txt | 2011-08-04 14:40 | 8 | TS28-001 | GRENOBLE | GRENOBLE | 5.06e-09 |  |  | 185 | 2.07 |  | Yes | Yes |  | 1496.8498 | [6 empty] | KDavis | KDavis |  | No HOMs performed on this cavity. | No | 1496.8496 | 29 | No | Overcoupled | A29ACE26-0D58-C474-0B5AA282D468EA2C | 4.65 | 13200000000 | 9510000000 | 1070000000000 | 11.7 | 0.007 |  |  | 29.5 | 9000000000 | 9950000000 | 0.0216 | 18.8 | 30 | 13700000000 | RF power | Unable to determine performance limitation with the 250W amplifier. | No |  | -5.17 | [17 empty] | A30F891B-BAA0-1F59-2CABDF18F5C200C9 | A311D5CF-A59C-3963-9C2988D737949B4E | A3144630-9D62-3B9D-722B0C4B6A40060F | A3168E42-FEFE-7A64-E766AAC8766292F8 |  | A319A481-9AF3-1BDF-BD9538A348759B8E | A31E2FAE-0360-C62E-7A12BFB8C7579FE2 | 1 | HAIPENG | 2012-02-27 11:24 | Yes | 29
79 | C100-RI-039 |  | No | No | 2 | C100_RI_39A | 2011-08-09 16:35 | 8 | TS28-009 | WESTBY | KDAVIS | 5e-09 |  |  | 162.67 | 2.07 |  | Yes | Yes |  | 1496.8701 | 1495.6478 | 1492.4055 | 1487.6541 | 1482.3504 | 1477.6775 | 1474.3241 | Westby | KDavis |  | No HOM's Performed |  | 1496.8668 | 29 | No | Undercoupled | 16A4F51F-D4E8-68A0-227FEE49407F0762 | 3.34 | 10400000000 | 13400000000 | 1400000000000 | 11.7 | 0.003 | 829000000000000 | 7040000000000 | 26.9 | 7270000000 | 8130000000 | 6.49 | 13 | 599 | 10200000000 | Other | Admin Limit | No |  | -5.9231 | [17 empty] | 16DE3D9D-F66E-0084-6D55E8CFC70F2F9C | 16E11677-CD6B-F425-55ABDAB6B991F847 | 16E54C3D-F070-19AE-458921D6DB86520F | 16E6DFF9-FA70-E39E-4A95AE51C9E040CF | 16E8BFF5-BB3D-9088-52B816973E9E59D1 | 16EAF6DF-E627-1478-F6DB8D0392FE29C5 | 16EF2F22-9706-2EE9-889B2AE829D3CF86 | 1 | HAIPENG | 2012-02-27 11:10 | Yes | 29
80 | C100-RI-043 |  | Yes | No | 2 | C100-RI-043.txt | 2011-08-16 14:37 | 8 | TS28-010 | STIRBET | GRIGORY | 5e-09 |  |  | 212.61 | 2.07 |  | Yes | Yes |  | 1496.5899 | 1495.3575 | 1492.0399 | 1487.4012 | 1482.1337 | 1477.6107 | 1474.2732 | Mircea | Bass | A201BA49-F3E3-51C0-65EC517DA92CEBC0 | HOM survey done using updated Polyfit program. Initial data are saved in HOM Survey Excel file, processed data in HOM processed Excel file. |  | 1496.5892 | 29 | Yes | Undercoupled | 7A854A2E-077B-EE7C-FBF234240659D768 | 4.71 | 12400000000 | 12800000000 | 1820000000000 | 11.8 | 0.000974 | 1980000000000 | 1030000000000 | 29 | 7960000000 | 9510000000 | 36.8 | 13 | 1360 | 12600000000 | RF power |  | No |  | -5.279 | [12 empty] | 9470000000 | 10300000000 | 10800000000 | 11800000000 | RF Power limited at 29MV/m | 7ACA6073-02AC-3AC1-6B2966EA0C54D150 | 7ACC9C4D-0CBC-2101-8B832048276514C7 | 7ACE84B6-9CB5-E00D-1C0046A00DA13638 | 7AD0B8E4-C784-6A7C-CD285CC1BEC76793 | 7AD26266-F661-C6DD-0DC391E22F00304C | 7AD4B380-0654-A9D2-7858B54A1AB4E422 | 7AD768A1-CAF8-9E54-F78F96A32C77CDF6 | 1 | HAIPENG | 2011-09-20 10:24 | Yes | 29
81 | C100-RI-044 |  | Yes | No | 2 | C100-RI-044.txt | 2011-08-18 16:15 | 4 | TS28-004 | GRIGORY | STIRBET | 5e-09 |  |  | 178 | 2.07 |  | Yes | Yes |  | 1496.721 | 1495.495 | 1492.2721 | 1487.4933 | 1482.3178 | 1477.7522 | 1474.4161 | Bass | Mircea | B181ABA4-EA5D-761F-E77955EA281D136B | Processed file was done with upgraded polyfit. |  | 1496.72 | 29 | No | Overcoupled | B1A2F3D0-F962-04E2-13B1B9E38F6283CD | 4.72 | 15200000000 | 10300000000 | 1230000000000 | 12.1 | 0.00287 | 863000000000 | 245000000000 | 28 | 9220000000 | 10400000000 | 2.41 | 15 | 118 | 15400000000 | RF power | The cavity is limited by the available power at 28 MV/m. The resonance frequency is 20 kHz off. HOM coupling is in 10e11 range (low?). | No |  | -5.2256 | [12 empty] | 10400000000 | 11700000000 | 12500000000 | 14300000000 |  | B1D52AA2-FBBC-8994-015EC4EA0DCA81B4 | 7CEBFE1A-DBC5-CE38-8DF6159FD36172E6 | B1E07096-C95F-DD0A-3D279FDC70058209 | B1E1D923-F699-E5FF-8E59DC2DFF8A244E | B1E3462F-AA92-ECE2-E94520F49B1CAD62 | B1E4F27D-EF99-4945-2EAAF447E022A83B | B1D8B2CA-B6AE-3013-210300B670FD0D0E | 1 | HAIPENG | 2011-09-20 10:35 | No | 29
82 | C100-RI-042 |  | Yes | No | 2 | C100-RI-042.txt | 2011-08-19 15:05 | 8 | TS28-001 | GRIGORY | STIRBET | 5e-09 |  |  | 184 | 2.07 |  | Yes | Yes | Peter took the network analyzer! | 1496.6625 | 1495.4297 | 1492.1546 | 1487.5596 | 1482.431 | 1477.7209 | 1474.3983 | Mircea | Eremeev | E2754B6D-BC66-81E2-B1267C267055113A | HOM survey done using upgraded version of polyfit | Yes | 1496.6625 | 29 | Yes | Critically coupled | E2864291-C8F9-A796-427136DDE13990A4 | 4.79 | 12300000000 | 10900000000 | 10700000000 | 11.7 | 0.00365 | 805000000000 | 65200000000000 | 17 | 4520000000 |  |  | 14 | 418 | 12500000000 | FE | It was observed that there was smoke inside the lab during this cavity assembly from wild fires in NC. The smoke also was felt in the cleanroom. NCR in the cleanroom has been generated. | Yes | Went through all the passbands and recorded the maximum accelerating gradient, quality factor (w/o recalibrating), and the radiation level. | -4.721 |  |  |  | 9.5 | 885 | 17.96 | 27 | 23 | 16.7 | 8.6 | 5 | Two quench events in 6pi/7 mode; one quench event in 4pi/7 mode. |  |  |  |  | Strong field emission, the maximum gradient - 17 MV/m. | E2E22003-94B9-793D-94A63D1A8B92EB7C | E2E53C5B-0868-A240-4E1511DF3BA1DC86 | E2F328F1-E905-4DAA-BFDC5F335ED3DBE5 | E2F65B72-A632-4DAA-C187D1E99CC0DE48 | E2F94410-BAAB-8973-C1FCEF99D1DCE522 | E2FC082A-F864-EEB4-C561E7A664901417 | E2FEC88B-C475-512A-71B0429F76E51E40 | 1 | HAIPENG | 2011-09-20 10:11 | No | 29
83 | C100-RI-040 |  | Yes | No | 2 | C100_RI_40A | 2011-08-22 17:42 | 7 | TS28-009 | TBASS | TBASS | 5e-09 |  |  | 221 | 2.07 |  | Yes | Yes |  | 1496.6382 | 1495.3934 | 1492.0754 | 1487.3248 | 1482.118 | 1477.3712 | 1473.9842 | Mircea | Eremeev | AB9E5722-90C8-2CFF-18F91619BD677048 |  |  | 1496.688 | 29 | Yes | Overcoupled | 829198F5-B52E-6C72-AA393C2840402A6B | 3.91 | 11700000000 | 630000000 | 805000000000 | 11.9 | 0.0085 | 1.92e-14 | 7.24e-13 | 16.5 | 9950000000 |  |  | 6.92 | 339 | 11700000000 | FE |  | No | Cavity was in the cleanroom during the smoke event and exhibits signs of field emission likely due to smoke inhalation. RF processing skipped. | -5.03 | [16 empty] | Cavity was FE Loaded at 16MV/m | 82AF8039-F11C-E8FA-8F47B7CEF2A292A5 | 82B0C538-0F9C-88C6-1CF25DB2EB0474B1 | 82CBC24F-E296-8EF1-E5F5F8EECFE22419 | 82CD2332-BE97-8B5D-2F57748285C788A6 | 82CE4C52-F6C6-C150-4D7B6614FB28D272 | 82CF6E03-F876-9B38-F2975794F174A657 | 82D0FCB3-E7C6-55E6-37B8DFCBE83550A3 | 1 | HAIPENG | 2011-09-20 10:04 | No | 29
84 | C100-RI-026 |  | Yes | No | 2 |  | 2011-08-23 12:58 | 8 | TS28-008 | TBASS | TBASS | 5e-09 |  |  | 188 | 2.07 |  |  |  |  | 1496.6319 | 1495.4598 | 1492.2512 | 1487.7591 | 1482.8044 | 1478.3663 | 1475.1876 | Bass | Bass | ABD45AA5-DBFD-A1E2-431AE8A221EE72CB |  |  | 1496.6311 | 29 | Yes | Overcoupled | B4EEA956-EA12-68E3-2C915919F9A87F6D | 5.1 | 13500000000 | 10600000000 | 1250000000000 | 11.6 | 0.00315 | 31000000000000 | 26700000000000 | 18.7 | 9920000000 |  |  | 16.43 | 183 | 13800000000 | Quench (non FE) |  | Yes | Cells 3 and 5 where limited at 22 MV/m, all other cells went higher in gradient per the spreadsheet. | -5.16 |  |  |  | 15.1 | 1.77 | 39 | 33.6 | 17.5 | 33.5 | 11.5 | [7 empty] | B528821D-AEE3-9229-F9193696B5340D13 | B529D2D6-F79F-A7CE-E6BEA2605C7A3C63 | B52B0C9C-F587-000D-B7C14B1A682584B4 | B52C1EEE-DF12-9D15-816A0786912FA439 | B52D8326-9684-2FE2-4C354FC311A5F357 | B52EE00F-EB21-D6B7-E3526AD42704839E | B5303B20-EA61-B0F0-7AB2531DA4850DDC | 1 | HAIPENG | 2011-08-25 13:48 | No | 29
85 | C100-RI-008 |  | Yes | No | 2 | C100_RI_008D.txt | 2011-08-25 09:41 | 7 | TS28-010 | TBASS | TBASS | 1.3e-08 |  |  | 197 | 2.07 |  | Yes | Yes |  | 1496.586 | 1495.354 | 1492.249 | 1487.603 | 1482.654 | 1477.934 | 1474.88 | Bass | Grenoble | 1377C325-9CB0-B5B2-8265236BA82867A1 |  |  | 1496.5821 | 29 | Yes | Overcoupled | 0C4FE401-B55F-4358-DBCD162D48EE21B8 | 4.66 | 17500000000 | 8840000000 | 604000000000 | 12.5 | 0.001 | 36800000000000 | 151000000000000 | 21 | 8250000000 | 9350000000 | 0.101 | 17 | 3990 | 17000000000 | FE |  | Yes | Cavity lost lock (FE?) at 32MV/m w/ Qo of 9.67E+09. Subsequent lock showed much lower Q value. | -4.9789 |  |  | 798 | 11.5 | 3990 | [11 empty] | Cavity was FE loaded at 21MV/m | 0E7968F7-B522-CD78-F4B17464D5C77038 | 0E7C2597-B7E4-4E02-6620F3C75B7A90EE | 0E7E8BA3-9F88-E88B-3D9FBBA5D5DA1D08 | 0E8088C9-A86F-6A8E-5BDB19D6C35D6389 | 0E82699E-96A8-EC35-B9DD037C4DF204EC | 0E841C49-A643-E8A3-9F921034C34249BC | 0E85ECF7-F7D3-0A91-4DE788843FE741A2 | 1 | HAIPENG | 2011-08-25 13:19 | Yes | 29
86 | C100-RI-042 |  | Yes | No | 2 | C100_RI_42.txt | 2011-08-25 | 4 | TS28-004 | KUSHNICK | KDAVIS | 1.3e-08 |  |  | 170 | 2.07 |  | Yes | Yes |  | 1496.666 | 1495.433 | 1492.158 | 1487.563 | 1482.434 | 1477.723 | 1474.4 | Mircea | Eremeev | 7C488042-C7D4-B2E4-05739D732DA076B0 | HOM Survey performed on previous HP RF test done on 19 Aug 2011 |  | 1496.6617 | 29 | Yes | Overcoupled | A8612599-A7CB-DB6D-12487C72833C5629 | 6.02 | 19700000000 | 8700000000 | 1650000000000 | 12.8 | 0.0014 | 1090000000000 | 64000000000000 | 34.4 | 10800000000 | 14100000000 | 1 | 20 | 100 |  | Quench (non FE) |  | No |  | -3.99 | [12 empty] | 13300000000 | 15800000000 | 17100000000 | 20200000000 | Quench limited at 34.4MV/m | A897E802-0B5D-5688-23A1E91CDDE647D4 | A8998CEA-E090-3AA4-E782014018BCBA82 | A89C61DF-B399-1DC2-9C2E36B444F42037 | A89DA1B4-BF85-4FE4-785A4ADEF0A6F1C8 | A89ECD91-CA91-8A8F-8E747ED200F26F8F | A8A0330C-96FE-F75B-0D192F7597B5BBD1 | A8A38197-A554-B45A-9B46F6C2394C6718 | 1 | HAIPENG | 2011-09-20 10:25 | Yes | 29
87 | C100-RI-032 |  | Yes | No | 2 | C100_RI_032B.txt | 2011-09-09 19:57 | 8 | TS28-009 | STIRBET | KDAVIS | 1.6e-06 |  |  | 176.99 | 2.07 |  | Yes | Yes |  | 1496.6792 | 1495.4909 | 1492.3786 | 1487.8099 | 1482.83 | 1478.2672 | 1475.2035 | Mircea | Mircea | 2661B83E-D475-A4C7-7A9635703E7F0481 |  |  | 1496.678 | 29 | Yes | Undercoupled | AFADB6C8-A1D7-DDC1-2742E59A0D5B4B7A | 5.04 | 11900000000 | 13200000000 | 1350000000000 | 11.5 | 0.001 | 81500000000000 | 35000000000000 | 16.47 |  | 5590000000 |  | 8.5 | 1290 | 12300000000 | RF power | radiation onset at 8.5 MV/m followed by Q drop. After a suddenrecovery, radiation went higher than 1 R/h associated with Q drop. Obs. Cavity vacuum was in 1 e-6 mbar at the begionning of this test. | No |  | -5.8751 | [17 empty] | ABBB99EC-DA69-B1D3-182B765CD501E045 | ABBFDFC9-C49D-50C8-0B81233F10B5CDBF | ABC311FF-EEAC-A1E3-439436C21516AA0B | ABC60242-BADE-A430-7C3F9CE6D537875C | ABC8AD8F-0B74-1AD1-62AEF0B81335D39F | ABCB29ED-FA47-1CE3-93F00C30070A0FB4 | ABCE698C-E4E6-1C72-1E6DBD0F2F1727B9 | 1 | HAIPENG | 2011-09-20 10:01 | No | 29
89 | C100-RI-045 |  | Yes | No | 2 |  | 2011-09-12 18:22 | 8 |  | KDAVIS | KDAVIS | [70 empty] | 1 | HAIPENG | 2012-02-27 11:41 | No
90 | C100-RI-046 |  | Yes | No | 2 | C100_RI_046.txt | 2011-09-14 | 8 | TS28-001 | GRENOBLE | GENG | 1e-09 |  |  | 169.83 | 2.07 |  | Yes | Yes |  | 1496.5689 | 1495.3256 | 1492.1027 | 1487.5161 | 1482.5583 | 1477.8539 |  | Geng | Bass | 53758106-F263-7C92-A135E239DF741633 |  |  | 1496.5677 | 29 | Yes | Undercoupled |  | 4.54 | 11000000000 | 13400000000 | 933000000000 | 11.6 | 0.006 | 143000000000000 | 15400000000000 | 26.9 | 6780000000 | 8140000000 | 14.8 | 13.47 | 1060 | 11700000000 | Operator | Radiation higher than 100 mR/h at 27 MV/m. Acctually it was about 1 R/h. | No | [14 empty] | 8290000000 | 7480000000 | 7810000000 | 8610000000 | Field emission was present at 20 MV/m, therefore Qo have been affected by fi9eld emitted electrons. | 53FE4270-E2EB-675E-BA10E223042602D2 | 5402129F-A981-A78A-E88B704D48CDB10F | 54083766-9118-10C5-59581073BEEC31A4 | 540AFC36-FC0D-5775-5CF749E9C3EE6178 | 540D5584-040C-159B-013B2B49406BB217 | 540FD72A-0E71-11F8-319237CAD0EC7F3E | 54143C3C-928A-4766-EA0743C9645EB7F9 | 1 | HAIPENG | 2012-02-13 17:25 | No | 29
91 | C100-RI-045 |  | Yes | No | 2 | C100_RI_045.txt | 2011-09-15 09:33 | 4 | TS28-004 | STIRBET | KDAVIS |  |  |  | 169.22 | 2.07 |  | Yes | Yes |  | 1496.6372 | 1495.4321 | 1492.1565 | 1487.4025 | 1482.2489 | 1477.6761 | 1474.5014 | Mircea | KDavis | 78B943D9-C921-C39D-2BFEAEE9901D5C0C | HOM trainny Tom goodman | Yes | 1496.636 | 29 | Yes | Overcoupled |  | 5.82 | 15700000000 | 11000000000 | 1380000000000 | 11.7 | 0.003 | 550000000000000 | 28000000000000 | 27.1 | 9130000000 | 10900000000 | 7.2 | 9.95 | 1170 | 16100000000 | FE | Admin. limit 27MV/m while Field emission > 100mR/hr | No |  | -4.4667 | [12 empty] | 10700000000 | 12100000000 | 12700000000 | 14500000000 |  | 793F6C7C-E7AD-FD2C-664903460930AAEB | 7945F5E7-FA2E-1AD6-C21CE1F78AD34A32 | 7950E26E-F74F-51E7-C060C602D653FC75 | 7958C174-A0E5-F6C0-B9CBB1D456027963 | 795C92B0-DF26-6B65-039F330A2E806571 | 7960D3B6-F817-D25C-E1698470289564BB | 796537D0-0040-7620-F36BF7C3FEE20D47 | 1 | HAIPENG | 2011-09-17 21:03 | Yes | 29
92 | C100-RI-041 |  | Yes | No | 2 | C100_RI_041B.txt | 2011-09-16 15:08 | 3 | TS28-007 | TBASS | TBASS | 8.6e-09 |  |  | 152 | 2.07 |  | Yes | Yes |  | 1496.6039 | 1495.3475 | 1492.0908 | 1487.4045 | 1482.3268 | 1477.4606 | 1474.3191 | Bass | Bass | EB89DDF7-9730-442E-1AABCDB767D6A0C8 |  |  | 1496.6011 | 29 | Yes | Undercoupled | EB9F8099-077F-6D2C-5305D790471B7F14 | 4.78 | 10500000000 | 11100000000 | 1570000000000 | 11.6 | 0.003 | 52300000000000 | 54200000000000 | 23.8 | 5910000000 | 7740000000 | 3.16 | 15 | 36.1 | 10800000000 | Quench (non FE) |  | No |  | -5.2714 | [16 empty] | Cavity began quenching at 20MV/m and was quench limited at 23.8MV/m. | EBD3C3F5-E431-E8BB-0415588AC9D16D63 | EBD6D6E2-CAF3-5115-E48DFFF8785350E3 | EBD88FC7-0F93-0FEC-B0076C841860D1BD | EBD9D833-A491-6269-43BD721116806F5F | EBDB97D1-95DD-5D57-FEA4B35AE023F9BD | EBDD5B48-B1D3-6F3D-05F7705716B22BAB | EC72A0E2-90C2-6811-A73BFD401F348DA8 | 1 | HAIPENG | 2011-09-18 17:06 | Yes | 29
93 | C100-RI-039 |  | Yes | No | 2 |  |  | 7 |  | KDAVIS | KDAVIS | [9 empty] | 1496.6242 | 1495.5378 | 1492.142 | 1487.3888 | 1482.0887 | 1477.4129 | 1474.0687 | Grenoble | Grenoble | 154C7D3E-C069-1967-D90E0AA7127C00CF | [5 empty] | Overcoupled | [45 empty] | 1 | HAIPENG | 2011-09-20 10:33 | Yes
94 | C100-RI-040 |  | Yes | No | 2 | C100_RI_040B.txt | 2011-09-16 | 7 | TS28-008 | KDAVIS | KDAVIS | 1e-09 |  |  | 213 | 2.07 |  | Yes | Yes |  | 1496.64 | 1495.396 | 1492.078 | 1487.327 | 1482.12 | 1477.373 |  | KDavis | Mircea | 466226BA-BFE0-1B3D-C8FD8CB085396E30 |  | Yes | 1496.638 | 29 | Yes | Overcoupled | 4676F42F-DCBE-1AFD-0A255077936B64C4 | 4.88 | 17100000000 | 6600000000 | 1300000000000 | 13.5 | 0.009 | 200000000000000 | 80000000000000 | 28 | 6850000000 | 10900000000 | 0.00147 | 23 | 0.82 | 18100000000 | Quench (non FE) |  | No |  | -5.5 |  |  | 0.00147 | 23 | 0.82 | [11 empty] | Q0 vs T not performed due to CTF limitation with concurrent CMTF operations | 46DB2AE3-BC56-B5C7-37FF932BD3ABA33A | 46DE4AFA-C7B5-85B1-D4C55494AF8AB6F4 | 46E1955D-E337-762E-B159BC795FE0E76A | 46E40E6C-AB20-A069-2DBF736927DC8A44 | 46E69527-D9C0-65E5-D54170FBC0DB9E42 | 46EC9C84-B6B7-4044-1E8D7F3E54BB6C5A | 1B3F8499-CA15-CB90-C975C4CCD42BD700 | 1 | HAIPENG | 2011-09-20 10:25 | Yes | 29
95 | C100-RI-039 |  | Yes | No | 2 | C100_RI_039B.txt | 2011-09-19 | 7 | TS28-010 | STIRBET | STIRBET | 1e-09 |  |  | 181.2 | 2.07 |  | Yes | Yes |  | 1496.623 | 1495.3769 | 1492.141 | 1487.3879 | 1482.0879 | 1477.412 | 1474.0677 | Mircea | Grenoble | 492B2BCD-E2A4-E5DC-EFCFB5A01BFD93DE |  |  | 1496.6222 | 29 | Yes | Undercoupled | 4AAC8491-9FE4-B7A1-C759C7E98BA8BDBA | 4.53 | 9980000000 | 11500000000 | 655000000000 | 11.6 | 0.001 | 25600000000000 | 17900000000000 | 27.1 | 9190000000 | 9420000000 | 0.161 | 11.54 | 291 | 10000000000 | Operator | Cavity had low Q, at limit of specification. Strange behavior in field probe channel - power jumps After initial low radiation, turn on even to 10R/hr per hour, and then decay over next 4 hours of testing to 100mR/hr at final run. Iritic behavior in field probe decayed over time, and was almost gone after final run. | Yes | RF process to check iritic behavior in field probe. RF process to reduce radiation after inital turn on above about 25MV/m. RF process to check max heat power into cavity. | -4.3761 |  |  | 291 | 14 | 291 | [7 empty] | 7980000000 | 9190000000 | 9370000000 | 9930000000 | Low Q but within spec, Beta slightly undercoupled, about .95 Admin limit to 27MV/m Some radiation precessing from 1R/Hr down to 300mR/Hr | 4A299834-9EBA-4574-D13A5589272B0D44 | 4A2D7C5D-AD49-82FC-5FDB821E5C7D0060 | 4A39A683-BA52-B9A2-09079D7B90A85310 | 4A43A9B5-F655-61E8-DE48C598BE617FD0 | 4A489BF5-B477-3083-FA499064B06635B3 | 4A4CBFC6-BFC7-3349-21BD88E6EEA91E6D | 4A4FEF0C-C780-AFF9-B80FCC259C521F3D | 1 | HAIPENG | 2012-02-27 11:07 | Yes | 29
96 | C100-RI-047 |  | Yes | Yes | 2 | C100_RI_047.txt | 2011-09-20 12:00 | 4 | TS28-003 | WESTBY | KDAVIS | 5.4e-09 |  |  | 167.83 | 2.07 |  | Yes | Yes |  | 1496.5224 | 1495.2783 | 1492.1108 | 1487.4688 | 1482.2721 | 1477.5396 | 1474.5062 | Mircea | Westby | 50E1B4CC-F658-BD94-7BC5CB0468A12768 | The mode at 2146.677MHz processed by Polfit with Ql=2.2e6 is not there in the recorded spectrum, but no harm to the spec. |  | 1496.5211 | 29 | Yes | Overcoupled | E466A5E8-AFC3-0D08-3E3A8B0633808847 | 5.79 | 15700000000 | 12100000000 | 1600000000000 | 11.6 | 0.001 | 122000000000000 | 13500000000000 | 24 | 6700000000 | 9180000000 | 54.4 | 12.03 | 717 | 15900000000 | Quench (non FE) |  | Yes |  | -4.6623 |  |  | 54.4 | 12.03 | 717 |  | 35 | 22 | 14 |  |  |  | 8920000000 | 9810000000 | 10300000000 | 11400000000 |  | E4A800B4-C80A-F857-18EFA604C442BECA | E4AA168D-C716-A0AE-8C55EB68067C99EB | E4AF6125-AD00-C3BA-36B732FAC3B98498 | E4B0E76A-97B1-22CD-32DF8F9AE326F961 | E4B2AB4B-081F-14CB-EEEE5A3D64E4D886 | E4B4F5C8-0B23-6308-2DC7A245746FEC59 | E4B9BB08-A4C4-1365-E753FDB838FF84C9 | 1 | HAIPENG | 2011-09-29 09:48 | Yes | 29
97 | C100-RI-041 | Re-test after early quench in Dewar 4 | Yes | No | 2 | C100-RI-041C.txt | 2011-09-21 07:33 | 8 | TS28-009 | KDAVIS | KDAVIS | 8.7e-08 |  |  | 159.4 | 2.04 |  | Yes | Yes |  | 1496.5929 | 1495.345 | 1492.087 | 1487.4 | 1482.322 | 1477.455 | 1474.313 | KDavis | KDavis | 7B68CE47-BA39-9112-28ED23B826FF2F2F | HOM survey done on 16 Sept 2011 |  | 1496.5973 | 29 | Yes | Undercoupled | A94E9158-F5CC-4111-4F441FEE4FFBEEF2 | 4.28 | 9040000000 | 10100000000 | 1670000000000 | 11.6 | 0.00181 | 65000000000000 | 41400000000000 | 27.3 | 6840000000 | 7390000000 | 0.0913 | 9.71 | 35.1 | 9640000000 | Quench (non FE) |  |  | Early Radiation onset FE blew up at 12MV/m Radiation dropped from 18mRem/hr to 1e-3mRem/hr Quench processing began at 21.5MV/m up through 27MV/m | -5.496 |  |  | 0.0913 | 9.71 | 11.5 | [12 empty] | A9E64223-097F-3FC6-BBB58F490751F2D8 | A9E8DDF8-C1BF-03EB-35E5C9401F940322 | A9EB4A7F-0BB9-ACA2-7FF186EE83E4AB4E | A9ECF826-0ACC-FF6A-5DD568115A5B6690 | A9EF01FA-DDC0-ABD8-13B247AD8B147D75 | A9F142C5-A26D-D993-B1C37AD79195C539 | A9F4CE45-E4EE-393C-26104D5211D5B6DC | 1 | HAIPENG | 2011-09-29 10:02 | No | 29
98 | C100-RI-032 | Cavity was originally damaged from fall (dented). Represents the follow up testing post-repair | Yes | No | 2 | C100_RI_32C.txt | 2011-09-17 12:00 | 8 | TS28-009 | WESTBY | KDAVIS | 3e-09 |  |  | 183.96 | 2.07 |  | Yes | Yes |  | 1496.6813 | 1495.4929 | 1492.3809 | 1487.8119 | 1482.832 | 1478.2687 | 1475.2056 | Westby | Mircea | E1F3C2C2-DC4F-64E4-4304B0FFFEEC735F |  |  | 1496.6792 | 29 | Yes | Undercoupled | E220EF16-00BF-C79C-6AA1AF5380761BD5 | 4.87 | 11500000000 | 13500000000 | 1170000000000 | 11.6 | 0.001 | 74700000000000 | 32300000000000 | 27 | 8690000000 | 9680000000 | 0.142 | 17.46 | 836 | 13300000000 | FE |  | Yes | Initial run to 30.5MV/m. Follow up runs to analyze post rad spike. | -5.5625 |  |  | 0.142 | 17.46 | 836 | [12 empty] | E3D51C5C-B659-09D2-464BAFA5F1C11DFD | E3DAF39B-A405-5F28-BDE3A130D00E7994 | E427531A-F538-B26B-C38196C62E175F16 | E4292651-AA9B-46ED-9F88C50C40B0DD20 | E42AFA6A-B95B-C034-2ACEB658CB308DDD | E42D346A-9A85-8944-04A0F0F5F9BDF088 | E4319A08-02EE-BB60-5D98AE3348DB696D | 1 | HAIPENG | 2011-09-29 09:58 | Yes | 29
99 | C100-RI-048 |  | Yes | No | 2 | C100_RI_048.txt | 2011-09-23 16:51 | 8 | TS28-001 | GRENOBLE | KDAVIS | 4e-06 |  |  | 175.7 | 2.07 |  | Yes | Yes |  | 1496.5973 | 1495.3556 | 1492.242 | 1487.6179 | 1482.5642 | 1477.9777 | 1474.574 | Grenoble | Mircea | F39A6954-9DD7-7D99-17FF745A37D1A818 |  |  | 1496.5945 | 29 | Yes | Undercoupled | F0E39042-C3AE-47E7-9DDF9192EB5987B3 | 4.48 | 9250000000 | 11700000000 | 943000000000 | 11.6 | 0.002 | 215000000000000 | 39200000000000 | 25.7 | 7050000000 | 7350000000 | 6.91 | 15.78 | 7000 | 9400000000 | Other | Superfluid He leak | No |  | -5.11 | [17 empty] | F11BD3B9-D154-EF5A-D418611F494D98FD | F11DEE87-E8EF-D1FE-B3690A068E1C07A6 | F11FDA5A-05FF-208B-2D7E441A7CDD002B | F1219F7A-0682-0D98-0D4F1C574C4EB606 | F122F006-0F3F-70AB-262A3D2F24490B2E | F12650BA-E19D-1A1A-50B29A982513B0E1 | F12955E5-F783-237C-CC9E8ED489F3D834 | 1 | HAIPENG | 2012-02-13 17:32 | No | 29
100 | C100-RI-050 |  | Yes | No | 2 | C100_RI_050.txt | 2011-09-28 18:44 | 8 | TS28-009 | GRENOBLE | KDAVIS | 1.7e-08 |  |  | 162.41 | 2.07 |  | Yes | Yes |  | 1496.6418 | 1495.3775 | 1492.1475 | 1487.5713 | 1482.4775 | 1477.6299 | 1474.3245 | Grenoble | KDavis | F6477D75-AB9F-BB9D-B7260E9A8C2331B2 |  |  | 1496.6409 | 29 | Yes | Overcoupled | 1B529AB0-C469-2E74-409AB2217375E3D4 | 4.95 | 11900000000 | 10500000000 | 1300000000000 | 11.6 | 0.005 | 127000000000000 | 1840000000000000 | 25.8 | 7350000000 | 9020000000 | 297 | 12.25 | 6580 | 12200000000 | Quench (non FE) |  | Yes | FE increases after 21MV/m. Qo reduced due to FE loading. Perform passband processing to clean up FE in cavity. | -5.22 |  |  | 51.2 | 14.21 | 324 | 30 | 35.1 | 30 | 23.8 | 13.25 | 7 | Repeated Q vs. E after performing the passband processing. | 9320000000 | 8680000000 | 9130000000 | 10500000000 | Above 21MV/m FE greater than 100mR/hr. Performed passband processing to clean up cavity. FE was reduced enough to measure Qo up to 26MV/m | 1EB01C56-A6D2-BC15-C8EF9CF6C2022E90 | 1F7C78A1-BD80-A9F6-F000B7837FC5D7FF | 1F817E20-0EB9-D7E4-A165F723BE9BB513 | 1F870B04-FC4B-E041-27A541C7ED381A9D | 1F96BE6E-B029-2984-D088ACBFB45CD3EA | 1F9B5494-A3B9-76DD-C68DA7E742A4D405 | 1FA008A6-D932-E788-26EC59D3C8AF4B61 | 1 | HAIPENG | 2011-09-29 13:05 | Yes | 29
101 | C100-RI-039 |  | Yes | No | 2 | C100_RI_039C.txt | 2011-09-27 | 7 | TS28-010 | ARI | STIRBET | 1.4e-08 |  |  | 185 | 2.07 |  | Yes | Yes |  | 1496.6284 | 1495.3826 | 1492.1465 | 1487.3937 | 1482.0941 | 1477.4186 | 1474.075 | Ari | Mircea | 13F6746D-FF5D-D22F-2A58C03D496001B8 | HOM survey done on previous test Sept 18th |  | 1496.6265 | 29 | Yes | Undercoupled | 1416D6E1-C1A8-F719-7BEADFB936B0364A | 5.1 | 11200000000 | 12400000000 | 709000000000 | 11.6 | 0.003 | 29400000000000 | 19400000000000 | 25 | 7950000000 | 9130000000 | 91.5 | 15.6 | 965 | 11500000000 | RF power | limited by operator due to high radiation | No |  | -4.9038 | [12 empty] | 9300000000 | 9710000000 | 10000000000 | 10900000000 | no processing occurred | 1461E6BB-0FA0-D1FB-D3691C7635BF218A | 14679C96-DA5B-6FE1-AFF4BDF660434E85 | 146D12D8-E228-E1D1-CC8E3B43632CE51F | 14702B99-0535-8186-9A5F914E11AE4C46 | 147B25D3-0079-3C16-FAE2D182CF3B9965 | 147F0D9E-E26D-725A-25901E453CDAEC87 | 14840637-A0D0-A2B2-BEE43D3AD607DCC4 | 1 | HAIPENG | 2011-09-29 09:37 | Yes | 29
102 | C100-RI-049 |  | Yes | No | 2 |  | 2011-09-29 18:45 | 7 | TS28-008 | KDAVIS | KDAVIS | 3.3e-05 |  |  |  | 2.07 |  |  |  |  | 1496.5661 | 1495.376 | 1492.1496 | 1487.6263 | 1482.5288 | 1477.8538 | 1474.5256 | Mircea | Mircea | 29AFD070-D3C6-B114-D4B951C6D724FFD1 | Cold Leak Developed as cavity passed below lambda, no high power test attempted. | Yes |  |  |  | Overcoupled | [45 empty] | 1 | HAIPENG | 2012-02-27 11:57 | Yes | 23
103 | C100-RI-051 |  | Yes | No | 2 | C100_RI_051.txt | 2011-10-01 11:40 | 4 | TS28-007 | GRIGORY | ARI | 2e-06 |  |  | 158 | 2.07 |  | Yes | Yes |  | 1496.6288 | 1495.4013 | 1492.2215 | 1487.5375 | 1482.3967 | 1477.6188 | 1474.304 | Ari | Eremeev | 844612FF-E279-C262-64B9460EAEFFED60 |  |  | 1496.6281 | 29 | Yes | Overcoupled | 81ADA71F-C4FF-EAF4-A425F379D75D469E | 4.05 | 12000000000 | 10200000000 | 997000000000 | 11.8 | 0.00285 | 9370000000000 | 15900000000000 | 23.5 | 6570000000 | 8000000000 | 3000 | 10 | 3000 | 11900000000 | Other | probably field emission induced quench | Yes |  | -5.37 |  |  | 193 | 11 | 1000 | 24.5 | 28 | 28 | 17.6 | 10.32 | 4.6 | 3pi/7 and 2pi/7 were power limited, the rest are limited by quench | 8180000000 | 8570000000 | 9070000000 | 9980000000 |  | 820291FF-CB4F-336E-8D05741C7C6FA66C | 82075009-926A-3E03-008082232142DF8E | 8211E62A-EDE7-52B1-A120B08E528CC9C5 | 821F2D75-F6EE-FAB1-D35E3D2CE2E1BAA0 | 82217DF8-D6BA-6D3E-88231E146416758F | 8223670D-047F-DD24-5C92C9EDE93DCFDD | 8231D130-091E-F83F-4EE5D30D48F28733 | 1 | HAIPENG | 2011-10-01 12:28 | Yes | 29
104 | C100-RI-052 |  | Yes | No | 2 | C100_RI_052.txt | 2011-10-03 | 4 | TS28-004 | KDAVIS | KDAVIS | 1e-09 |  |  | 172 | 2.07 |  | Yes | Yes |  | 1496.6055 | 1495.4095 | 1492.1353 | 1487.4536 | 1482.4921 | 1477.7021 | 1474.2982 | GCiovati | KDavis | 2131BB9B-D3F5-ECED-05438E8152823E1C | 3 modes in TM111 passband had QL of 1-5x10^6. Haipeng looked at those modes and verified they are quadrupoles with QL below the threshold | Yes | 1496.6051 | 29 | Yes | Overcoupled | 217B6D34-059C-E880-7699C2AF4A8B4E23 | 5.62 | 16300000000 | 11000000000 | 1030000000000 | 11.8 | 0.0043 | 10200000000000 | 21400000000000 | 34.1 | 7590000000 | 11600000000 | 0.00392 | 16 | 221 | 17000000000 | Quench (non FE) |  | No |  | -4.93 | [12 empty] | 11600000000 | 11200000000 | 12300000000 | 14100000000 | quench | 21B336FF-B698-1400-F6712A742FCDBF6C | 21DCF069-900E-A219-828D2F602630A548 | 21E49583-03F4-60CD-F50782B00AD29C3E | 21E974F0-9B4D-0168-5EDDC9C27A415D5C | 21ECAC64-A4D8-C9AC-0A28DA8EF5191038 | 21F022F9-C68E-2621-5914DE1263A453F8 | 221CDD81-B450-BAA5-D5049C1290584C29 | 1 | HAIPENG | 2012-08-15 16:16 | Yes | 29
105 | C100-RI-053 |  | Yes | No | 2 | C100_RI_053.txt | 2011-10-04 18:16 | 7 | TS28-001 | GRIGORY | GRIGORY | 5e-09 |  |  | 204 | 2.07 |  | Yes | Yes |  | 1496.5452 | 1495.2761 | 1491.9895 | 1487.3542 | 1482.0649 | 1477.345 | 1474.2423 | Eremeev | Eremeev | 2A4964B5-E274-819E-E1A97FA2B3B79CC6 |  |  | 1496.5444 | 29 | Yes | Undercoupled | 2B7BEDB8-C77C-4D21-7260C71730FCBAE3 | 4.9 | 15600000000 | 18300000000 | 1960000000000 | 11.6 | 0.00702 | 426000000000000 | 281000000000000 | 27 | 10600000000 | 12600000000 | 3.9 | 15 | 905 | 16300000000 | Operator | radiation exceeded 100 mR/h at Eacc = 27 MV/m | Yes |  | 3.8 |  |  | 29.1 | 13.5 | 905 | [7 empty] | 12500000000 | 13100000000 | 14500000000 | 16000000000 |  | 2B8A069D-AA2D-2447-12B434A2579D9BB8 | 2B939666-CF73-72C9-60AD86C1F7C54B73 | 2B952DAC-F446-1F98-4F87D0457460B702 | 2B97101C-989A-14B0-3D50428A532F9D07 | 2B9C7B99-C4D0-FE79-BECB48A2E7B784E8 | 2B9EB786-F135-7343-4061557DBB7021B7 | 2BA8FFD5-E32F-312F-D8EE502ED7D47002 | 1 | HAIPENG | 2012-08-15 16:23 | Yes | 29
106 | C100-RI-048 |  | Yes | No | 2 | C100_RI_048A.TXT | 2011-10-05 | 8 | TS28-009 | ABURRILL | TGOODMAN | 5e-08 |  |  | 157 | 2.07 |  | Yes | Yes |  | 1496.5999 | 1495.3593 | 1492.2456 | 1487.6207 | 1482.567 | 1477.978 | 1474.575 | Grenoble | Mircea | 7E343329-0260-A6A3-9B71C70148CE981E | measured previously on 23-Sep-2011 |  | 1496.5978 | 29 | Yes | Overcoupled | 7B527FC4-B231-46DF-60274CFA76E598A6 | 5.33 | 14400000000 | 10500000000 | 1500000000000 | 11.7 | 0.002 | 964000000000 | 31100000000000 | 26 | 9880000000 | 10700000000 | 0.543 | 10 | 6000 | 14300000000 | Quench (non FE) |  | Yes | processed in 6/7 pi mode early on, FE decreased significantly, see processed data file | -5.52 |  |  | 0.0543 | 16.63 | 1150 | [11 empty] | not enough liquid to carry out low temperature Qo measurements. | 7B0BFB4E-DCF8-4D78-B44E778942A7FD15 | 7B0D6719-BF56-2951-CB2AE28EB1BFBEA7 | 7B0EFE3D-0974-93C5-8678EAD8BEFB6F0F | 7B100DD3-DF30-53F8-30C64738804A6363 | 7B1179DC-CEF7-8F90-96DC8F37AC3A4DD3 | 7B12B991-B988-057A-F6ACAAAD38A9405A | 7B1425A7-DC97-909D-F4213A8C52469C9B | 1 | HAIPENG | 2011-10-07 09:07 | Yes | 29
107 | C100-RI-054 |  | Yes | No | 2 | C100_RI_054 | 2011-10-18 10:30 | 3 | TS28-003 | WESTBY | TBASS | 1e-08 |  |  | 159.9 | 2.07 |  | Yes | Yes |  | 1496.57 | 1495.2813 | 1491.9976 | 1487.3788 | 1482.1314 | 1477.3098 | 1473.9969 | Westby | Bass | BAF63853-B78E-D122-ED78A68AC48C0EA5 |  |  | 1496.5689 | 29 | Yes | Overcoupled | BB08E7E2-D2D1-3F9A-3B20199075C80874 | 5.69 | 17900000000 | 10800000000 | 1280000000000 | 12 | 0.001 | 18200000000000 | 1160000000000 | 27.2 | 9470000000 | 10800000000 | 0.00915 | 21.08 | 48.9 | 16900000000 | Other | Admin Limited | No | [14 empty] | 10600000000 | 11700000000 | 12600000000 | 14200000000 |  | BBE00308-A29A-52E9-BF426A12DA8EEC41 | BBE29782-E501-51CA-BB6F49DBAAB5DCB3 | BBE62BA9-C8BF-5A7F-B7B97250DD3F9359 | BBE8BB52-9A81-F1E7-CD4660F7F9E10603 | BBEA7C49-9C4C-2F60-B01D520A80225F41 | BBED7A02-D6BA-0479-CC25CD8FA41199B1 | BBF02826-B137-004D-B8758FBFF596BA27 | 1 | HAIPENG | 2012-08-15 16:24 | Yes | 29
108 | C100-RI-055 |  | Yes | No | 2 | C100_RI_055.txt | 2011-10-12 19:07 | 8 | TS28-001 | ABURRILL | TBASS | 6.81e-09 |  |  | 188 | 2.07 |  | Yes | Yes |  | 1496.569 | 1495.3732 | 1492.2372 | 1487.6061 | 1482.6187 | 1477.9562 | 1474.7453 | ABurrill | Bass | C81BC407-C14E-4260-FFF3F9E287B8EBE5 |  |  | 1496.5664 | 29 | Yes | Undercoupled | E4448D9A-09D5-DB53-429914C289FACCE0 | 4.71 | 9730000000 | 11000000000 | 690000000000 | 11.6 | 0.001 | 10600000000000 | 37500000000000 | 28 | 6530000000 | 7340000000 | 175 | 6.93 | 1080 | 9660000000 | Other | Admin | Yes |  | -5.43 |  |  | 2.21 | 15.5 | 1080 | 34 | 25.8 | 23.3 | 18 | 12 |  |  | 7340000000 | 6700000000 | 6850000000 |  | early field emission onset cleared up. Very low Qo, barely meeting spec. Admin limited at 28 MV/m | D0A1EFE6-E4EC-CE05-76BA6F19B0BE7A11 | D0A6D830-07EA-6A29-DD508F868DF9F7BA | D0A8377D-F1AF-2817-899935FBE47F8E90 | D0A9950A-F6D0-7CBB-4331C3D7F5E15A57 | D0AB00FA-E669-010B-A68EAE69D23A2C86 | D0AC59ED-9BDA-B6D1-4696724ED77DBB5F | D0ADC436-D21D-321F-8EA22D4C1801A9F5 | 1 | HAIPENG | 2012-08-15 16:31 | Yes | 29
109 | C100-RI-056 |  | Yes | No | 2 | C100-RI-056.txt | 2011-10-14 20:20 | 7 | TS28-010 | TBASS | TBASS | 8.2e-09 |  |  | 160 | 2.07 |  | Yes | Yes |  | 1496.5399 | 1495.3639 | 1492.0956 | 1487.5818 | 1482.4162 | 1478.0256 |  | Bass | Bass | B255CCFA-F253-651F-653730191A21A391 |  |  | 1496.5372 | 29 | Yes | Undercoupled | B259D164-C94A-4D11-75DB81C4925818FD | 5.92 | 14000000000 | 16200000000 | 803000000000 | 11.6 | 0.00097 | 57300000000000 | 13500000000000 | 28.5 | 12800000000 | 11000000000 | 0.00316 | 25 | 0.0448 | 14500000000 | Cable | Observed unstable Ptrans at Eacc>24MV/m | No |  | -4.47 | [16 empty] | Cavity tested to 25MV/m with very little radiation. Radiation Onset at 25MV/m. Ptrans cable became unstable at higher gradients. Stopped test at 27MV/m | B35E0D08-F7B2-C4F5-BE53A7B836385C7D | B3624EC2-9C40-AEF5-C7652800438FEF64 | B3720931-E3F5-A02A-23BCDD90D462345E | B37336C3-A5FA-53E4-ECF0EAA4ADDF36B1 | B37439D4-E28C-5623-F248F16A046EA795 | B3756D4B-BA2C-9005-A3C0019B320697A9 | B376C9CB-F013-251C-555EEBE93C15DE21 | 1 | HAIPENG | 2012-08-15 16:35 | Yes | 29
110 | C100-RI-049 | NA | Yes | No | 2 | C100_RI_049 | 2011-10-15 07:00 | 8 | TS28-008 | ARI | ARI | 2e-08 |  |  | 175 | 207 |  | Yes | Yes | I re-did the cable cal three times because the cavity was over-coupled but really close to critically coupled. There was also a high Q0, and large change in Ci with power. | 1496.5684 | 1495.381 | 1492.1514 | 1487.6287 | 1482.5304 | 1477.8561 | 1474.5235 | Ari | Ari | E9BB38C3-B1BB-7D4E-E64F9EB724F78F33 |  | Yes | 1496.5661 | 29 | Yes | Overcoupled | E9BE9596-A07D-90FA-04D29C2996742A73 | 5.23 | 17900000000 | 12400000000 | 1370000000000 | 11.8 | 0.001 | 3660000000000000 | 9390000000000 | 27.2 | 8000000000 | 11400000000 | 50 | 13 | 1000 | 17800000000 | FE |  | No |  | -4.1 |  |  | 6 | 13 | 1000 | [6 empty] | did not do pass band processing | 12100000000 | 13400000000 | 14000000000 | 15900000000 | some processing initially and possible some more after time- see final curve for rad and Q change. | EA0D0CB8-FF2D-4CCF-9984C7EB729E4CF6 | EA12627C-DB19-C384-73D808543854FB89 | EA15B50B-EF4D-6EB6-108CC2FF6CEA6CF2 | EA16E1F6-B6F1-665F-43E2EC0041D5C4B9 | EA1814DF-C492-1A2E-91D1A9F514D9EF84 | EA197FF5-0C29-91F7-3E2BA382AE0253DF | 1B23EA13-B360-2F40-BDE5F18E16AE926C | 1 | HAIPENG | 2011-10-19 23:07 | Yes | 29
111 | C100-RI-057 |  | Yes |  | 2 | C100_RI_057 | 2011-10-18 13:18 | 4 | TS28-007 | WESTBY | TBASS | 1.94e-08 |  |  | 159.4 | 2.07 |  | Yes | Yes |  | 1496.613 | 1495.3856 | 1492.1756 | 1487.6176 | 1482.5305 | 1477.9782 | 1474.8442 | Mircea | Bass | F09B7806-01E3-1CDA-8DCADCE58BBC9600 |  | Yes | 1496.6121 | 29 | Yes | Overcoupled | F0C1B3E2-A57B-2753-7C47C5411650F17D | 4.87 | 12300000000 | 9560000000 | 1170000000000 | 11.6 | 0.001 | 31100000000000 | 19000000000000 | 27.1 | 8360000000 | 9540000000 | 19.2 | 14.1 | 591 | 12500000000 | FE |  | No |  | -5.2717 | [17 empty] | F13D75BC-B94C-763A-D799FEE5991A0807 | F1423855-B1C0-9FF1-A78920215A75A7AF | F14573E8-AC3A-9A5D-C2874078E6FFE701 | F147A3FE-F39F-F0B9-6B0CC73FCA9705AE | F149703F-FAD1-3AC9-F6F48E3A7A8F4A87 | F14C062D-C856-3E47-A118E72AC9B00F8A | F1509509-0ED3-4C71-E9098E0CE5F0B285 | 1 | HAIPENG | 2011-10-19 23:03 | Yes | 29
112 | C100-RI-059 |  | Yes | No | 2 | C100_RI_059 | 2011-10-19 19:49 | 8 | TS28-009 | ARI | ARI | 5e-09 |  |  | 170 | 2.07 |  | Yes | Yes |  | 1496.5413 | 1495.2918 | 1492.0971 | 1487.5344 | 1482.3982 | 1477.878 | 1474.7263 | Mircea | Mircea | 4CDE9FE8-E8BF-8377-B7AA4D1D13B994BA |  |  | 1496.538 | 29 | Yes | Undercoupled | 320F8BCE-BBFA-FF04-D7329A238CA530FA | 4.59 | 10500000000 | 12000000000 | 1100000000000 | 11.6 | 0.001 | 19300000000000 | 27500000000000 | 27 | 7760000000 | 8320000000 | 0.02 | 19.5 | 1 | 10900000000 | RF power | NA | No | NA | [13 empty] | 8290000000 | 8910000000 | 9270000000 | 10000000000 | NA - good cavity low FE | 32398BE0-BA6F-CCCD-AB3C06ECC480288E | 3240ECDB-0837-4896-5DC42EF913E1176C | 324474F5-E3DA-61F5-2D59BB9E57109981 | 32461541-9B5D-3495-6095D2733DEB9952 | 3247F58B-F5DF-D188-FBBC59D0CB4CEE9E | 3249A65D-057D-7276-DFE1292AD923A327 | 4D03D9DC-F8FB-7900-55DEF4B7B646066A | 1 | HAIPENG | 2011-10-19 23:00 | Yes | 29
113 | C100-RI-060 |  | Yes | No | 2 | C100_RI_060.TXT | 2011-10-21 20:33 | 4 | TS28-004 | STIRBET | ABURRILL | 7e-09 |  |  | 170 | 2.07 |  | Yes | Yes |  | 1496.575 | 1495.3155 | 1492.18 | 1787.515 | 1482.376 | 1477.845 | 1474.613 | Mircea | Mircea | 9A951602-C2FF-C9E0-F80C4ACC3F400B1A |  |  | 1496.5743 | 29 | Yes | Overcoupled | 9A9F5349-FCBC-1AC1-20B01F844EFB011C | 5.61 | 12900000000 | 10600000000 | 1030000000000 | 11.6 | 0.002 | 81300000000000 | 30900000000000 | 27.2 | 8130000000 | 9060000000 | 687 | 10.93 | 687 | 13200000000 | RF power |  | No | Cavity had a field emitter blow up at 18.9 MV/m | -5.56 |  |  | 0.001 | 25.4 | 0.045 | [7 empty] | 8240000000 | 9460000000 | 10000000000 | 11300000000 |  | 9B0BDD55-B3D1-44BD-AFABF662CAB7BD19 | 9B0D0963-AFA1-357A-C15093B6E0136CEE | 9B0E3BFF-98DF-8759-99DF41D2E54AB437 | 9B0F45BC-BD2A-4B62-E84605128EE93D61 | 9B11656F-E315-422B-C03F01ECBA07330A | 9B1280B1-A6AD-0B82-22C5231C2D885DDC | 9B13B620-BD88-15D9-66667E2ACC683C0A | 1 | HAIPENG | 2011-10-22 16:03 | Yes | 29
114 | C100-RI-058 |  | Yes | No | 2 | C100_RI_058.txt | 2011-10-22 16:27 | 8 | TS28-001 | STIRBET | STIRBET | 1e-09 |  |  | 193.34 | 2.07 |  | Yes | Yes |  | 1496.5957 | 1495.3582 | 1492.1201 | 1487.6001 | 1482.4858 | 1477.5757 | 1474.3654 | Mircea | Mircea | C571459B-9F5E-D6CF-656595A05EA2C4C4 |  |  | 1496.5937 | 29 | Yes | Undercoupled | C57A6CB0-D42F-A68D-9888154C673E99A3 | 5.1 | 10900000000 | 13100000000 | 624000000000 | 11.58 | 0.001 | 15400000000000 | 26100000000000 | 24.21 | 5540000000 | 7250000000 | 10.9 | 13 | 112 | 11100000000 | Quench (non FE) | Cavity with radiation onset at 13 MV/m , Qo bellow specifications at 19.2 MV/m ( only 7.25e9), cavity quench limited at 24 MV/m |  |  | -5.74 | [12 empty] | 7820000000 | 8130000000 | 8570000000 | 9040000000 | Radiation onset at 13 MV/m, cavity quench limited at 24 MV/m, radiation at 24 MV/m about 100mR/h (clean cavity). Qo at 19.2 MV/m bellow specifications. | C5C3F9A7-0B7D-2698-0ADBA28D6470B1B9 | C5C61E3E-07FC-CC7A-09F453D851D2B0F9 | C5C85F56-C08B-6DAD-D41FC4A5A9B748C2 | C5CA3336-974D-7FD0-C03D048349A610C2 | C5CC04CA-D8B7-6AB1-BBF94679FBE13FE4 | C5CDFEC5-C593-F161-717F7D9F3132B3EE | C5D0CB49-F9F7-6CE0-22AB8B4C9BFE97AC | 1 | HAIPENG | 2012-08-15 16:41 | No | 29
115 | C100-RI-061 | NA | Yes | No | 2 | C100_RI_061.txt | 2011-10-26 12:51 | 8 | TS28-010 | ARI | ARI |  |  |  | 180 | 2.07 |  | Yes | Yes |  | 1496.6133 | 1495.3681 | 1492.1053 | 1487.4377 | 1482.2832 | 1477.7978 | 1474.5809 | Ari | Ari | 8B92967A-FAD2-D806-58F5E079C7BEB152 | many (5-10) above 10^6 but all I think they were quadruples. | Yes | 1496.6112 | 29 | Yes | Undercoupled | 8E5FBBFF-CF4A-A5AA-EEE238F5E254A09F | 5.55 | 12100000000 | 12300000000 | 1110000000000 | 12 | 0 | 26700000000000 | 53700000000000 | 23 | 8.9 | 8.9 | 12.5 | 12 | 10000 | 12500000000 | FE | Radiation limit to 23MV/m at 10R/hr after a turn on event at 23.5 MV/m - no complete final power rise, only from 12Mv/m to 23 after processing | Yes | process all pass bands to 27MV/m max in any cell, with no change in radiation. | -5.1 |  |  | 3000 | 12 | 4000 | 27 | 25 | 22 | 18 | 12.5 | 6.5 | kept all cells below 27MV/m no quench or processing events - low radiation in pass band data | 8180000000 | 8800000000 | 9220000000 |  | FE limited to 23MV/m just passes spec at 19.5MV/m, but with 2R/hr radiation. The plan is to re-raise the cavity and test again even though it technically passes spec. | 90A29B94-CD4A-90C5-4E9F051A1426DE1D | 90A6389F-0D7A-AB20-59492216F2D6485A | 90A988A8-D9BE-D2FC-CA0B20C006E638DE | 90AE35D7-C2A3-FF25-6C62906068E709DC | 90B09D4E-C3F3-9C5D-4D9A8BF86D4D5A7D | 90B2D85F-F71B-0B11-3D21F8F2F7A078A7 | 9036C41C-0AC9-1E03-3ECDC31CDA78CDFA | 1 | HAIPENG | 2012-08-15 16:51 | No | 29
116 | C100-RI-063 |  | Yes | No | 2 |  | 2011-10-27 19:06 | 4 | TS28-007 | GENG | TGOODMAN | 1.5e-08 |  |  | 173 | 2.07 |  | Yes | Yes |  | 1496.5807 | 1495.3305 | 1491.9969 | 1487.3119 | 1482.1056 | 1477.2737 | 1473.9514 | Geng | TGoodman | CC9D4223-A8C2-25CB-5B30B1BB7B109062 |  | Yes | 1496.5794 | 29 | Yes | Overcoupled | CCD744EF-E7B9-7D09-12358B39A61D9C60 | 5.45 | 17000000000 | 7600000000 | 1180000000000 | 12.8 | 0.001 | 16300000000000 | 8650000000000 | 26.9 | 12600000000 | 13300000000 | 12.3 | 15.1 | 92.8 | 17400000000 | Other | admin limit per CEBAF 12 GeV project | No |  | -4.5 | [12 empty] | 13500000000 | 15100000000 | 15900000000 | 18400000000 | is a pleasure to measure | CD4F124E-00E1-9A11-7AFCC64A5129CC37 | CD57456B-CB8F-6D14-9D054A1F141A841B | CD5BBD7B-F674-0153-652ADDE899864C01 | CD5FC3BE-025C-8C5F-AC7962852E24F6C6 | CD632CC9-D595-68D9-38C4788D680531E4 | CD6923FD-E5DB-EDA1-7A41690AB3FC07EB | CD7158FB-E3BA-FE0A-BB0196B875CD466C | 1 | HAIPENG | 2012-08-15 17:04 | Yes | 29
117 | C100-RI-065 |  | Yes | No | 2 | C100_RI_065.txt | 2011-10-31 | 4 | TS28-004 | GCIOVATI | GCIOVATI | 6.9e-09 |  |  | 168 | 2.07 |  | Yes | Yes |  | 1496.5174 | 1495.2962 | 1492.1313 | 1487.333 | 1482.2101 | 1477.7937 | 1474.5465 | KDavis | KDavis | BB81AC39-058E-98C5-4033735B2A214CCA |  | Yes | 1496.5167 | 29 | Yes | Overcoupled | BB92B514-AA49-6A2F-7E9008A41FA0E1DD | 5.76 | 14300000000 | 10500000000 | 1120000000000 | 11.6 | 0.00144 | 119000000000000 | 78000000000000 | 27.4 | 9200000000 | 10100000000 | 0.00287 | 50 | 0.0115 | 15000000000 | Other | Admin limit | No |  | -5.378 | [12 empty] | 10200000000 | 11500000000 | 12300000000 | 14200000000 | no FE, admin limit | BBCBDCFF-B64B-0A55-39BAEB94D116A4EC | BBD02890-B7BA-2469-26F04AA6FB58CE18 | BBD4308B-F611-0003-68F9BBE78E3BF9C2 | BBD6BEE7-F093-A636-1252578BDE17BFE8 | BBD96A39-9E6B-3553-786765E6EF78D7E9 | BBDC1E75-FFF1-5369-941DEE06C8EEF4BB | B8B48E9E-B297-7736-6DC53FD9390707ED | 1 | HAIPENG | 2012-08-15 17:12 | Yes | 29
118 | C100-RI-061 | Rinsed cavity after last test produced high FE. | Yes | No | 2 | C100_RI_061A.txt | 2011-11-01 10:50 | 7 | TS28-001 | TGOODMAN | TBASS | 7.96e-09 |  |  | 175.7 | 2.07 |  | Yes | Yes |  | 1496.6146 | 1495.3703 | 1492.1069 | 1487.4431 | 1482.2845 | 1477.7988 | 1474.5811 | Ari | Ari | B800F495-9186-2C86-3115FFF830B4CF1B | HOM measurement preformed previously on 26-Oct-2011. |  | 1496.6121 | 29 | Yes | Overcoupled | EB2222B5-BAC0-C016-9FF2A26CB75F9616 | 5.91 | 14400000000 | 12400000000 | 1010000000000 | 11.6 | 0.005 | 26900000000000 | 42700000000000 | 27 | 9480000000 | 10900000000 | 0.873 | 19 | 900 | 15500000000 | Other | Administrative limit at 27 MV/m | No |  | -5.3 | [16 empty] | Cavity was close to critically coupled. The coupling data file shows the cavity to be undercoupled. The cavity is slightly overcoupled when on removes the reflected power overshoot. | EB916A2E-C1C6-1825-5F2364FE6195ACA1 | EB932592-AE3F-1C04-CD874601683D3E17 | EBB4582A-F61B-B450-235912F2A3F13F4D | EBB59243-9DCF-27D7-58F0435604B2B343 | EBB6C847-9734-AD47-781E8C40C53A4E69 | EBB85BB8-C52F-DFDD-B122C802E6920E96 | EBB9A1FF-E8ED-7173-93E73B75D7B9C192 | 1 | HAIPENG | 2012-08-15 16:47 | Yes | 29
119 | C100-RI-058 |  | Yes | No | 2 | C100_RI_058.txt | 2011-10-22 11:50 | 8 | TS28-001 | STIRBET | KDAVIS | 1e-09 |  |  | 222 | 2.07 |  | Yes | Yes |  | 1496.596 | 1495.358 | 1492.12 | 1487.6 | 1482.486 | 1477.576 | 1474.365 | Mircea | Mircea | F11E9AFA-B1F8-BB67-E79C615E38375F98 |  | Yes | 1496.594 | 29 | Yes | Undercoupled | EE5B4485-F6F6-145D-9EEF43A3DA282804 | 5.27 | 11400000000 | 12200000000 | 708000000000 | 11.6 | 0.0085 | 17000000000000 | 29000000000000 | 25.3 | 5250000000 | 8010000000 | 13.2 | 13 | 125 | 11500000000 | Quench (non FE) |  | No |  | -5.3 |  |  | 5.56 | 14 | 125 | [7 empty] | 7820000000 | 7800000000 | 7560000000 | 8430000000 | quenching above 25.3 MV/m | EF8BB69C-A926-0196-B95B37DE00AF8688 | EF8E66B0-BA33-5007-0EFAE46219B77C5E | EF912D83-9B93-CFC5-FDB78A7E33A35525 | EF936407-0D29-33FE-F8E970EC366F3744 | EF962208-FDA0-DDF8-A5287645460215E6 | EF9CE7BF-036D-CA42-B69A52F94D34087E | EFA3821F-E630-30C0-70ACFD7D17B6828D | 1 | HAIPENG | 2012-08-15 16:39 | Yes | 29
120 | C100-RI-064 |  | Yes | No | 2 | C100_RI_064.txt | 2011-10-28 18:00 | 8 | TS28-009 | KDAVIS | KDAVIS | 6.5e-09 |  |  | 178 | 2.07 |  | Yes | Yes |  | 1496.54 | 1495.31 | 1492.0339 | 1487.186 | 1481.962 | 1477.267 | 1474.133 | Mircea | Mircea | F0F96B31-F0B7-54DD-D5256A3932C1980D |  | Yes | 1496.539 | 29 | Yes | Overcoupled | F49C6780-B9FE-5FA3-8599BA0F6D42ACEC | 5.29 | 12300000000 | 9030000000 | 1190000000000 | 11.6 | 0.0017 | 18000000000000 | 34000000000000 | 27 | 9860000000 | 8330000000 | 0.0399 | 20 | 573 | 12600000000 | Other | 27 MV/m admin limit | No |  | -5.4 |  |  | 0.0399 | 20 | 573 | [7 empty] | 8330000000 | 9220000000 | 9870000000 | 11300000000 |  | F4D6BA80-D3B2-F179-0B3FB2423014EFF8 | F4D95BF6-9814-90BE-671DE431B7DF1D40 | F4DD0475-F789-E6A7-295615D67D0EA8FE | F4DF6584-9230-528C-3EDB247AE3D996CB | F4E29BBB-AF34-9949-8A4099B24F2BC6B0 | F4E4A415-CF57-7D5D-5458E10055A1DE25 | F4E6E7D1-A0EB-B9ED-8FAA4A3542B5EAC2 | 1 | HAIPENG | 2012-08-15 17:08 | Yes | 29
121 | C100-RI-066 |  | Yes | No | 2 | C100_RI_066.txt | 2011-11-03 18:59 | 8 | TS28-010 | GRIGORY |  | 5.4e-09 |  |  | 191 | 2.07 |  | Yes | Yes |  | 1496.5447 | 1495.3915 | 1492.2397 | 1487.6416 | 1482.8785 | 1478.4 | 1475.1339 | Eremeev | Eremeev | 34B9673D-9DD5-7E7B-4C0E05510EEB9A5D |  |  | 1496.544 | 29 | Yes | Undercoupled | 34C2CC31-9FE4-0D61-413C7FBEFA350B5C | 5.07 | 11000000000 | 11200000000 | 1100000000000 | 11.8 | 0.00175 | 15600000000000 | 23500000000000 | 27 | 7770000000 | 8410000000 |  | 50 | 0 | 11400000000 | Operator | The cavity reached 27 MV/m admin limit | No |  | -5.47 | [12 empty] | 8380000000 | 9000000000 | 9330000000 | 10300000000 | no detectable X-rays up to 27 MV/m | 34F07242-B0C2-D726-ECDBB2121C16B741 | 34F2B405-C89D-48F1-A2FE7F3ABAC804FA | 34F45D57-F55B-1F3F-64CF5F23179E09E6 | 34F5FBFD-F1E1-2517-4509CD85C29FBB30 | 34F7EC4B-D72D-2EBC-FA89EDD9A7DE43C8 | 34F98872-B50E-1E25-43C8E7821480FD6F | 34FDE8B8-B062-E964-CC340B881D119E3E | 1 | HAIPENG | 2012-08-15 17:22 | Yes | 29
122 | C100-RI-067 |  | Yes | No | 2 | C100-= |  | 8 | [73 empty] | 1 | HAIPENG | 2012-08-15 17:28 | No
123 | C100-RI-067 |  | Yes | No | 2 | C100-RI-067.txt | 2011-11-10 15:18 | 4 | TS28-007 | TGOODMAN | WESTBY | 1.6e-08 |  |  | 182.71 | 2.07 |  | Yes | Yes |  | 1496.5156 | 1495.2943 | 1492.1262 | 1487.4981 | 1482.3262 | 1477.7987 | 1474.5597 | TGoodman | Mircea | 9B1B2593-DD89-0A94-59CC9386C5C38E11 |  |  | 1496.5134 | 29 | Yes | Overcoupled | 97A2DDA3-C365-59C2-03B77B800D2D9776 | 5.28 | 15800000000 | 9260000000 | 1010000000000 | 12 | 0.001 | 1620000000000000 | 10700000000000 | 27 | 11500000000 | 12500000000 | 0.001 | 50 | 0.0517 | 16300000000 | Other | Admin. Limit @ 27MV/m | No |  | -5.1 | [12 empty] | 12700000000 | 13900000000 | 14700000000 |  |  | 980EA9F5-F7A3-DF8E-95C4845D74059266 | 9818A48E-A45A-D543-7C1200F083E490D8 | 981CF329-9832-B432-E67F7B30CFD501E9 | 981FECB2-CAF4-578F-36A83DA101396FE1 | 98235F3F-ACA6-BD1A-C3693D97D8BC0F3E | 982871EB-B54F-C603-0D246D86D7927151 | 983184F7-0F06-42D3-DC538375FFE08FBA | 1 | HAIPENG | 2012-08-15 17:27 | Yes | 29
124 | C100-RI-069 |  | Yes | No | 2 |  | 2011-11-11 | 8 | TS28-008 | GENG | GENG | 9.1e-09 |  |  | 164 | 2.07 |  | Yes | Yes |  | 1496.4019 | 1495.1975 | 1491.8797 | 1487.1512 | 1481.896 | 1477.2781 | 1473.9631 | GCiovati | GCiovati | F483C9A9-D90E-625C-EE5324F309535D50 |  | Yes | 1496.4015 | 29 | Yes | Overcoupled | 5835184A-9B00-DF55-6619FC83709017E3 | 5.34 | 13200000000 | 10400000000 | 1800000000000 | 11.6 | 0.001 | 36600000000000 | 20500000000000 | 26.9 | 9520000000 | 10200000000 | 0.002 | 22.7 | 0.1 | 13600000000 | Operator |  | No |  | -4.89 |  |  | 0.001 | 22.4 | 0.125 | [7 empty] | 10200000000 | 11100000000 | 11600000000 |  | is about average in Q0 values, admin limited to 27 MV/m gradient, 19 Torr dewar pressure was unable to reach due to CMTF testing. | 58AA62AA-00E2-BE42-513CE00AA965CD59 | 58AE4FA1-AA86-31EB-94CCFC1C7531D7C2 | 58B35B3C-F958-4D15-1E8853ED93E4DF86 | 58B89BDC-CCF2-F2C1-E979C898FF3325F1 | 58C45EEA-0A76-EA75-883D8C23D8002640 | 58C738D5-EECE-A450-F253CDA26FB9BE66 | 58CDD7E2-E342-BEEE-E5FEDA10E7D2099D | 1 | HAIPENG | 2012-08-15 17:37 | Yes | 29
125 | C100-RI-070 |  | Yes | No | 2 |  | 2011-11-15 | 7 | TS28-010 | GENG | GENG | 6.9e-09 |  |  | 154 | 2.07 |  | Yes | Yes |  | 1496.5916 | 1495.3491 | 1492.0984 | 1487.4537 | 1482.2703 | 1477.6339 | 1474.3624 | Geng | Geng | 97C96C80-02D7-3592-C90982F3DF2257E8 |  | Yes | 1496.5893 | 29 | Yes | Overcoupled | 97D9DFA9-B161-1D18-D56CED8D9ECE304D | 5.92 | 15100000000 | 11500000000 | 1140000000000 | 11.6 | 0.001 | 20900000000000 | 65200000000000 | 27 | 11000000000 | 11600000000 | 0.02 | 19.5 | 2 | 15300000000 | Operator |  | Yes | rf power increased slowly starting at 20 MV/m to allow processing | -4.57 |  |  | 0.001 | 21 | 2 | [7 empty] | 11600000000 | 13100000000 | 13800000000 |  | typical low field Q0, first x-ray at 19.5 MV/m, RF processing as incident power was raised, max. gradient 27 MV/m, admin limit, minimum field emission. | 9824AEFF-969A-68FB-75A413DDDC704104 | 9828258F-F2D1-9288-F0908ECC4323435C | 982B530A-0C7A-2D59-A28EEB833F08889C | 982E12B1-A1A5-6F4D-A4570AE475360558 | 9830A0CF-BBF0-1EFF-A1A882A5D8D6DE6D | 9832EC5B-AB5C-736C-36A717EF3ED1C87A | 98361D39-CCDE-D709-ABC0F46A8B1CBC90 | 1 | HAIPENG | 2012-08-15 17:40 | Yes | 29
126 | C100-RI-068 |  | Yes | No | 2 | C100_RI_068.txt | 2011-11-16 11:00 | 4 | TS28-004 | GCIOVATI |  | 6.4e-09 |  |  | 168 | 2.07 |  | Yes | Yes |  | 1496.6021 | 1495.3781 | 1492.0981 | 1487.368 | 1482.2032 | 1477.6543 | 1474.4389 | GCiovati | GCiovati | D37F3E82-0FEB-6129-D39C41F635105225 |  | Yes | 1496.602 | 29 | Yes | Overcoupled | D388C689-D608-16C9-3876CF2058A89552 | 5.63 | 15300000000 | 11100000000 | 1140000000000 | 11.7 | 0.00428 | 46500000000000 | 42300000000000 | 25.5 | 9410000000 | 10200000000 | 0.644 | 19.7 | 25 | 16000000000 | Quench (non FE) |  | No |  | -5.45 | [12 empty] | 9400000000 | 10900000000 | 11600000000 |  |  | D44A7F4D-A13A-E30E-D5B1B6D94458DC26 | D44C95B3-F79D-C899-279FD9047C5C1F39 | D44EBF2B-DADB-54E4-135A00520D3D945B | D45047A8-D5DD-D9EF-5819D7428A68DF07 | D45217C2-9833-8BCB-D4996FFD8B4093F0 | D454A072-AF85-8580-B4E10D224889F651 | F13A5FCB-09D3-3735-5CFD85503BCADA7E | 1 | HAIPENG | 2012-08-15 17:33 | Yes | 29
127 | C100-RI-071 |  | Yes | No | 2 | C100_RI_071.txt | 2011-11-19 10:24 | 7 | TS28-001 | WESTBY |  | 9.3e-08 |  |  | 191.5 | 2.07 |  | Yes | Yes |  | 1496.5603 | 1495.314 | 1492.0446 | 1487.1089 | 1482.0014 | 1477.4081 | 1473.9924 | Mircea | Westby | C37C5F70-E384-DE05-D7224189AEE06F0F |  | Yes | 1496.5597 | 29 | Yes | Undercoupled | C388497B-D358-E120-678BBD10F5D35D1B | 6.12 | 13800000000 | 14300000000 | 634000000000 | 11.6 | 0.007 | 174000000000000 | 1070000000000000 | 27 | 9280000000 | 9880000000 | 0.133 | 18.49 | 26.5 | 13500000000 | Other | Admin via VTRF Procedure | No | [19 empty] | C3BE2C73-F5B5-634C-730AE35F9342DE24 | C3C0EE15-0402-5542-8E5801011355AC99 | C3C31746-0D98-289E-5D693DFE7B67C26B | C3C52E82-0F8C-0880-6A8321344C82F9A7 | C3C7A977-9E6B-0508-35B24C10461C7E01 | C3CA7D4F-C462-D352-94A470A457232E18 | C3CD0326-975E-ABEF-7710DE0BFA299A8A | 1 | HAIPENG | 2012-08-15 18:28 | Yes | 29
128 | C100-RI-073 |  | Yes | No | 2 | C100_RI-073.txt | 2011-11-21 | 8 | TS28-009 | STIRBET | STIRBET |  |  |  | 167.7 | 2.07 |  | Yes | Yes |  | 1496.5821 | 1495.34 | 1492.149 | 1487.6015 | 1482.2734 | 1477.7567 | 1474.0075 | Mircea | Mircea | DA9A4547-DEC2-CFFC-3A7EE316CCD13662 |  |  | 1496.5798 | 29 | Yes | Overcoupled | 287B574C-BED2-81E4-D897D242CAAB4D7E | 5.19 | 13100000000 | 10800000000 | 1740000000000 | 11.6 | 0.005 | 61900000000000 | 24900000000000 | 27.3 | 9050000000 | 10300000000 | 0.0609 | 20.01 | 52.7 | 13800000000 | Operator | Administrative limit 27 MV/m | No |  | -5.2846 | [17 empty] | DB1D51A6-E07B-D090-3E1AEFE0A5C1E120 | 288B93B0-A042-84A6-710605D229534037 | 28914821-B02C-1BF1-C2A26F2B75265024 | 289530C5-ADB1-0BBB-441AFE3F9BAD8795 | 28994194-F6BF-7F2D-F0BD12B8B2FF17C1 | 289CB8CE-B153-8C29-8C98B2AF340A9841 | 28A0CE45-CB79-BDAA-2A54461C0F560C3A | 1 | HAIPENG | 2012-08-15 18:39 | Yes | 29
129 | C100-RI-072 |  | Yes | No | 2 | C100_RI_072.txt | 2011-11-19 09:15 | 4 | TS28-007 | WESTBY |  | 2.7e-08 |  |  | 176.6 | 2.07 |  | Yes | Yes |  | 1496.5256 | 1495.3631 | 1492.1375 | 1487.5425 | 1482.3187 | 1477.7306 | 1474.9153 | Westby | Westby | 282A7DEF-0E50-2980-4E8E5D611C25C8BB |  |  | 1496.5235 | 2.7e-08 | Yes | Overcoupled | 28324DF8-A3C3-8416-509B9289BBE3D131 | 5.41 | 14800000000 | 11600000000 | 1420000000000 | 11.6 | 0.004 | 200000000000000 | 21600000000000 | 11.6 | 4820000000 |  |  | 6.46 | 54.7 | 15300000000 | Other | Severe Q loading in conjunction with FE | Yes | Did not resolve loading | [18 empty] | 28642859-AC43-F2ED-80A775771A2A2E89 | 286676C2-9CEE-6E2F-64B4DD18E0BB3781 | 28682DC9-0D89-AEAF-920FD93AEE1887FC | 286A5FEF-E715-BFBF-FCA34A954625D400 | 286DE9D7-B430-631C-169FAD4528DA4CAC | 2870C66B-AB90-BCA8-21CBC1172D14C84F | 2872F34F-B43D-2218-D5EA014297466CC0 | 1 | HAIPENG | 2012-08-15 18:33 | No | 29
130 | C100-RI-074 |  | Yes | No | 2 | C100_RI_074.txt | 2011-11-22 09:37 | 4 | TS28-004 | WESTBY |  | 6.4e-09 |  |  | 162.5 | 2.07 |  | Yes | Yes |  | 1496.5698 | 1495.2936 | 1491.9827 | 1487.404 | 1482.0415 | 1477.3586 | 1474.0674 | Westby | Bass | 28B6D77F-C337-855D-14906DB83FB56EC9 | Single out of spec at 2940.962987MHz proven to be Quadrapole through manual measurement. | Yes | 1496.5674 | 6.4e-09 | Yes | Overcoupled | 28E067DF-B664-B19B-1B6B4740C496B3E5 | 5.59 | 13600000000 | 10300000000 | 1940000000000 | 11.6 | 0.0004 | 42300000000000 | 30900000000000 | 24.5 | 6080000000 | 8240000000 | 19.3 | 12.53 | 328 | 13400000000 | Quench (non FE) | Quench limited at 24.5 MV/m | Yes |  | 5.4378 |  |  | 19.3 | 12.53 |  | 27 | 27 | 24.75 | 22.13 | 12.52 | 6.5 | [6 empty] | 29419F06-BE5A-5A8D-6A003B6A7A8A52C9 | 2944763B-9042-A5AE-78D417781B0DE117 | 29467627-C7BA-F8B3-BCD9810F92BD3ED3 | 29483367-AFDD-FD18-942AEDCCDE963B9C | 294BAA5F-BE35-A24B-C674DD206D9389C5 | 294DE263-0F20-99C5-839D3194CC906CFC | 294FF69C-95DB-A21D-3F6E0941A334929A | 1 | HAIPENG | 2012-08-15 18:43 | Yes | 29
132 | C100-RI-075 |  | Yes | No | 2 | C100_RI_075.txt | 2011-11-29 09:41 | 8 | TS28-010 | WESTBY |  | 5.06e-09 | Yes |  | 177.4 | 2.07 | 29 | Yes | Yes |  | 1496.4651 | 1495.2269 | 1491.9478 | 1487.2862 | 1482.0206 | 1477.4371 | 1474.0258 | TBASS |  | 5DD070DE-09A2-4EA1-0D7DFF852EFB5799 |  |  | 1496.4644 | 29 | Yes | Overcoupled | 60CACAB2-BE8E-0D52-EEC31F47AE51F1CA | 5.55 | 13300000000 | 9550000000 | 1120000000000 | 11.7 | 0.0066 | 150000000000000 | 20900000000000 | 27.09 | 9370000000 | 10100000000 | 0.00318 | 50 | 0.0115 | 13800000000 | Other | Admin Limit 27 MV/m | No |  | -4.94 | Yes | [16 empty] | 60EB8642-9960-9617-B79C8941932D94A8 | 60EE63FB-D7F3-CAB1-C586C4195D84E085 | 60F0F59F-BC29-6973-4399960905A5A2F5 | 60F4366C-A820-2BA0-0F43D4D28BAE5D49 | 60F66041-C125-46FE-43BF1E3543E3779B | 60F9373B-02C0-41EA-37080912E7AA9EBA | 60FC3B7A-9454-E68A-6A9616E92620ABD5 | 1 | HAIPENG | 2012-08-16 09:42 | Yes
133 | C100-RI-072 |  | Yes | No | 2 | C100_RI_072B.txt | 2011-11-30 13:13 | 4 | 28-007 | WESTBY |  | 1.29e-08 | Yes |  | 163 | 2.07 | 29 | Yes | Yes |  | 1496.5242 | 1495.3629 | 1492.1369 | 1487.5425 | 1482.3185 | 1477.7302 | 1474.9153 | STIRBET |  | FFF377FA-D803-9586-E5837021C0320B4F |  | Yes | 1496.5243 | 29 | Yes | Overcoupled | FFFD5FFC-A871-55D0-EDC9BB82F116169D | 5.95 | 17300000000 | 12000000000 | 1500000000000 | 11.7 | 0.003 | 283000000000000 | 23200000000000 | 27.01 | 12500000000 | 13000000000 | 0.00332 | 50 | 0.03 | 18600000000 | Other | Admin | Yes |  | -4.65 | Yes |  | 0.655 | 15.46 | 29.4 | 27 | 24.75 | 22.13 | 17.37 | 11 |  | Above limits are admin, not quench | [5 empty] | 0036D9FA-FA74-ADB8-0DE97DCD6936DEFB | 00395D5A-F078-5B03-9BE50523C97C1EA6 | 003B69B6-FCEB-5420-85C61F58CFACA0E6 | 003D0507-AD33-3F0B-D6CC30E5A9875FA0 | 003F51B6-011C-2EDD-C97B6F4D1BB23802 | 0041DA4A-CB7C-F9F1-04F0AFD843399D57 | 0044A4B3-99A5-49AA-2A95EC8F0E5DB224 | 1 | HAIPENG | 2012-08-15 18:31 | Yes
134 | C100-RI-076 | Field probe cable TDR shows an open near the top plate. HOM-B cable used for final power rise measurement. | Yes | No | 2 | C100_RI_076.txt | 2011-12-02 14:05 | 8 | TS28-008 | ARI | GRENOBLE | 5.06e-09 | Yes |  | 169 | 2.07 | 29 | Yes | Yes |  | 1496.5762 | 1495.3585 | 1491.9878 | 1487.338 | 1482.031 | 1477.3272 | 1474.2686 | ARI |  | 01D56DC6-B366-60F7-2CAA5B47BE0B210D |  | Yes | 1496.574 | 29 | Yes | Undercoupled | 03508F01-B67B-5E95-E5B957D5D4034A25 | 4.95 | 13400000000 | 11600000000 | 11300000000000 | 11.5 | 0.001 | 17600000000000 | 11500000000000 | 27.02 | 9640000000 | 10200000000 | 0.00101 | 25.01 | 25.1 | 13400000000 | Other | Admin | No |  | -4.66 | Yes | [11 empty] | 10300000000 | 9950000000 | 10600000000 | 11800000000 |  | C2E9A701-95A0-A944-7A3FEBCE31E7A275 | C2EB1E05-0C93-D279-413CD3B8A7F8BAA1 | C2ECBB1E-E984-658B-DF499EE0F127C16A | C2EE0DD6-DE55-BEDB-B36F4B45632A0A47 | C2F0DC8C-D180-349C-3F5E481765ED423D | C2EF5231-B2D8-EEE7-C09F11E6A4873200 | C2F2C053-AF7E-D01B-D65637FCD5740F23 | 1 | HAIPENG | 2012-08-16 09:46 | Yes
136 | C100-RI-079 |  | Yes | No | 2 | C100_RI_079 | 2011-12-10 19:18 | 4 | TS28-007 | KDAVIS |  | 2.4e-08 | Yes |  | 154 | 2.07 | 29 | Yes | Yes |  | 1496.5451 | 1495.325 | 1492.013 | 1487.313 | 1482.19 | 1477.476 | 1474.399 | TBASS |  | A8F364DF-9B46-40C4-1B34659658B54BEB |  |  | 1496.5432 | 29 | Yes | Overcoupled | 9D659F13-98F6-4988-BC471CE187FC1E49 | 5.08 | 14200000000 | 11300000000 | 1040000000000 | 11.6 | 0.004 | 16000000000000 | 3900000000000000 | 20 | 4790000000 | 4790000000 | 192 | 13.3 | 192 | 14400000000 | FE |  | Yes | Power processing in 6/7, 5/7, 4/7, 3/7, and 2/7 modes resulted in little or no decrease in radiation | -5.4 | Yes |  | 387 | 10.6 | 387 | [11 empty] | Cavity failed due to FE loading. Q0 out of spec at 19.5 MV/m. Quenching at 21.5 MV/m. Q0 vs Temp not performed. | 9E1D78C7-E037-1AAC-60B0B5C0F6B11C5D | 9E21B790-927B-6E42-C89B58277DAC4547 | 9E2580E6-CE08-06EA-69542D735E1B2163 | 9E2BD0C1-B4EF-8080-02F5AB659BD99B1E | 9E310809-92DB-B7C8-6E93F34045C1C7F4 | 9E33C246-C1B9-2875-F85318EDC6B9D3EF | 9E3711C2-BAA7-BBD9-EEFBEC151FF525A6 | 1 | HAIPENG | 2012-08-16 10:16 | No
137 | C100-RI-077 |  | Yes | No | 2 | C100_RI_077 | 2011-12-09 16:04 | 8 | TS28-009 | GCIOVATI |  | 6.1e-09 | Yes |  | 161.8 | 2.07 | 29 | Yes | Yes |  | 1496.5744 | 1495.3591 | 1492.0906 | 1487.4084 | 1482.2425 | 1477.678 | 1474.541 | GCIOVATI |  | FD97B072-95BB-765B-ED5978D3E8936098 |  | Yes | 1496.5722 | 29 | Yes | Overcoupled | FD9F8DA5-D356-4244-4487AD355D399C23 | 5.33 | 12600000000 | 11300000000 | 1670000000000 | 11.5 | 0.009 | 42800000000000 | 35700000000000 | 26.92 | 6910000000 | 9030000000 | 0.00524 | 50 | 0.00866 | 13000000000 | Other | Admin limit | Yes | Field emitter turned on, max rad. 2.9 R/h. Tried 20min pulsed-power RF processing and test on passband modes | -5.37 | Yes |  | 30.9 | 13.5 | 2580 | [11 empty] | Cavity was clean after first RF ramp-up to 27 MV/m then emitter turned on. Still achieves 27 MV/m but with high radiation | FDD9BA18-9ABF-E9DE-67B36F078E1880E5 | FDF838EB-EB62-70B3-AE3BA2B940CB98A9 | FDFC8776-FB61-B90D-5EF7B7420FE003A4 | FDFFBCC7-B0D9-224C-70C663C707A08505 | FE02E428-02AF-589E-C22619E91B5FE47E |  | FE05BADC-CCF2-B844-37E052DE78FCEA42 | 1 | HAIPENG | 2012-08-16 09:50 | Yes
138 | C100-RI-081 | NA | Yes | No | 2 | C100_RI_081 | 2011-12-13 17:22 | 8 | TS28-008 | ARI |  | 8.6e-09 | Yes | NA | 180 | 2.09 | 29 | Yes | Yes |  | 1496.56 | 1495.2728 | 1492.0054 | 1487.1578 | 1481.8969 | 1477.1796 | 1473915219 | ARI |  | 3F577012-D86E-0820-E174538C315C8D3C | all good | Yes | 1496.5581 | 29 | Yes | Overcoupled | 3F5F8FCC-F6FE-458A-42AC50C033356058 | 5.4 | 13300000000 | 10600000000 | 1290000000000 | 11.6 | 0 | 10600000000000 | 23400000000000 | 27.1 | 9610000000 | 10300000000 | 0 | 50 | 0 | 13900000000 | Other | admin limit to 27 | No |  | -5.56 | Yes |  | 0 | 0 | 0 | [7 empty] | 10100000000 | 11200000000 | 11800000000 | 13200000000 | perfect | 3F8A5C45-C364-63E2-0D754A75EF7F1E16 | 3F8D1356-B9EC-E6B9-878E36BE99B74B64 | 3F8F53E3-B18D-AC89-8A45620A272BC84D | 3F91604E-BEFC-1B9D-11F8547C7F230761 | 3F938243-CFAD-4821-F5B4391CA0C99479 | 3F9587BF-FE04-6BC5-15DFBF07F5B11FD9 | 3F9CCB34-D9B8-C987-03ACF6D34468E46C | 1 | HAIPENG | 2012-08-16 10:34 | Yes
139 | C100-RI-080 |  | Yes | No | 2 | C100_RI_080 | 2011-12-15 15:04 | 7 | 28-010 | WESTBY | TGOODMAN | 8.5e-09 | Yes |  | 197 | 2.07 | 29 | Yes | Yes |  | 1496.6112 | 1495.3541 | 1492.0261 | 1487.1627 | 1481.8862 | 1477.081 | 1473.8289 | WESTBY | TGOODMAN | A1631386-E17E-8E47-D4DF3EB78016A56F |  | Yes | 1496.6086 | 29 | Yes | Overcoupled | A185AD40-0025-C3B4-BF52010A3222741E | 5.79 | 12800000000 | 8010000000 | 1530000000000 | 11.8 | 0.055 | 302000000000000 | 4410000000000000 | 16.37 | 8070000000 | 0 | 0 | 6 | 342 | 13500000000 | FE | Q loading under 15 MV/m | No |  | -5.13 | Yes | [16 empty] | A1D5348A-9143-6687-CF5B2F1F22DD8E56 | A1D786ED-A07A-7985-6AA6C01C82E6B891 | A1DD743B-F24F-5BFB-DD1690D5EBAD02E1 | A1E10F01-AF9E-4462-6CB2C82C6A46B6C3 | A1E4A353-D4E6-A51F-1E84FFE13C6AE3DB | A1E6CC6C-E23E-995A-2665C73C165AAAC0 | A1E98C88-9396-AA26-6AD6A65AC108B01E | 1 | HAIPENG | 2012-08-16 10:26 | No
140 | C100-RI-082 |  | Yes | No | 2 | C100_RI_082 | 2011-12-16 11:13 | 8 | TS28-001 | WESTBY | TGOODMAN | 8.5e-09 | Yes |  | 185.6 | 2.07 | 29 | Yes | Yes |  | 1496.5254 | 1495.2764 | 1491.9949 | 1487.1445 | 1481.8816 | 1477.0425 | 1473.8526 | STIRBET | WESTBY | CCB90470-FF02-F93D-076D01528BAFF695 |  | Yes | 1496.5244 | 29 | Yes | Overcoupled | CCC32E31-C4E5-FF7F-4DAC69F7EF63FE91 | 5.53 | 12200000000 | 11500000000 | 1110000000000 | 11.6 | 0.001 | 753000000000000 | 217000000000000 | 27.18 | 9180000000 | 9770000000 | 0.0108 | 20.46 | 9.42 | 12500000000 | Other | Admin | No |  | -4.87 | Yes | [16 empty] | CCE8963B-A639-20ED-53D17F2F6F5551D4 | CCEAF422-9243-FF15-C4FA2A34EBCB1A35 | CCED741E-A157-447A-691081C5DB5EA7D2 | CCEFDF4B-E4EF-FA3E-6435E240714774C1 | CCF6D396-D2D7-CA45-9B80A93AED031EF4 | CCF9025F-A4C6-1435-161066119F0EC79F | CCFB3002-DF2C-74C6-D9DD89F1B2F537AA | 1 | HAIPENG | 2012-08-16 10:39 | Yes
141 | C100-RI-079 |  | Yes | No | 2 | C100_RI_079B.txt | 2011-12-16 16:39 | 8 | TS28-007 | GRIGORY |  | 7.75e-08 | Yes |  | 159.3 | 2.07 | 29 | Yes | Yes |  | 1496.5447 | 1495.3253 | 1492.0133 | 1487.313 | 1482.1891 | 1477.4753 | 1474.3984 | TBASS |  | 60CA1E1E-E24E-6F82-8FE6FABD0044347B |  | Yes | 1496.5466 | 29 | Yes | Overcoupled | D848C2B7-EA79-33EB-6A545F9576F30800 | 5.36 | 15700000000 | 11500000000 | 1150000000000 | 11.7 | 0.00141 | 14800000000000 | 7600000000000000 | 27.02 | 10200000000 | 10900000000 | 6.52 | 14.5 | 324 | 15700000000 | Other | Admin limited at 27 MV/m | No |  | -5.28 | Yes | [11 empty] | 10900000000 | 12100000000 | 12800000000 | 14800000000 |  | D88088BE-90C9-60CC-08ED8C559DD2EEB6 | B49E12A6-E2A6-1332-7C70CCED16282624 | D8844438-BC9A-08CD-2A3740636C0D1F3B | D8897AE3-E351-5130-B3E25221552FA0D0 | D88C3CF6-DE79-9F6C-BE58587E79A825BB | D8901133-FD82-D3CF-E921A479C744F214 | D8920E96-9202-B3B8-ED20407D9B1A65E4 | 1 | HAIPENG | 2012-08-16 10:13 | Yes
142 | C100-RI-078 |  | Yes | No | 2 |  | 2011-12-17 | 7 | TS28-009 | STIRBET |  | 5.8e-08 | Yes | Vacuum instabilities during cavity tests (During HOM qualification went up and down from 6.3 e-8 to 5.5 e-8mbar). | 190.87 | 2.07 | 29 | Yes | Yes | Performed cable calibration twice... | 1496.4875 | 1495.2342 | 1491.979 | 1487.265 | 1482.0011 | 1477.3128 | 1474.0687 | STIRBET |  | 6330FC4B-00F4-0D02-DB1FD55D7E8EEA87 |  |  | 1496.4855 | 29 | Yes | Overcoupled |  | 6.11 | 1.59 | 24800000000 | 9030000000000 | 11.9 | 0.0015 | 9050000000000 | 1820000000000000 | 9.5 | 4190000000 |  |  | 9 | 118 | 16600000000 | FE | Strong FE and qo drop at low cavity fields. | No |  | -5.261 | Yes | [16 empty] | 63E2BFC7-FF2A-32B0-36AAED8B28D02842 | 63E52622-CAF3-BEEC-8DC8DFA5D13A2AC6 | 63E7D217-B7D1-7698-6806D8B45107139E | 63EA0875-AD9F-60C1-585E506AEA810FED | 63EC38D7-9ADB-133D-CCD6ED75F1C74E6F | 63EEA109-C58C-7EC6-1642712B5B30F2A0 | 63F109F8-0723-2F03-0EC1EEC78D284DA5 | 1 | HAIPENG | 2012-08-16 10:05 | No
143 | C100-RI-080 | cavity was tested on 15-Dec-2011 (VTRF # 139) and was FE loaded @ 16MV/m | Yes | No | 2 | C100_RI_080A.txt | 2011-12-19 16:58 | 4 | TS28-004 | WESTBY | TBASS | 1.17e-08 | Yes |  | 168 | 2.07 | 29 | Yes | Yes |  | 1496.6139 | 1495.357 | 1492.0289 | 1487.1659 | 1481.8893 | 1477.0846 | 1473.832 | WESTBY | TGOODMAN | A9FDF956-A878-09C0-8F328C744170F3AF | HOM survey performed on 15-Dec-2011 (VTRF # 139) |  | 1496.6117 | 29 | Yes | Overcoupled | 97A21E49-EDA3-EA3C-F05F037836230C38 | 5.95 | 16800000000 | 8230000000 | 928000000000 | 12.4 | 0.001 | 214000000000000 | 21000000000000000 | 27 | 10200000000 | 11100000000 | 369 | 7.48 | 875 | 18000000000 | Other | Admin Limited @ 27MV/m | Yes | Initial cavity run had a few cavity quenches. Clean Q0 vs Eacc data curve. | -4.78 | Yes |  | 5.94 | 16.1 | 882 | [12 empty] | 76385D7B-D17E-2147-84D16B4B3F7E3EFE | 763A717F-B457-874B-B23016A95AF2F8A0 | 763C4283-082D-10CD-87ACE77C98413D8C | 763E2F95-EE60-8FC4-AAE2BCA93DD2BE22 | 763FCE8A-0863-C13A-D0B263DA523E0B59 | 7641AF77-DFB0-3A09-5EB460F3A8682FD6 | 76436807-D5F0-40AB-D968A071445F1D7D | 1 | HAIPENG | 2012-08-16 10:22 | Yes
144 | C100-RI-083 |  | Yes | No | 2 | C100_RI_083.txt | 2012-01-11 10:34 | 8 | TS28-008 | GRENOBLE |  | 5.06e-09 | Yes |  | 152 | 2.07 | 29 | Yes | Yes |  | 1496.5268 | 1495.2435 | 1491.9127 | 1487.204 | 1481.9842 | 1477.1406 | 1473.8989 | TBASS |  | 0651883E-05C7-5CE0-0CBEF15C049F8AD7 |  | Yes | 1496.5263 | 29 | Yes | Overcoupled | 07802B99-B099-6A33-DC363A76BF322B4F | 5.72 | 14600000000 | 13400000000 | 1700000000000 | 11.6 | 0.006 | 14400000000000 | 138000000000000 | 26.92 | 10200000000 | 10200000000 | 0.001 | 21.9 | 7.37 | 14900000000 | Other | administration limit at 27MV/m | No | Cavity had no FE. | 5.04 | Yes | [15 empty] | Cavity passed at administrative limit of 27MV/m. No FE greater than 7.37mR/hr Unable to perform Q-vs-T. Pumping on dewar caused pressure problems with cryo. | 0DE680A8-DAE0-F287-8AFDED93872CE9D1 | 0DE9626C-BFEC-387D-A437A9F575AD5053 | 0DEBEA77-B194-3B50-3A0F6D0CC821A435 | 0DEEF1A8-FF3C-A77D-EE9E43ECE9B58E2C | 0DF1F26B-9DE8-D777-EAE80248B26DE512 | 0DF46CED-C4E9-29A7-C64014496700FA83 | 0DF73110-FEF4-2B33-C0DD2CA8DE382249 | 1 | HAIPENG | 2012-01-11 14:30 | Yes
145 | C100-RI-078 |  | Yes | No | 2 | C100_RI_078B | 2012-01-11 14:04 | 4 | TS28-004 | TGOODMAN | KDAVIS | 1.52e-08 | Yes |  | 175.6 | 2.07 | 29 | Yes | Yes |  | 1496.4893 | 1495.236 | 1491.9806 | 1487.2184 | 1482.0128 | 1477.3149 | 1474.0706 | STIRBET | STIRBET | A9A98042-EE6F-94BD-6D010DFEE4C2DBE8 | HOM survey for C100-RI-078 was done on 17 Dec 2011 |  | 1496.4869 | 29 | Yes | Overcoupled | 36DC9FBA-A278-D2CC-D15F14DA515DD099 | 5.74 | 14900000000 | 11300000000 | 942000000000 | 11.6 | 0.00144 | 15000000000000 | 66000000000000 | 27.14 | 9330000000 | 9370000000 | 84.5 | 17.22 | 573 | 15900000000 | Other | Admin limit 27 MV/m | No |  | -5.05 | Yes | [11 empty] | 9490000000 | 10700000000 | 11700000000 |  |  | 374AB3DE-F8DC-D3C8-0B61221E94C07DEB | 3761A957-CF8D-8E65-D2891B82C22EA18A | 37686A01-C99D-2A41-D2358BC1B0D4B652 | 376DA335-9997-79B2-C94AF3AA5FF1E1C6 | 37713556-B72D-6A7D-000A1B0F2F9AD447 | 3775685A-B29E-0070-57500916F73941AF | 37848342-00C9-C42F-EDACD82983AF0874 | 1 | HAIPENG | 2012-08-16 09:57 | Yes
146 | C100-RI-001 |  | No | No | 2 | C100_RI_001F.txt | 2010-06-07 11:03 | 8 | TS28-007 | GRENOBLE | KDAVIS |  |  | cavity vacuum not monitored | 170.5 | 2.07 | 29 | Yes | Yes |  | 1496.8493 | 1495.6338 | 1492.3839 | 1487.6461 | 1482.6795 | 1477.9729 | 1474.8173 | GRENOBLE |  | 3AD9C074-020A-1CAD-BD86FC85B58E431E | HOMs measured manually. | Yes | 1496.8437 | 29 | No | Overcoupled | 3AFF1B26-0B92-8927-A18C655CD8CA9A80 | 3.19 | 13500000000 | 5120000000 | 1000000000000 | 16 | 0.1 |  |  | 37.32 | 9070000000 | 14700000000 | 0.00373 |  | 0.0084 | 16100000000 | RF power |  | No |  | 5.55 | Yes | [11 empty] | 8900000000 | 9600000000 | 10000000000 | 10800000000 |  | 3C8A3DB0-AC6A-9D7E-4B335C838805E792 | 3C8EDAB4-B777-C1AA-2D58FE5C77156EB5 | 3C915BF2-C3DB-A463-99B5B13198C51C2F | 3C95771E-9220-FBFA-1B34CF5595572E71 | 419BD973-F773-8F79-A56C9D405E51AC49 | 3C9B41EA-DFBE-5E65-D1F6A5903C78FE66 | 3C9E0D4B-E2AC-830D-763219C347835077 | 1 | HAIPENG | 2012-01-12 15:03 | Yes
147 | C100-RI-084 |  | Yes | No | 2 | C100_RI_084 | 2012-01-12 11:27 | 3 | 28-007 | WESTBY |  | 7.6e-09 | Yes |  | 156.56 | 2.07 | 29 | Yes | Yes |  | 1496.5717 | 1495.322 | 1492.0768 | 1487.2441 | 1481.9497 | 1477.3218 | 1473.9354 | STIRBET |  | 6F186EC1-0957-6610-89A8629E4CAC6773 |  |  | 1496.5709 | 29 | Yes | Overcoupled | 6F3B50AB-F324-6459-DAFA6C174EBB6322 | 5.39 | 15300000000 | 9790000000 | 1140000000000 | 11.8 | 0.001 | 68900000000000 | 34800000000000 | 10.5 | 10900000000 |  |  | 9 | 136 | 16200000000 | FE |  | No |  | -5.1 | Yes | [15 empty] | FE Loaded @ 10.5MV/m, test aborted | 4058830C-AF43-E336-DB0B02AACF9D7780 | 453F6FC0-A250-9A3E-5DEE05BC9B99D910 | 405CB9F9-FE82-79C1-F4A370867FB417B3 | 405DC1B5-F445-557C-21261EC1D8B7AA02 | 405EE134-F1C7-3C90-2D366E73A03234FF | 406083F7-C9CC-2A4B-2656381CC94C006D | 40634048-9542-7C2D-9A2693AE9AF4E76F | 1 | HAIPENG | 2012-01-17 16:20 | No
148 | C100-RI-084 |  | Yes | No | 2 | C100_RI_084A | 2012-01-19 14:58 | 4 | 28-007 | WESTBY |  | 7.7e-08 | Yes |  | 177.8 | 2.07 | 29 | Yes | Yes |  | 1496.5717 | 1495.0768 | 1492.0768 | 1487.2441 | 1481.9497 | 1477.3218 | 1473.9354 | STIRBET |  | ABA883D3-C2BA-A9F5-FA3FAFCFD476D345 |  | Yes | 1496.5694 | 29 | Yes | Overcoupled | ABB05A82-FEFE-020E-F3FDFA0A333B15A1 | 5.17 | 15000000000 | 9270000000 | 1080000000000 | 11.9 | 0.003 | 36200000000000 | 39300000000000 | 27.45 | 10500000000 | 11100000000 | 3.69 | 12.48 | 65.8 | 14500000000 | Other | Admin limited at 27 MV/m.Added by Mircea. | Yes |  | -5.47 | Yes |  | 0.245 | 17.95 | 67 | [11 empty] | During previous VTA qualification attempt done on C100-RI-084 on 12 Jan 2012 the cavity has Q drop at 10 MV/m and Rad max ~150 mR/h. After warming up the cavity has been recovered by HWP rinsing. Added by Mircea. | ABE43772-E7E4-6922-2B848F3C2873464E | ABE6DB4C-CF67-3356-218FD8167DD3CE83 | ABE92836-CA7F-0C20-2E17FDD0912ABB6A | ABEB3243-F2B3-EBC5-07E3A0C3D3FEE10A | ABED31CA-E75B-8F00-640F06298D804CF0 | ABEF88DB-D167-3F3B-CEC2EA71790CA383 | ABF2A98F-B84A-DF92-925C049ED4AAF858 | 1 | HAIPENG | 2012-08-16 10:50 | Yes
149 | C100-RI-085 |  | Yes | No | 2 | C100_RI_085.txt | 2012-01-28 | 7 |  | KDAVIS |  | 6.5e-09 | Yes |  | 200 | 2.07 | 29 | Yes | Yes |  | 1496.483 | 1495.282 | 1492.225 | 1487.65 | 1482.681 | 1478.202 | 1475.132 | STIRBET |  | D9287DD3-9068-568D-059E40A0BB94914F |  |  | 1496.4814 | 29 | Yes | Overcoupled | 75362256-FBC1-AA01-ECBA849726E0C1BD | 5.53 | 12400000000 | 11600000000 | 1380000000000 | 11.7 | 0.005 | 11000000000000 | 15000000000000 | 27 | 8720000000 | 8820000000 | 0.01 | 19 | 0.01 | 12900000000 | Other | Admin limit of 27 MV/m | No |  | -5.12 | Yes |  | 0.01 | 19 | 0.01 | [7 empty] | 8820000000 | 10000000000 | 10700000000 |  | beginning at 19 MV/m, FE was just above background and then cleaned up as gradient was increased in 0.5 MV/m steps | 773692B8-026A-40C3-C4A73DA32D712C4C | 77459F26-0BB1-C8AC-3890A91859AA085D | 774F6214-CBCD-F97A-D04CF97DF3811CBC | 7751E10E-B49A-9795-608740EF36FBC0D7 | 77547C3D-0A1B-F7BA-866D5F4913CCD9C4 | 7756F7AC-B559-3B31-023B2BEE9EBE06E8 | 77598E73-9ECF-0DA6-A989FE4E922C6B9C | 1 | HAIPENG | 2012-08-16 10:58 | Yes
150 | C100-RI-062 |  | Yes | No | 2 | C100_RI_062.txt | 2012-01-31 | 8 | 28-010 | STIRBET | KDAVIS | 5.5e-09 | Yes |  | 169.3 | 2.07 | 29 | Yes | Yes |  | 1496.5075 | 1495.2506 | 1491.9774 | 1487.325 | 1482.1012 | 1477.5188 | 1474.3961 | STIRBET |  | 41983A96-E603-57A0-42E3B524919E4728 |  |  | 1496.5055 | 29 | Yes | Undercoupled | 48C7449B-EABE-4299-C0B4B5C7CE80B277 | 5.32 | 10700000000 | 13000000000 | 1180000000000 | 11.5 | 0.001 | 19300000000000 | 1010000000000000 | 26.18 | 7640000000 | 8260000000 | 9.54 | 14.99 | 10000 | 11000000000 | Operator | At 25.5 MV/m lost cavity, after re-locking cavity the radiation measured was in excess of 1 R/h at cavity gradient higher than 25.5 MV/m. At 26.2 MV/m radiation was over 10 R/m. C100-RI-062 will be warmed up , rinsed then VTA retested. | No |  | -5.2 | Yes | [11 empty] | 8050000000 | 8600000000 | 8930000000 | 9710000000 | The cavity C100-RI-062 had Qo near to the qualification limit from 15 M/m up to 26 MV/m. Dewar 8 related? Perforemd Qo vs temperature measurements at two supplementary pressures: 16 Torr and 13 Torr. Enter the Qo values for each temperature in the yellow boxes below: Pressure (Torr) Temp (K) Qo 1/Qo Calc 29 2.07 8.05E+09 1.24E-10 25 2.03 8.60E+09 1.16E-10 23 1.99 8.93E+09 1.12E-10 19 1.93 9.71E+09 1.03E-10 16 1.88E+00 1.03E+10 9.68E-11 13 1.82E+00 1.11E+10 9.05E-11 | 42137869-DF0B-406F-FC7F41928F4291FB | 4939C10B-F226-B46D-E528B4D4FC51C68B | 4224A2FF-EEE3-068F-AA12258B0D1CDAEB | 42278A4B-BDC0-742A-7B9820D6045FEF98 | 42295FBA-ACAB-1A8A-E6D217D2D1F370AD | 422B136B-B356-4711-0B44B69813C31107 | 422CD5B8-D9F8-C0EE-74817B40365C2691 | 1 | HAIPENG | 2012-08-15 16:59 | Yes
151 | C100-RI-009 | Cavity originally had physical repairs to HOM ports early in C100 run. | No | No | 2 | C100_RI_009 | 2012-02-03 09:03 | 8 | TS28-009 | WESTBY |  | 8.5e-09 | Yes | Cavity Vacuum did increase during HP RF, remained within specifications. | 185.5 | 2.07 | 29 | Yes | Yes |  | 1496.7142 | 1495.5712 | 1492.4075 | 1487.921 | 1482.8937 | 1478.4256 | 1475.2881 | STIRBET |  | 71779F3F-FC85-8467-1223C4585E914FDB |  | Yes | 1496.7138 | 29 | No | Overcoupled | 7181A68E-B9B7-06CA-3D80A3567E7CB070 | 4.67 | 11400000000 | 7030000000 | 1180000000000 | 11.9 | 0.004 | 7700000000000 | 35500000000000 | 33.12 | 6240000000 | 9090000000 | 0.00351 | 22.05 | 151 | 11500000000 | Quench (non FE) |  | No |  | -5.69 | [12 empty] | 79100000000 | 8600000000 |  |  | Admin limit of 27MV/m exceeded with Persmission of John Hogan | 71AA518D-A56C-E3C8-41AAA211DF931FCC | 71AD9D5B-A804-CCD0-B391B8D855B7248C | 71AFF8DA-E913-FA83-2FEE14194BED595C | 71B25297-F5CA-5DF6-874855BA926D3E8C | 71B56076-C263-4FFF-7106895A384FF3FC | 71B7CFE2-DDE3-4F0A-0AE98C2DF9BEFBB6 | 71BA5298-9EE2-6E41-05D59F8E1DF8221D | 1 | HAIPENG | 2012-02-24 13:51 | Yes
152 | C100-RI-062 |  | Yes | No | 2 | C100_RI_062A.txt | 2012-02-04 | 7 | TS28-010 | KDAVIS |  | 5.9e-08 | Yes |  | 177 | 2.07 | 29 | Yes | Yes | FP cable flaky. After initial RF qualification test done on 4 Feb,on 5 Feb perform retest using HOMA as FP. | 1496.511 | 1495.255 | 1491.981 | 1487.329 | 1482.106 | 1477.523 | 1474.4 | STIRBET |  | A3802944-FC59-C424-BD7B7F5305E741A3 | HOM survey done on 31 Jan 2012 |  | 1496.5092 | 29 | Yes | Undercoupled |  | 6.57 | 15200000000 | 16800000000 | 12000000000000 | 11.5 | 0.003 | 12000000000000 | 430000000000000 | 27.05 | 10800000000 | 12000000000 | 0.0717 | 17.49 | 67.5 | 16500000000 | Operator | Administrative limit 27 MV/m | No |  | -4.87 | Yes | [11 empty] | 11200000000 | 13000000000 | 14600000000 |  | Administrative limit at 27 MV/m, low radiation levels during VTA qualification tests. | A3E4E5D7-DB39-F04C-88BB446B6ED89B53 | A8044D9B-EE94-D170-8FBEAD62D5A2EA0F | A418EB14-B98E-3929-6D58930174F5F3EB | A41DB3CF-C84D-E17B-C2F5AAC298F29D99 | A42012A9-FC53-9166-9328437C485C4258 | A4226C30-F3ED-E1C4-8BD0C0E7D41A7DBC | A424A62C-06ED-6D01-0074FE29A422C82F | 1 | HAIPENG | 2012-08-15 16:56 | Yes
153 | C100-RI-086 |  | No | No | 2 | C100_RI_086.txt | 2012-02-10 | 4 | TD28-003 | STIRBET |  | 1e-09 | Yes |  | 175.91 | 2.07 | 29 | Yes | Yes | Questionable FPC cable under dear lid. | 1496.7079 | 1495.5438 | 1492.4026 | 1487.901 | 1482.9101 | 1478.4292 | 1475.3031 | STIRBET |  | A5D7EE38-F52D-0E5D-0E4632FE1F79F820 |  |  | 1496.7055 | 29 | No | Overcoupled |  | 5.53 | 13400000000 | 9050000000 | 1450000000000 | 11.7 | 0.008 | 612000000000000 | 33100000000000 | 30.59 | 7500000000 | 9540000000 | 0.00214 | 23.06 | 1.6 | 13900000000 | Quench (non FE) |  | No |  | -5.32 | Yes |  | 0.00214 | 23.06 | 1.6 | [7 empty] | 9570000000 | 10500000000 | 11000000000 |  | C100-RI-086 has been tested twice: the first test done on 9-Feb-2012 had high values for Qo and was quench limited at 31.5 MV/m. Radiation onset at 23 - 24 MV/m, low radiation yield up to 31.5 MV/m. During atempt to measure temperature dependence of Qo at 2.07K a sudden drop in Qo. Repeat VTA qualification on 10-Feb-2012, cavity had normal values for all parameters, displlayed again low radiation and was quench limited at 30.5 MV/m. | A666D6EC-E97E-0D7A-A7D10E1219D222C0 | A7F077ED-EA2E-927B-0EADA8374DC987C1 | A66C70A3-D1F0-EB25-9FE1D233D27FC409 | A66E29EC-A903-D6E3-859E6D6DAC5DC1E9 | A6701B28-CF0A-DE2D-A2C615E6DFB30387 | A672B289-ADF3-07B7-9A44ABFC5216D60B | A6747E30-E5E5-E3CF-FCF30A1453963B4E | 1 | HAIPENG | 2012-08-16 11:04 | Yes
157 | C100-RI-086 | Plasma processed | Yes | No | 2 | C100_086_201217 | 2020-12-17 13:58 | 4 | TS-28-006 | GRENOBLE |  |  |  | Vacuum readout showed under range. | 157 | 2.1 | 29 | Yes | Yes |  | 1496.0319 | 1494.7358 | 1491.6636 | 1487.2474 | 1482.0152 | 1477.5029 | 1474.5748 | POWEN | GRENOBLE | A61FC459-E213-AE7C-DCD52E7D5E964AB7 |  | Yes | 1496.0174 | 29 | No | Overcoupled |  | 3.7 | 14000000000 | 10700000000 | 1830000000000 | 11.39 | 0.0038 | 1830000000000 | 21500000000000 | 9.99 | 2420000000 |  |  | 5.69 | 771 | 13400000000 | FE |  | Yes | Processed during the first Q vs E curve measurement. Noticeable improvement in Q roll off after initial quenches and radiation bursts. | -0.47 | Yes | Value above measured before Q roll off and radiation. | [15 empty] | A460BB2A-AEB2-41EA-6DB4C54C985C674B | A467A8D7-0771-CB46-2EF9DDAA05190B3B | A469FF9E-B098-2606-BA57C3CD183CE3E8 |  | A4701687-91A3-83ED-DFE97584761D1089 | A4721A28-0B5C-3266-DE4B32CFB87A16DB | A474A7A2-B198-7E01-4C2691915ADAADBE |  |  |  | No
158 | C100-RI-086 | Plasma treatment | No | No | 2 |  | 2021-01-21 19:03 | 4 | TS-28-006 | GRENOBLE |  |  | Yes | Vacuum gauge displayed 'ur' | 162 | 2.1 | 29 | Yes | Yes |  | 1496.022 | 1494.747 | 1491.744 | 1487.187 | 1482.015 | 1477.499 | 1474.55 |  |  |  | HOM measurement was not required in test plan. |  | 1496.2 | 29 | No | Undercoupled |  | 2.54 | 7.559 | 9600000000 | 1300000000000 | 11.6 | 0.004 | 19000000000000 | 6300000000000 | 9.6 | 2510000000 |  |  | 5.96 | 928 | 7300000000 | FE | FE limit at 9.6MV/m. |  |  | -3.3 | No | [16 empty] | B1420A86-B1A7-8378-66E713824F0E4EFC | B1452B93-0385-96A0-A698DC11DA3F89EB | B1477AD7-DCC2-5E10-1E767ECD0F472CB2 | B1490143-A0DE-4618-1886CFCF01CCFBCC | B14AC371-08D6-B1B9-B197EFBFF98028CC | B14CA9BF-D5A2-BC80-D970D53CFF546CD0 |  |  |  |  | No
159 | C100-RI-086 | Post plasma processing round 2, retest with fast cooldown | Yes | No | 0 | C100-RI-086_210210A | 2021-02-10 12:45 | 4 | TS28-006 | POWEN |  | 1e-09 | Yes |  | 161 | 2.07 | 29 | Yes | Yes | [13 empty] | 1496.0171 | 29 | Yes | Overcoupled |  | 4.8746 | 14.2913 | 10100000000 | 13.6086 | 11.4 | 0.001 |  |  | 9.39 | 3530000000 |  |  | 5 | 786 | 14000000000 | Operator | FE Onset at ~ 5MV/m, Eacc ~ 9.5MV/m with Rad approaching 1E3 mR/h. | No | [18 empty] | FE Onset at ~ 5MV/m, similar to recent tests. Eacc ~ 9.5MV/m with Rad approaching 1E3 mR/h. Qo on this test with a fast cool down was improved over the previous test. | 87405365-0CBD-E0BF-C6A28C9DDAC6E4C9