JLEIC Crab Cavity Trim Tuning Measurements and Cable Calibration Procedure

- Purpose and Scope

This procedure describes the steps to measure the frequency, and Qext of the Input Probe and Pick Up Probe for the JLEIC crab cavity.

- References

Coupling-cal – Spreadsheet to calculate Qext

JL0127752 – JLEIC Crab Cavity Frequency Measurement Drawing

- Process Details

  - Trim Tuning

Fig. 1: Trim Tuning Set Up

Assembly Steps:

Fig. 2: Top and bottom plates with opening to insert auxiliary probes.

- Assemble the 4 rods to the bottom plate. Make sure the plate is leveled (Fig. 2).

- Assemble a 6” blank flange with a crushed 6” Cu gasket to the end cap at the bottom.

- Stack the 3 sub-assemblies (Fig. 1).

- Temporary blanks, auxiliary probes will supply RF power to the cavity while measuring RF frequency of the stacked assembly (Fig. 3).

Fig. 3: Temporary blanks with auxiliary probes.

- The threaded holes for the feed-throughs are slightly off center to couple to the fundamental mode through the beam pipe. The temporary antenna will be cut to a length that will ensure that it is under coupled and not affecting the field in the cavity (approximately 5”-6” in length).

- Insert the temporary blank that fits on a 6” CF flange with the auxiliary probe (with an offset of 1 cm as shown in the diagram) and a 6” Cu gasket. Clamp the temporary blank to the CF flange.

- Place the top plate and tighten the set up.

- Use the temporary blank that fits a 2¾” CF flange with the feed-through. Use a temporary antenna of approximately 2”-3” in length. Insert the temporary blank with auxiliary probe centered on one of the 2¾” flanges. Clamp the temporary blank to the CF flange.

Frequency Measurement Steps:

- Connect Vector Network Analyzer (VNA) and attach cables to the input (Port 1) and pick up (Port 2) probes.

- Set the VNA for an S21 measurement, Format = Log Mag, Number of points = 1600, IF Bandwidth = 1 kHz.

- Set the Start Frequency to 950 MHz and Stop Frequency to 960 MHz. There should be a single peak displayed corresponding to each mode to the fundamental crabbing mode.

- Set 3dB BW S21 for measurement. Make sure QL is > 6500 (Fig. 4). Tighten the assembly to determine the QL doesn’t vary. This verifies the good electrical contact between the sub-assemblies in the cavity.

Fig. 4: VNA screenshot to measure QL.

- Set the VNA for an S11 and S22 measurement, Format = Polar  Real Imag.

- Check that S11 and S22 to ensure it is critically coupled and not over coupled (Fig. 5).

Fig. 5: VNA screenshots to check the coupling.

- Record the frequency of the mode.

- Follow the above steps to measure the cavity frequency at different processing steps on the cavity.

  - Probe Calibration

Fig. 6: Input and Pick Up Probes for VTA Test.

The JLEIC crab cavity test in VTA requires an Input Probe to couple forward power in to the cavity and a Pick Up Probe to measure the transmitted power (Fig. 6).

Input Probe Qext Measurement and Tuning Steps:

- Setup the VNA as follows:

  - Measure: S21

  - Trigger: Continuous

  - Format: Log Mag

  - Span: 1 MHz

  - Number of Points: 1601

  - IF BW: 300 Hz

  - Set the Center Frequency to be the frequency of the crabbing mode (952 MHz)

- Calibrate the VNA: connect the cables to the Keysight N4431B electronic calibration module. Select Cal, E-CAL.

- Verify the calibration using an S21 measurement with the cables connected together: the attenuation should be ≤ ±0.01dB when measuring calibrated cables.

- A launching antenna should be inserted in a feedthrough with a temporary blank installed on the Pick Up Probe flange (Fig. 6). Connect Port 1 of the VNA to the feed-through.

- Attach the shorter Cu probe (Fig. 7) to the 2¾” CF flange with the threaded feed-through (Fig. 8). Connect the flange to the cavity Input Probe flange with a crushed 2¾” Cu gasket. Connect Port 2 of the VNA to the feed-through.

Fig. 7: Cu probes for Pick Up Probe (left) and Input Probe (right).

Fig. 8: 2¾” CF flange for Input Probe

- Set the VNA to Measure S11, Format: Polar, to check the input antenna for coupling and reflection coefficient. The VNA screen should be similar to the one shown in Fig. 9: the signal will not include the origin if the antenna is properly under-coupled. Too short of an antenna will result in noisy signal.

Fig. 9: VNA screenshots to check the coupling.

Fig. 10: VNA screenshot to measure the reflection coefficient of the launching antenna.

- Set marker to search for minimum. Using the Scale: Phase Offset button, roll the signal around until the point of the marker is on the positive x-axis.

- Change the Format to Real, use the Reference Value button to roll the “off-resonance” (detuned) signal to the reference line (Fig. 10). Record the values of “S11_Detuned” and “S11_Resonant” in mU by placing the marker off resonance and on resonance (Use the Mrk Search: Min function), respectively. The length of the launching antenna should be such that S11_Resonant is ~ 0.8-0.9 mU.

The reflection coefficient can be calculated as: Γ = S11_Resonant/ S11_Detuned.

- Change the measurement to S21, Format: Log Mag. Set Mrk Search to Max, Tracking: On, set Width: On (default setting for marker widths is -3 dB), as shown in Fig. 11. Record the values of “QL” and “loss” displayed on the screen. The Qext can be calculated as:

The measured quantities can be typed in the Qext calculation spreadsheet to automatically calculate the Qext. The Qext of the Input Probe needs to be set at 9.0×109 ± 15% (7.7×109 – 10.4×109).

Fig. 11: VNA screenshot of the Input Probe insertion loss measurement.

- If the Qext lower than the range mentioned above, the Input Probe is strongly coupled. Disassemble and disconnect the Cu probe from the 2¾” CF flange.

- Trim the Cu probe by 1-3 mm. And repeat the steps 5-9 until the desired Qext is achieved.

Pick Up Probe Qext Measurement and Tuning Steps:

- Repeat steps 1-3 given above.

- A launching antenna should be inserted in a feedthrough with a temporary blank installed on the Input Probe flange (Fig. 6). Connect Port 1 of the VNA to the feed-through.

- Attach the longer Cu probe (Fig. 7) to the 6” CF flange with the threaded feed-through that is designed with an offset 1 cm (Fig. 12). Connect the flange to the cavity Input Probe flange with a crushed 6” Cu gasket. Connect Port 2 of the VNA to the feed-through.

Fig. 12: 6” CF flange for Pick Up Probe.

- The Qext of the Input Probe needs to be set at 5.0×1011 ± 15% (4.3×1011 – 5.8×1011).

- Repeat steps 6-11 given above.