|  |  |
| --- | --- |
| Traveler Title | L2HE Nine Cell Cavity Transfer to Test Stand for Dewar 5 |
| Traveler Abstract | This traveler verifies proper transfer of L2HE 9-cell cavities to test stands for dewar 5 in preparation for VTA testing. |
| Traveler ID | L2HE-CLNRM-CAV-TSTD5 |
| Traveler Revision | R3 |
| Traveler Author | A. Grabowski |
| Traveler Date | 19-Dec-22 |
| NCR Informative Emails | weinmann,cheng,forehand |
| NCR Dispositioners | hannesv,adamg |
| D3 Emails | weinmann,cheng,forehand,adamg |
| Approval Names | A. Grabowski | C. Dreyfuss | J. Vennekate | M. Bevins |
| Approval Signatures |  |  |  |  |
| Approval Dates |  |  |  |  |
| Approval Title | Author | Reviewer | Reviewer | Project Manager |

|  |  |
| --- | --- |
| References | List and Hyperlink all documents related to this traveler. This includes, but is not limited to: safety (THAs, SOPs, etc), drawings, procedures, and facility related documents. |
| L2PRO Cavity Transfer to Test Stand Procedure (link not available) |  |  |  |  |
|  |  |  |  |  |

|  |  |
| --- | --- |
| Revision Note |  |
| R1 | Initial release of this Traveler. |
| R2 | Revise step 3 & 4 to leave RAV closed for leak check and test on new from vender cavities |
| R3 | Revised RGA scan from 2 to 65 to 2 to 100 |

|  |  |  |
| --- | --- | --- |
| Step No. | Instructions | Data Input |
| 1 | At this point three cavities are ready for transfer to a dewar 5 test stand in the VAA. If a cavity has arrived from the vendor, it had incoming RF and dimensional inspections, has been wiped down by the chemistry staff and passed into the clean room on a CTV. Otherwise, it had HPR and assembly.Hang the first cavity in the dewar 5 test stand using the two-arm lifting fixture and secure the plates to the square blocks attached to the helium vessel.\*\*Repeat this two more times for the other cavities. \*\*The test stand will need to be rotated in the hole to accommodate the installation of the third cavity.* Verify that the cavity isolation valves are in the fully closed position.
* Remove four bolts from the bellows blank on one of the test stand down-tubes.
* Spray the flange and holes with ionized nitrogen while watching a particle counter. Wait until the counts are below fifty at all sizes for a minimum of ten seconds.
* Repeat this procedure for the blank flange located on the right angle cavity isolation valve.
* Remove the remaining two bolts on the valve blank and spray the inside of the right angle valve (bellows area) until the counts are less than ten at all sizes for ten seconds.
* Remove the remaining two bolts from the test stand bellows blank on the down-tube.
	+ DO NOT spray the bellows flange after the blank is removed.
* Place a clean copper gasket onto the joint and connect the test stand flange to the right angle cavity isolation valve. Tighten all fasteners.
* Repeat all these steps for the remaining cavities.
 | [[CAVSNPosition1]] <<CAVSN>>[[CAVSNPosition2]] <<CAVSN>>[[CAVSNPosition3]] <<CAVSN>>[[CAVSN]] <<TEXT>>[[CAVSN = CAVSNPosition1 + CAVSNPosition2 + CAVSNPosition3]] <<NOTE>>[[VTATSN]] <<VTATSSN>>[[Technician1]] <<SRFCVP>>[[Technician2]] <<SRFCVP>>[[CAVPosition1VLVClosed]] <<YESNO>>[[CAVPosition2VLVClosed]] <<YESNO>>[[CAVPosition3VLVClosed]] <<YESNO>>[[CavComment]] <<COMMENT>> |
| 2 | Open all the cavity isolation valves on the top-plate of the test stand. If the cavity is not already under vacuum from the vendor, start the pumping system on the test stand top plate. Slow pump-down is not necessary at this point if this cavity is already under vacuum from the vendor. | [[IsolationValve1Open]] <<YESNO>>[[IsolationValve2Open]] <<YESNO>>[[IsolationValve3Open]] <<YESNO>>[[IsoValveComment]] <<COMMENT>> |
| 3 | Test stand flange to cavity isolation valve joints are ready to be leak checked after the vacuum is below 8e-7 mbar. Set the RGA for an analog scan looking at mass 2 thru 65. Leak check this joint while the valve is still closed using the RGA system attached to the test stand. There shall be no detectable leak on this joint with an MDL less than 4E-10 std cc/sec. Upload lk check file.Isolate the turbo-pump from the test stand and slowly open only cavity isolation valves that do not have red tape indicating they are new from vedor cavities, while observing test stand pressure. Record the highest pressure reached with all reworked cavities opened and no active pumping on stand.Carefully open the turbo-pump isolation valve to allow for active pumping on all reworked cavities. | [[StandLkChkTechnician]] <<SRFCVP>>[[ObservedCavitiesPressure]] <<SCINOT>>mbar[[StandLkChkComment]] <<COMMENT>>[[StandLkChkDate]] <<TIMESTAMP>> |
| 4 | Only for rework cavities, i.e. cavities WITHOUT tape on they right angle valve (RAV): Isolate the turbo-pump from the test stand and slowly open the RAV on the reworked cavities. Perform a slow pump down until the system is below 8e-7 mbar. Set the RGA for an analog scan looking at mass 2 thru 100. Upload scan. Leak check all mechanical joints on cavities with open RAV using the RGA system attached to the test stand. There shall be no detectable leak with an MDL less than 4E-10 std cc/sec. Upload leak check file. In the case of a leak and more than one re-worked/non-taped cavity, isolate the affected cavities and leak checked individually. The files for these leak checks shall be uploaded to a D3 associated with this traveler.After the leak check, close all open RAVs and remove the tape from closed RAVs."If any leak shows up on the RGA, all three cavities must be isolated and leak checked individually. The files for these leak ckecks shall be uploaded to a D3 associated with this traveler.After all three cavities are proven to be leak tight, the test stand can be moved to the VSA for sensor attachment, HOM tuning and VTA test. | [[CavitiesAnlogScan]] <<FILEUPLOAD>>[[CavitiesLkChk]] <<FILEUPLOAD>>[[CAVPosition1LkTight]] <<YESNO>>[[CAVPosition2LkTight]] <<YESNO>>[[CAVPosition3LkTight]] <<YESNO>>[[CavitiesLkChkTechnician]] <<SRFCVP>>[[CavityLkComment]] <<COMMENT>>[[CavityLkDate]] <<TIMESTAMP>> |