|  |  |
| --- | --- |
| Traveler Title | C75 Cavities VTA RF Test |
| Traveler Abstract | Cryogenic RF testing of C75 5-cell CEBAF cavities. Cavities can be tested single or as part of a hermetic pair.Cavity |
| Traveler ID | ER5C-VTA-CPR-VTRF |
| Traveler Revision  | R6 |
| Traveler Author | C. Wilson |
| Traveler Date | 17-May-2023 |
| NCR Informative Emails | areilly |
| NCR Dispositioners | kdavis,gciovati,grenoble |
| D3 Emails | kdavis,areilly,gciovati,grenoble |
| Approval Names | C. Wilson | G. Ciovati | T. Reilly | K. Davis |
| Approval Signatures |  |  |  |  |
| Approval Dates |  |  |  |  |
| Approval Title | Author | Reviewer | Project Manager | VTA facility manager |

|  |  |
| --- | --- |
| References | List and Hyperlink all documents related to this traveler. This includes, but is not limited to: safety (THAs, SOPs, etc), drawings, procedures, and facility related documents. |
| [VTA SOP](https://jlabdoc.jlab.org/docushare/dsweb/Get/Document-27461/A-09-001-SOP%20Operation%20of%20the%20Test%20Lab%20VTA%20Document-21542.pdf) | [CP-C75-CPR-VTRF-R4](https://jlabdoc.jlab.org/docushare/dsweb/Get/Document-272511/C75-PR-CPR-VTRF-R4.pdf) – VTA cryogenic RF test procedure for C75 cavities |  |  | Drawing |
|  | [Excel macro for C75-CPR-VTRF](https://jlabdoc.jlab.org/docushare/dsweb/Get/Document-254362/VTRF_Data_Processing_Spreadshet_Templete.xlsm) |  |  |  |

|  |  |
| --- | --- |
| Revision Note |  |
| R1 | Initial release of this Traveler. |
| R2 | Added performance limit |
| R3 | Updated link to the test procedure and added fields for Multipacting and Lorentz force detuning coefficient |
| R4 | Updated the high and low frequency specifications |
| R5 | Add: Hold point, highlight in yellow item that will cause an NCR to be created, administrative limit of 21 MV/m, list all cavity qualifications, correct link to RF data processing macro. |
| R6 | Step 18 updated step for performing an Unscaled Quench field Measurement with the step number in the procedure. |

|  |  |  |
| --- | --- | --- |
| Step No. | Instructions | Data Input |
| 1 | Select C75 5-cell cavity serial number (SN). (ex. 5C75-RI-001)Note any special handling, processing (chemistry or bake) or off-normal conditions associated with this cavity before test. | [[CAVSN]] <<CAVSN>>[[SpecialHandling]] <<COMMENT>> |
| 2 | Is this cavity part of a hermetic pair? If YES, note cavity serial number and cavity position on the test stand. | [[PartOfPair]] <<YESNO>>[[CPRSN]] <<CPRSN>>[[C75CavPosition]] {{Top,Bottom}} <<SELECT>> |
| 3 | Record if cavity is mechanically constrained (i.e. tuner attached, etc.). | [[TunerAttached]] <<YESNO>> |
| 4 | Enter the LabView file name, without special characters. | [[LabviewFile]] <<TEXT>> |
| 5 | Record Test Date, Dewar No, VTA Test Stand Serial Number and Operator(s). | [[TestDate]] <<TIMESTAMP>>[[Dewar]]{{8,4,7,3}} <<SELECT>>[[TSTDSN]]<<TSTDSN>>[[TestOperator1]] {{KDavis,CWilson,Liang,POwen,Reece,Powers,GCiovati,AdamG,Stirbet,None}} <<SELECT>>[[TestOperator2]] {{KDavis,CWilson,Liang,POwen,Reece,Powers,GCiovati,AdamG,Stirbet,None}} <<SELECT>> |
| 6 | If testing a single cavity, record cavity vacuum pressure. Cavity pairs are isolated per procedure and are not instrumented for vacuum readback. If the cavity pressure is <2e-5 mbar an RF test can be attempted. | [[CavityVacuum]] <<SCINOT>> [[CavityVacuumUnits]]{{(mbar),(Torr),(Pa)}} <<SELECT>>[[CavityVacuumOK]] <<YESNO>>[[CavityVacuumComment]] <<COMMENT>> |

|  |  |  |
| --- | --- | --- |
| Step No | Instructions | Data Inputs |
| 7 | Record Dewar helium bath liquid level, temperature and baratron pressure. Start cavity testing at (29 +/-0.1) Torr. | [[DewarLHeLevelcm]] <<FLOAT>> (cm)[[DewarTempK]]<<FLOAT>>(K)[[DewarPressureTorr]]<<FLOAT>>(Torr) |
| 8 | Zero power meters then calibrate cables at cavity fundamental frequency. If NO option is chosen launch D3 and record pertinent information. | [[PowermetersZeroed]] <<YESNO>>[[CableCalibrationOK]] <<YESNO>>[[CableCalibrationComment]] <<COMMENT>> |
| 9 | Perform low power measurements using a network analyzer (measure the five cavity mode frequencies). Record the cavity mode frequencies at the right. Example of cavity mode frequencies:Pi = 1497.388 MHz4Pi5 = 1492.711 MHz3Pi5 = 1480.413 MHz2Pi5 = 1466.088 MHz1Pi5 = 1454.720 MHz | [[Freq\_Pi]] <<FLOAT>> MHz[[Freq\_4Pi5]] <<FLOAT>> MHz[[Freq\_3Pi5]] <<FLOAT>> MHz[[Freq\_2Pi5]] <<FLOAT>> MHz [[Freq\_1Pi5]] <<FLOAT>> MHz |
| 10 | At 2.07 K determine and record dewar pressure (baratron) and cavity Pi-mode lock frequency precisely with LLRF frequency counter.Lock frequency specifications: Low: 1497.375 MHz High: 1497.575 MHzLaunch NCR if frequency specification is not met | [[LockFrequency]] <<FLOAT>> (MHz)[[LockFreqMeetsSpec]] <<YESNO>> |
| 11 | At 2.07 K and cavity field of 4-5 MV/m, determine cavity coupling state. | [[CavityCoupling]] {{Overcoupled,Undercoupled,Critically\_coupled}} <<SELECT>> |
| 12 | Perform decay measurements and record Eacc, Qo, Qext2, Qext1, %error, Tau, radiation, chosen for CW high power tests as specified. More information regarding these parameters can be found in CavID raw data.txt or CavID processed data.xlsx.Typical values during decay measurements for:Eacc: (4-5) MV/mQo: 1.0-1.6 e10Qext1: 0.5-2.4 e10Qext2: 0.8-3 e12error: 8-13%Tau : 0.3-1.2 sRadiation : 1e-3 mR/hr | [[Eacc]] <<FLOAT>> MV/m[[Qo]] <<SCINOT>>[[QextIn]] <<SCINOT>>[[QextFp]] <<SCINOT>>[[QextFpError]] <<FLOAT>> %[[Tau]] <<FLOAT>> Seconds[[Rad]] <<SCINOT>> mR/hr |

|  |  |  |
| --- | --- | --- |
| Step No | Instructions | Data Inputs |
| 13 | In Pi mode, test the cavity performance over its full dynamic range at 2.07 K. Be sure to capture a clean final-state data set from which to generate the post-processing Qo-vs-Eacc, Rad-vs-Eacc and f-vs-Eacc2 curves.At 2.07 K record:Final Maximum cavity gradient stably achieved Emax.Final Qo value at maximum cavity gradient.Final Qo at (20.5+/-0.3) MV/m. Acceptance criteria Qo >= 8e9.Cavity gradient at Qo specification (8e9)Initial value for Radiation at (20.5+/-0.3) MV/m.Initial FEonset: onset of field emission (FE onset, defined to be the first measured gradient where measured radiation is >= 3e-2 mR/hr). RadMax: value for the highest radiation level inside Dewar lid.Final Qo at Eacc= (4 +/- 0.3)MV/m. | [[EmaxMVm]] <<FLOAT>> MV/m[[QoAtEmax]] <<SCINOT>>[[QoAt20\_5MVm]] <<SCINOT>>[[EmaxAtQo\_Spec]] <<SCINOT>> MV/m[[Init\_RadAt20\_5MVm]] <<SCINOT>> mR/h[[Init\_FEonsetMVm]] <<FLOAT>> MV/m[[RadMax]] <<SCINOT>> mR/h[[QoAt4MVm]] <<SCINOT>> |
| 14 | Enter the value of the Lorentz force coefficient KL. Acceptable range: -3.7 to -4.2 Hz/(MV/m)2 | [[LorentzForceCoeff]] <<FLOAT>> Hz/(MV/m)^2 |
| 15 | Record if multipacting (MP) was found during the test, the Eacc at which occurred and if it was successfully processed.The C75 cavities have a "soft" MP barrier at 12-15 MV/m and a strong MP barrier at 17-24 MV/m. | [[MultipactingFound]] <<YESNO>>[[MultipactingGradient]] <<FLOAT>> MV/m[[MultipactingProcessed]] <<YESNO>>[[MultipactingComment]] <<COMMENT>> |
| 16 | At 2.07 K, was cavity RF processed to achieve performances over 20.5 MV/m. Record pertinent information at the right and lunch NCR if cavity RF processing was unsuccessful in achieving 20.5 MV/m with a Qo of 8e9. | [[RFProcessing]] <<YESNO>>[[RFProcessingComment]] <<COMMENT>> |
| 17 | Record after RF processing at 2.07 K:Post-processing Radiation at (20.5+/-0.3) MV/m.Post-processing FEonset: onset of field emission (FE onset, defined to be the first measured gradient where measured radiation is >= 3e-2 mR/hr).Post-processing RadMax value for the highest radiation level inside Dewar lid. | [[PP\_RadAt20\_5MVm]] <<SCINOT>> mR/h[[PP\_FEonsetMVm]] <<FLOAT>> MV/m[[PP\_RadMax]] <<SCINOT>> mR/h |
| 18 | At 2.07 K, if cavity is quench limited below Eacc 20.5 MV/m, find the unscaled Quench fields for each member of the fundamental passband as described in Section 8, Step 10 of C75-PR-CPR-VTRF and Launch NCR. | [[EaccQuench\_4Pi5]] <<FLOAT>> MV/m[[EaccQuench\_3Pi5]] <<FLOAT>> MV/m[[EaccQuench\_2Pi5]] <<FLOAT>> MV/m[[EaccQuench\_1Pi5]] <<FLOAT>> MV/m[[QuenchStudyComment]] <<COMMENT>> |

|  |  |  |
| --- | --- | --- |
| Step No | Instructions | Data Inputs |
| 19 | The cavity criteria to pass:Eacc >= 20.5 MV/m with a Qo >= 8e9.If the cavity does not meet the above criteria, record performance limitation.Definition of performance limitation terms:* Administrative Gradient: 21 MV/m
* Administrative Radiation: 1R/h admin limit
* FieldEmission: FE loaded Qo curve
* Quench: non-FE related quench limit
* Multipacting
* RF power: test limited by available RF power
* Cable: test aborted due to cable breakdown or broken cable
* Other: RF Tester defined

If Other is selected, record pertinent information in the Comment box at the right. | [[PerformanceLimit]] {{Quench,FieldEmission,Multipacting,RFPower,Cable,Admin,Other}} <<SELECT>>[[PerformanceLimitComment]] <<COMMENT>> |
| 20 | Upload the raw data file with VTA RF testing results using file name: CavID raw data.txt. | [[RF\_test\_raw\_data]] <<FILEUPLOAD>> |
| 21 | Upload processed (Excel) data file results using file name: CavID processed data.xlsx | [[RF\_test\_processed]] <<FILEUPLOAD>> |
| 22 | Upload processed Qo-and-Rad -vs-Eacc graph (in PDF format) using file name: QoandRadvsEacc.pdf Upload processed f-vs-Eacc2 graph (in PDF format) using file name: CavID\_FreqvsEacc2.pdf | [[UploadQvsE\_Rad]] <<FILEUPLOAD>>[[UploadFvsE2]] <<FILEUPLOAD>> |
| 23 | For a cavity pair test, measure the frequencies and loaded Q of all HOMs in the frequency range 1.74 GHz – 2.20 GHz for each cavity of the pair.Upload an Excel Spreadsheet with the HOM data using filename: CavID\_HOM.xlsx. | [[UploadHOMFile]] <<FILEUPLOAD>> |
| 24 | Cavity passed all specifications for this traveler? If NO option is chosen issue an NCR from this traveler.Issues that require an NCR opened:1. Cavity does not meet frequency specification.
2. Cavity RF processing was unsuccessful in achieving 20.5 MV/m with a Qo of 8e9.

Cavity is quench limited below Eacc 20.5 MV/m. | [[CavityMeetsSpecifications]] <<YESNO>>[[HOLDPOINT]] {{gciovati,kdavis}} <<HOLDPOINT>>  |