## TEN
| Componente | Operatore | Data-Ora
| BBF008 | M. Cleofi | 2022-01-17 10:42
QUALITY CONTROL SERVICE
CERTIFICATO DI PROVA DI TENUTA CON SPETTROMETRO DI MASSA metodo a vuoto | [16 empty] | N°. Documento |  |  | #VALUE!
Document No.
MASS SPECTROMETER LEAK TESTING vacuum method | [16 empty] | Pagina |  |  | 1 | di |  | 1
[17 empty] | Page |  |  |  | of
Cliente |  | FNAL | [7 empty] | Ordine N° |  |  |  |  | 647590
Customer | [9 empty] | P.O. No
Commessa |  | Z20008 | [7 empty] | Disegno |  |  |  |  | 3326.12.212.100
Mfr’ work No. | [9 empty] | Drawing
Apparecchio |  | RFD CRAB CAVITY | [7 empty] | N° Fabbrica |  |  |  |  | BBF008
Equipment | [9 empty] | Serial No.
P.C.Q. N°. |  | Z20008.QCP.001 | [7 empty] | Step |  |  |  |  | 11.13
Q.C.P. No. | [9 empty] | Fase
Reference Specification |  |  |  |  | Z20008.LT
Mass spectrometer |  |  |  |  | Oerlikon Phoenix XL300 Modul
Auxiliary pumping group |  |  |  |  | Scrollvac SC30D and Turbovac SL300
Vacuum gage |  |  |  |  | TTR91 (Pirani) + ITR90 (Bayard Alpert)
Pumping system |  |  |  |  | 4
Calibrated leak (qo) |  |  |  |  | 6.5e-09 |  | mbar · l · s-1 | [8 empty] | Building date |  |  |  | 2017 |  | @ | 23
Allowable leak |  |  |  |  | 1e-10 |  | mbar · l · s-1 | [8 empty] | Test date (dd/mm/yyyy) |  |  |  | 2022-01-17 10:42
Test gas |  |  | Helium |  | Concentration |  |  |  |  | 50% | [5 empty] | Test Temperature |  |  |  | 22 |  |  | °C
Correction coefficients |  |  |  |  | for the temperature K2= | [5 empty] | 0.035 |  | /°C |  | – over the years K1= |  |  |  | 0.0008 |  | /year
Calibration leak in testing conditions |  |  |  |  | qe=qo · (1-K1a)·[1+K2(te-to)]= | [6 empty] | 6.2474e-09
Tracer Probe |  |  |  |  | LACO Technologies Mod. LHSP01 Lot/N 11143 (ext)
Response time |  |  |  |  | 30 |  | seconds
TEST PARAMETERS
Background signal in the mass spectrometer (n1) | [5 empty] | 1e-11 | [7 empty] | mbar · l · s-1
Signal with calibrate leak opened (n2) | [5 empty] | 5.4e-09 | [7 empty] | mbar · l · s-1
Sensitivity in the mass spectrometer |  |  |  |  | K= qe/(n2-n1) = |  |  |  | 1.1591
Background signal at the end of the test (n) | [5 empty] | 1e-11 | [7 empty] | mbar · l · s-1
Pressure in the vessel |  |  |  |  | 9.7e-06 | [8 empty] | mbar
CALCULATION OF THE TOTAL LEAK : |  |  |  |  | Q=(n –n1) x qe/(n2-n1) x 1 / C= | [12 empty] | 1e-11
No leak greater than |  |  |  |  | 1e-11 |  |  | mbar · l · s-1 was found.
Note/Remarks
Risultato del test / Test result |  |  |  |  | Conforme / Satisfactory
Conforme a specifica / Conform to specification |  |  |  |  | Si / Yes
Eseguito da/Performed by |  |  |  |  | Liv./Level |  |  |  | Loc./Test place |  |  |  | Data/Date |  |  |  |  | Firma/Signature
Zanon R&I - |  |  | M. Cleofi |  | 2° (II) |  |  |  | Schio |  |  |  | 2022-01-17 10:42
Accettato/Accepted |  |  |  | Non Accettato/NotAccepted
Emesso da QCD |  |  |  | Verificato da QM | [5 empty] | Ispettore | [7 empty] | Ispettore
Issued by QCD |  |  |  | Reviewed by QM | [5 empty] | Inspector | [7 empty] | Inspector
M. Cleofi |  |  | Lev.1 | L. Lotti
|  |  | Lev.2 |  |  |  | Lev.2
|  |  | Lev.3 |  |  |  | Lev.3
Data/Date |  |  | 2022-01-17 10:42 | Data/Date |  | 2022-01-17 10:42

## Elenchi (hidden)
Personale | Ditta | Livello | Risultato | Strumenti | Gruppo pompaggio | Vacuometri | Gruppo | fughe calibrate | Anno | T | K1 | K2 | Tracer probe | Conform to specification
M. Ongaro | Zanon R&I | 2° (II) | No Recordable Indications. | Oerlikon Phoenix XL300 Modul | Scrollvac SC30D and Turbovac SL300 | TTR91 (Pirani) + ITR90 (Bayard Alpert) | 1 | 4.4e-09 | 2011 | 23 | 0.0015 | 0.035 | LACO Technologies Mod. LHSP01 Lot/N 11143 (ISO4) | Si / Yes
M. Cleofi | Zanon R&I | 2° (II) | Nonrejectable Indications were found. See Below. |  |  |  | 2 | 4.2e-09 | 2011 | 23 | 0.0014 | 0.035 | LACO Technologies Mod. LHSP01 Lot/N 11143 (ext) | No
M. Fanchin | Zanon R&I | 2° (II) | Rejectable Indications were found. See Below. |  |  |  | 3 | 4.3e-09 | 2013 | 24.4 | 0.0014 | 0.035 | LACO Technologies Mod. LHSP01 Lot/N 11143 (ext)
M. Festa | Zanon R&I | 2° (II) |  |  |  |  | 4 | 4.3e-09 | 2011 | 23 | 0.0015 | 0.035 | LACO Technologies Mod. LHSP01 Lot/N 11143 (ext)
[7 empty] | 5 | 5.5e-09 | 2011 | 23 | 0.0019 | 0.035 | LACO Technologies Mod. LHSP01 Lot/N 11143 (ISO4)
[7 empty] | 6 | 4.3e-09 | 2013 | 23 | 0.0014 | 0.035 | LACO Technologies Mod. LHSP01 Lot/N 11143 (ISO4)
[7 empty] | V | 5.3e-09 | 2015 | 23 | 0.0017 | 0.035 | LACO Technologies Mod. LHSP01 Lot/N 11143 (ISO4)
M. Ongaro | Zanon R&I | 2° (II)
M. Cleofi | Zanon R&I | 2° (II)
M. Fanchin | Zanon R&I | 2° (II)
M. Festa | Zanon R&I | 2° (II)
L. Lotti | Zanon R&I