| Traveler Title | NbSn Quarter Cryomodule Acceptance Testing |  |  |  |
| --- | --- | --- | --- | --- |
| Traveler Abstract | NbSn Quarter Cryomodule Acceptance Testing. This traveler covers cool down, low power and high power testing and the warm up |  |  |  |
| Traveler ID | NB3SN-CMTF-CM-ACTS |  |  |  |
| Traveler Revision | R1 |  |  |  |
| Traveler Author | M. Drury |  |  |  |
| Traveler Date | 09-Apr-24 |  |  |  |
| NCR Informative Emails | areilly |  |  |  |
| NCR Dispositioners | drury,fischer,michaelm,uttar |  |  |  |
| D3 Emails | drury,areilly,fischer,uttar |  |  |  |
| Approval Names | M. Drury | G. Ciovati | M. McCaughan | A, Reilly |
| Approval Signatures |  |  |  |  |
| Approval Dates |  |  |  |  |
| Approval Title | Author | Reviewer | Reviewer | Project Manager |

| References | List and Hyperlink all documents related to this traveler. This includes, but is not limited to: safety (THAs, SOPs, etc), drawings, procedures, and facility related documents. |  |  |  |
| --- | --- | --- | --- | --- |
| SRF-22-131703-OSP |  |  |  |  |

| Revision Note |  |
| --- | --- |
| R1 | Initial release of this Traveler. |

| Step No. | Instructions | Data Input |
| --- | --- | --- |
| 1 | Record the Cryomodule serial number. | [[CMSN]] <<CMSN>> |
| 2 | Record the Cavity SN's for each cavity position. (Note: Cavity 7-Supply side, Cavity 8-Return side) | [[CAVSN7]] <<CAVSN>> [[CAVSN8]] <<CAVSN>> |
| 3 | Verify that the ground cable is connected to the cryomodule. | [[GroundCableInspector]] <<SRF>> [[GroundCableTime]] <<TIMESTAMP>> [[GroundCableVerified]] (Yes) <<CHECKBOX>> |
| 4 | Verify that the Insulating Vacuum gauge is connected to data acquisition and the correct readout is available in the Control Room. Record the Insulating Vacuum pressure (in torr). | [[WarmInsVacInspector]] <<SRF>> [[WarmInsVacTime]] <<TIMESTAMP>> [[WarmInsVacPressure]] <<SCINOT>> (torr) |
| 5 | Verify that the Beamline ion pump is connected to data acquisition and the correct readout is available in the Control Room. Record the Beamline Vacuum pressure (in torr). | [[WarmBLVacInspector]] <<SRF>> [[WarmBLVacTime]] <<TIMESTAMP>> [[WarmBLVacPressure]] <<SCINOT>> (torr) |
| 6 | Verify that the Waveguide Vacuum ion pumps are connected to data acquisition and that the correct readouts are available in the Control Room. Record the Waveguide Vacuum pressures (in torr). | [[WarmWGVacInspector]] <<SRF>> [[WarmWGVacTime]] <<TIMESTAMP>> [[WarmWG7VacPressure]] <<SCINOT>> (torr) [[WarmWG8VacPressure]] <<SCINOT>> (torr) |
| 7 | List any problems involving vacuum pumps or vacuum instrumentation. | [[VacuumProblems]] <<COMMENT>> |

| Step No | Instructions | Data Inputs |
| --- | --- | --- |
| 8 | Verify that the cavity temperature measurement Cernox sensors are connected to data acquisition and that the correct read outs are available in the Control Room. List any sensors, by epics PV and physical location, that do not read correctly in comment block. | [[CavityTempSensorInspector]] <<SRF>> [[CavityTempSensorTime]] <<TIMESTAMP>> [[CavitySensorsVerified]] <<YESNO>> [[DeadCavitySensorList]] <<COMMENT>> |
| 9 | Verify that End Can Diodes are connected to data acquisition and the correct readouts are available in Control Room. List any diodes, by epics PV and physical location, that do not read correctly in comment block. | [[EndCanDiodeInspector]] <<SRF>> [[EndCanDiodeTime]] <<TIMESTAMP>> [[SupplyEndCanDiodesVerified]] <<YESNO>> [[ReturnEndCanDiodesVerified]] <<YESNO>> [[DeadEndCanDiodeList]] <<COMMENT>> |
| 10 | Verify that Liquid Level Sensors on both Supply and Return end cans are connected to data acquisition and the correct readouts are available in the Control Room. Note any problems in the Comment block. | [[LiqLevelProbeInspector]] <<SRF>> [[LiqLevelProbeTime]] <<TIMESTAMP>> [[SupplyLiqLevelProbeVerified]] <<YESNO>> [[ReturnLiqLevelProbeVerified]] <<YESNO>> |
| 11 | Verify that Pressure Transducers (50 torr, 100 torr and 5000 torr) are connected to data acquisition and the correct readouts are available in the Control Room. Note any problems in the Comment block. | [[PressureTransducerInspector]] <<SRF>> [[PressureTransducerTime]] <<TIMESTAMP>> [[TransducerVerified50TorrPress]] <<YESNO>> [[TransducerVerified100TorrPress]] <<YESNO>> [[TransducerVerified5KTorrPress]] <<YESNO>> |
| 12 | Verify that all Heaters are connected and operational. List any problems in the Comment block. (This inspection should be completed after liquid Helium level reaches 20%.) | [[CryoHeaterInspector]] <<SRF>> [[CryoHeaterInspectorTime]] <<TIMESTAMP>> [[CryoHeaterInspectorVerified]] <<YESNO>> |
| 13 | Verify that JT valve actuator is installed and set up properly. Verify that the LVDT is correctly calibrated through the epics control system. List JT valve readings that correspond to completely closed and fully opened. | [[JTValveInspector]] <<SRF>> [[JTValveTime]] <<TIMESTAMP>> [[JTValveInformation]] <<COMMENT>> |
| 14 | List any problems involving the process instrumentation and controls listed in steps 8 -13 above. | [[ProcessInstrProblems]] <<COMMENT>> |

| Step No | Instructions | Data Inputs |
| --- | --- | --- |
| 15 | Verify that all Arc Detector assemblies are installed and are connected to power and interlock controls. | [[ArcDetectorInstaller]] <<SRF>> [[ArcDetectorInstTime]] <<TIMESTAMP>> [[ArcDetectorsVerified]] <<YESNO>> [[DeadArcDetectorList]] <<COMMENT>> |
| 16 | Verify that all IR Detectors are connected to data acquisition and the correct readouts are available in the Control Room. List any problems in the comment block. | [[IRDetectorInspector]] <<SRF>> [[IRDetectorTime]] <<TIMESTAMP>> [[IRDetectorsVerified]] <<YESNO>> [[DeadIRDetectorList]] <<COMMENT>> |
| 17 | Verify that the Stepper Motors and Limit Switches are connected. | [[StepperInspector]] <<SRF>> [[StepperTime]] <<TIMESTAMP>> [[StepperVerified]] <<YESNO>> [[StepperProblems]] <<COMMENT>> |

| Step No | Instructions |  |  |  | Data Inputs |
| --- | --- | --- | --- | --- | --- |
| 18 | Record cavity warm frequencies (in MHz) in table below. |  |  |  | [[WarmFreqOperator]] <<SRF>> [[WarmFreqTime]] <<TIMESTAMP>> [[WarmFreqComments]] <<COMMENT>> |
| Cavity Number | 1/5 Pi | 2/5 Pi | 3/5 Pi | 4/5 Pi | Pi |
| 7 | [[Cav7Warm15Pi]] <<FLOAT>> | [[Cav7Warm25Pi]] <<FLOAT>> | [[Cav7Warm35Pi]] <<FLOAT>> | [[Cav7Warm45Pi]] <<FLOAT>> | [[Cav7WarmPi]] <<FLOAT>> |
| 8 | [[Cav8Warm15Pi]] <<FLOAT>> | [[Cav8Warm25Pi]] <<FLOAT>> | [[Cav8Warm35Pi]] <<FLOAT>> | [[Cav8Warm45Pi]] <<FLOAT>> | [[Cav8WarmPi]] <<FLOAT>> |

| Cool Down Section |  |  |
| --- | --- | --- |
| 19 | Verify that U-tube installation is complete. List any problems in the comment block. | [[UTubeInstallTech]] <<SRF>> [[UTubeInstallTime]] <<TIMESTAMP>> [[UTubeInstallComments]] <<COMMENT>> |
| 20 | Verify that all U-Tube Diodes are connected to data acquisition and the correct readouts are available in the Control Room. List any diodes, by epics PV and physical location, that do not read correctly in comment block. | [[UTubeDiodeInspector]] <<SRF>> [[UTubeDiodeTime]] <<TIMESTAMP>> [[UTubeDiodeComments]] <<COMMENT>> |
| 21 | Return U-tube pneumatic valve: - Verify that air-line is connected. - Verify that purge line is connected. - Verify that valve is bagged. - Verify that solenoid switch in control room is in open position. - Verify that pressure bleed valve on cave wall is open. - Verify operation of the valve. List any problems in the Comment block. | [[RTValveTech]] <<SRF>> [[RTValveTime]] <<TIMESTAMP>> [[RTVerificationComplete]] <<CHECKBOX>> [[RTValveProblems]] <<COMMENT>> |
| 22 | Install guard vacuum on circle seal and sub-atmospheric U-tube. | [[GuardVacInstTech]] <<SRF>> [[GuardVacInstTime]] <<TIMESTAMP>> [[GuardVacInstProblems]] <<COMMENT>> |

| 4K Cool Down |  |  |
| --- | --- | --- |
| Step No | Instructions | Data Inputs |
| 23 | Record date and time of cool down start. Reference temperature sensors on midplane HOM’s and at either end of the cryomodule to determine time. Note: Cool down rate must be less than 30K / hour from room temperature to 4K as measured by referenced sensors. Thermal gradient across cavity pair, as measured by Cernox sensors on the cavity beam tubes, must be regulated to T < 0.3K between 20K and 10K. Note: He vessel pressure should remain below 950 Torr. | [[NBSNCoolDownOperator]] <<SRF>> [[NBSNCoolDownStartTime]] <<TIMESTAMP>> |
| 24 | Record date and time at which cool down to 4 K was completed. Reference Cernox sensors on beam tubes at either end of the cavity pair to determine time. | [[NBSNCoolDownOperator4k]] <<SRF>> [[NBSNCoolDown4KTime]] <<TIMESTAMP>> |
| 25 | Note any problems associated with the cool down or the pump down and fill. | [[NBSNCMCoolDownProblems]] <<COMMENT>> |
| 26 | Upload logfiles generated by Labview. Logfiles should cover the entire period from start of cool down to end of pump down and fill. | [[NBSNCMCoolDownFiles]] <<FILEUPLOAD>> |

| 4K Cool Down |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Step No | Instructions |  |  |  |  |  | Data Inputs |  |  |
| 27 | Record date and time of cool down start. Record Pi Mode frequency (in MHz) of each cavity every 50K during cool down, as well as at 4K and 2K (after initial 4K testing). Measured frequencies should be +/- 100kHz of target. If not, notify test director / NbSn SME’s |  |  |  |  |  | [[CoolDownTempMonitor]] <<SRF>> [[CoolDownStartTime]] <<TIMESTAMP>> |  |  |
| Cavity Number | Room Temperature | 250K | 200K | 150K | 100K | 50K |  | 4K | 2K |
| 7 (5C75-RI-NbSn1) | Target: 1494.174 MHz [[NBSNCav7RTFreq]] <<FLOAT>> | Target: 1494.52 MHz [[NBSNCav7250KFreq]] <<FLOAT>> | Target: 1494.91 MHz [[NBSNCav7200KFreq]] <<FLOAT>> | Target: 1495.32 MHz [[NBSNCav7150KFreq]] <<FLOAT>> | Target: 1495.71 MHz [[NBSNCav7100KFreq]] <<FLOAT>> | Target: 1496.11 MHz [[NBSNCav750KFreq]] <<FLOAT>> |  | Target: 1496.47 MHz [[NBSNCav74KFreq]] <<FLOAT>> | Target: 1496.62 MHz [[NBSNCav72KFreq]] <<FLOAT>> |
| 8 (5C75-RI-04) | Target: 1494.024 MHz [[NBSNCav8RTFreq]] <<FLOAT>> | Target: 1494.37 MHz [[NBSNCav8250KFreq]] <<FLOAT>> | Target: 1494.76 MHz [[NBSNCav8200KFreq]] <<FLOAT>> | Target: 1495.16 MHz [[NBSNCav8150KFreq]] <<FLOAT>> | Target: 1495.56 MHz [[Cav8CDFreq100K]] <<FLOAT>> | Target: 1495.96 MHz [[NBSNCav850KFreq]] <<FLOAT>> |  | Target: 1496.32 [[NBSNCav84KFreq]] <<FLOAT>> | Target: 1496.47 [[NBSNCav82KFreq]] <<FLOAT>> |
| 28 | Upload any associated files. |  |  |  |  |  | [[NBSNTempVFreqFiles]] <<FILEUPLOAD>> |  |  |

| Step No | Instructions | Data Inputs |
| --- | --- | --- |
| 29 | Begin the Integrated Insulating Vacuum Leak Check: Isolate the insulating vacuum from the pumping station after the pump down and fill are complete and the cryomodule primary pressure and liquid level are stable. Monitor the insulating vacuum pressure for at least 1 week. Record start, time, completion time and the vacuum pressure (in torr) at start and finish. | [[InsVacLeakTech1]] <<SRF>> [[InsVacLeakCheckStartTime]] <<TIMESTAMP>> [[InsVacLeakCheckStartPressure]] <<SCINOT>> [[InsVacLeakTech2]] <<SRF>> [[InsVacLeakCheckEndTime]] <<TIMESTAMP>> [[InsVacLeakCheckEndPressure]] <<SCINOT>> |

| High Power Checkout |  |  |
| --- | --- | --- |
| Step No | Instructions | Data Inputs |
| 30 | Complete the RF Cable Calibration Measurements for the cryomodule under test. | [[RFCableCalTech]] <<SRF>> [[RFCableCalTime]] <<TIMESTAMP>> [[RFCableCalWorksheet]] <<FILEUPLOAD>> |
| 31 | Test Arc Detectors for Cavities 1-8. Verify that each detector generates a fault and disables RF. Check off each working arc detector interlock. **Do Not Attempt to Supply High Power RF to Cavity if its Arc Detector and Interlock are not Functioning Correctly!** | [[ArcDetectorInspector]] <<SRF>> [[ArcDetectorTime]] <<TIMESTAMP>> [[C7ArcDetectorIntlkGood]] <<YESNO>> [[C8ArcDetectorIntlkGood]] <<YESNO>> |
| 32 | Test Warm Window IR Detectors for both cavities. Verify that each detector generates a fault and disables RF. Check off each working IR detector interlock. Note any problems in the Comment block. **Do Not Attempt to Supply High Power RF to Cavity if its IR Detector and Interlock are not Functioning Correctly!** | [[IRDetectorTestInspector]] <<SRF>> [[IRDetectorTestTime]] <<TIMESTAMP>> [[C7IRDetectorIntlkGood]] <<YESNO>> [[C8IRDetectorIntlkGood]] <<YESNO>> |
| 33 | Use the Comment block to list any problems associated with Arc and/or IR detectors. | [[ArcAndIRProblems]] <<COMMENT>> |

| Step No | Instructions | Data Inputs |
| --- | --- | --- |
| 34 | Test the Beamline Vacuum Interlock. Verify that a fault is generated and RF is disabled. Check the box if working. ** Do Not Attempt to Supply High Power RF to any Cavity if Beamline Vacuum Interlock is Not Working Properly ** | [[BLVacIntlkInspector]] <<SRF>> [[BLVacIntlkInspectTime]] <<TIMESTAMP>> [[BLVacIntlkGood]] <<YESNO>> |
| 35 | Test the Waveguide Vacuum. Verify that a fault is generated and RF is disabled. Check off each working interlock ** Do Not Attempt to Supply High Power RF to Cavity if its Waveguide Vacuum Interlock is Not Working Properly ** | [[WGVacIntlkInspector]] <<SRF>> [[WGVacIntlkInspectTime]] <<TIMESTAMP>> [[WG7VacIntlkGood]] <<YESNO>> [[WG8VacIntlkGood]] <<YESNO>> |
| 36 | Use the Comment block to list any problems associated with vacuum interlocks. | [[VacuumIntlkProblems]] <<COMMENT>> |

| Step No | Instructions | Data Inputs |
| --- | --- | --- |
| 37 | Verify that DecaRad Geiger-Mueller tubes are connected and correctly positioned around cryomodule. Ensure that read outs are active in epics. Typical setup for this type of cryomodule:: - Channel 5 at Top Center of Cryomodule - Channel 7 at Top Hat between FPC’s - Channel 8 at Supply Side of Beamline - Channel 9 at Return Side of Beamline. Deviations from this layout should be noted in the Comment box. Any problems such as dead channels should be noted in the Comment box. | [[DecaRadTech]] <<SRF>> [[DecaRadSetupTime]] <<TIMESTAMP>> [[DecaRadDeviations]] <<COMMENT>> |
| 38 | Inspect all waveguide connections in the Test Cave. All waveguide sections must be in place with all flange bolt holes filled. Bolts must not be loose. **No waveguide that is capable of delivering RF power into the cave may be open during High power RF. Waveguides must be terminated either by connection to cryomodule or by shorting plate or an appropriate load. Do not perform PSS Sweep of Test Cave until this step has been completed.** | [[WaveGuideTech]] <<SRF>> [[WaveGuideCheckTime]] <<TIMESTAMP>> [[WaveGuideComments]] <<COMMENT>> |
| 39 | Inspect all RF heliax cable connections. A proper connection means at least hand tightened and connected to the appropriate connector. Ensure that: - All Ptrans cables are properly connected to the appropriate field probe connectors. - All RF Heliax cables hanging from RF patch panels 1A, 2A, 3A, 4A are properly connected. - All cables coming from the directional couplers are properly connected. - All cables are properly connected to waveguide switching chassis located on South wall of cave. **Do not perform PSS Sweep of Test Cave until this step has been completed.** | [[RFCableTech]] <<SRF>> [[RFCableCheckTime]] <<TIMESTAMP>> [[RFCableComments]] <<COMMENT>> |

| Step No | Instructions | Data Inputs |
| --- | --- | --- |
| 40 | PSS Function Check: After completing a sweep of the cave, bring the RF HPA into the high voltage state. Drop the PSS state from Run to Open Access using the PSS key and verify that HPA has dropped out of the High Voltage. Note any problems in comment block. **If PSS fails this test, contact PSS on-call personnel immediately. HPRF must not be supplied to cavities until this test is completed successfully.** | [[PSSFunctionTech]] <<SRF>> [[PSSFunctionCheckTime]] <<TIMESTAMP>> [[PSSProblems]] <<COMMENT>> [[PSSFunctionTestPassed]] <<YESNO>> |

| Cavity 7 Gradient and Field Emission at 4K |  |  |
| --- | --- | --- |
| Step No | Instructions | Data Inputs |
| 41 | Record QextFPC and QextFP for Cavity 7 and the gradient at which measurement was made (in MV/m). If QextFPC for either cavity does not meet the requirement of 2.0E7 +/- 20%, write an NCR | [[C74KQextTech]] <<SRF>> [[C74KQextMeasTime]] <<TIMESTAMP>> [[C74KQextFPC]] <<SCINOT>> [[C74KQextFP]] <<SCINOT>> [[C74KQextMeasGradient]] <<FLOAT>> |
| 42 | Record the Maximum gradient (Emax) for Cavity 7 and the gradient limiting condition. | [[C74KEmaxTech]] <<SRF>> [[C74KEmaxMeasTime]] <<TIMESTAMP>> [[C74KEmax]] <<FLOAT>> (MV/m) [[C74KEmaxLimit]] {{Admin Limit,Quench,FE related,ArcFault,Window Temp Fault,BLVacuum Fault,WG Vacuum Fault,RF Power,Heat Load,End Group Quench}} <<SELECT>> |
| 43 | Record the gradient at which a successful One Hour Run was completed for Cavity 7. Upload spreadsheet containing data on the One Hour Run. If the One Hour Run Gradient is lower than 14.0 MV/m, write an NCR. | [[C74KOneHourRunTech]] <<SRF>> [[C74KOneHourRunTime]] <<TIMESTAMP>> [[C74KEmaxop]] <<FLOAT>> [[C74KOneHourRunFile]] <<FILEUPLOAD>> |
| 44 | Record Static Lorentz value for Cavity 7 calculated from data collected during the Field Emission Measurement gradient ramping. | [[C74KStaticLorentz]] <<FLOAT>> [[C74KStaticLorentzFile]] <<FILEUPLOAD>> |
| 45 | Record the Field Emission Onset gradient for Cavity 7. Upload the file containing Field Emission data. | [[C74KFETech]] <<SRF>> [[C74KFEMeasTime]] <<TIMESTAMP>> [[C74KFEOnset]] <<FLOAT>> (MV/m) [[C74KFEMaxDoseRate]] <<FLOAT>> (R/hr) [[C74KFEFile]] <<FILEUPLOAD>> |
| 46 | Record any information about the performance of Cavity 7 in the Comment block. | [[C74KEmaxTestComments]] <<COMMENT>> |

| Cavity 8 Gradient and Field Emission at 4K |  |  |
| --- | --- | --- |
| Step No | Instructions | Data Inputs |
| 47 | Record QextFPC and QextFP for Cavity 8 and the gradient at which measurement was made (in MV/m). If QextFPC for either cavity does not meet the requirement of 2.0E7 +/- 20%, write an NCR | [[C84KQextTech]] <<SRF>> [[C84KQextMeasTime]] <<TIMESTAMP>> [[C84KQextFPC]] <<SCINOT>> [[C84KQextFP]] <<SCINOT>> [[C84KQextMeasGradient]] <<FLOAT>> |
| 48 | Record the Maximum gradient (Emax) for Cavity 8 and the gradient limiting condition. | [[C84KEmaxTech]] <<SRF>> [[C84KEmaxMeasTime]] <<TIMESTAMP>> [[C84KEmax]] <<FLOAT>> (MV/m) [[C84KEmaxLimit]] {{Admin Limit,Quench,FE related,Arc Fault,Window Temp Fault,BL Vacuum Fault,WG Vacuum Fault,RF Power,Heat Load,End Group Quench}} <<SELECT>> |
| 49 | Record the gradient at which a successful One Hour Run was completed for Cavity 8. Upload spreadsheet containing data on the One Hour Run. If the One Hour Run Gradient is lower than 8.0 MV/m, write an NCR. | [[C84KOneHourRunTech]] <<SRF>> [[C84KOneHourRunTime]] <<TIMESTAMP>> [[C84KEmaxop]] <<FLOAT>> [[C84KOneHourRunFile]] <<FILEUPLOAD>> |
| 50 | Record Static Lorentz value for Cavity 8 calculated from data collected during the Field Emission Measurement gradient ramping. | [[C84KStaticLorentz]] <<FLOAT>> [[C84KStaticLorentzFile]] <<FILEUPLOAD>> |
| 51 | Record the Field Emission Onset gradient for Cavity 8. Upload the file containing Field Emission data. | [[C84KFETech]] <<SRF>> [[C84KFEMeasTime]] <<TIMESTAMP>> [[C84KFEOnset]] <<FLOAT>> (MV/m) [[C84KFEMaxDoseRate]] <<FLOAT>> (R/hr) [[C84KFEFile]] <<FILEUPLOAD>> |
| 52 | Record any information about the performance of Cavity 8 in the Comment block | [[C84KEmaxTestComments]] <<COMMENT>> |

| 30K Cryocycle |  |  |
| --- | --- | --- |
| Step No | Instructions | Data Inputs |
| 53 | Lower Liquid Level to 20 % Record time at which JT valve is closed. Record Power Level for HeVessel Heater. Record time at which liquid level reaches 20%. | [[NBSN30KOperator]] <<SRF>> [[NBSN30KStartTime]] <<TIMESTAMP>> [[NBSN30KBoilOffHeaterLvl]] <<FLOAT>> [[NBSN30KLvl20Time]] <<TIMESTAMP>> |
| 54 | Cryomodule Warm Up Heaters must be turned off when temperatures reach 25K Record Heater Setting for warm up portion. Record time when liquid level reaches approximately 0%. Record time when all cavity temperature sensor read 25K or higher or when any single cavity temperature sensor reaches 45K. | [[NBSN30KWarmUpHtrSetting]] <<FLOAT>> [[NBSN30KZeroLvlTime]] <<TIMESTAMP>> [[NBSN30K25KTime]] <<TIMESTAMP>> (MV/m) |
| 55 | Record any observations about the boil off and warm up portions of the cryocycle in the comment block. Attach any files relevant to Warm up portion of Cryocycle. | [[NBSN30KWarmUpComments]] <<COMMENT>> [[NBSN30KLogFiles]] <<FILEUPLOAD>> |
| 56 | Cryomodule Cool Down Record time at which JT valve is opened. Record time at which any cavity temperature reaches 4K. Record time at which liquid level reaches 90%. | [[NBSN30KCoolDownStartTime]] <<TIMESTAMP>> [[NBSN30KCoolDown4KTime]] <<TIMESTAMP>> [[NBSN30K90LvlTime]] <<FLOAT>> |
| 57 | Record any observations about the boil off and warm up portions of the cryocycle in the comment block. Attach any files relevant to Cool Down portion of Cryocycle. | [[NBSN30KCoolDownComments]] <<COMMENT>> [[NBSN30KCoolDownFiles]] <<FILEUPLOAD>> |

| Cavity 7 Q0 at 4K |  |  |
| --- | --- | --- |
| Step No | Instructions | Data Inputs |
| 58 | Record QextFPC and QextFP for Cavity 7 and the gradient at which measurement was made (in MV/m). If QextFPC for either cavity does not meet the requirement of 2.0E7 +/- 20%, write an NCR. | [[C74KQextTech]] <<SRF>> [[C74KQextMeasTime]] <<TIMESTAMP>> [[C74KQextFPC]] <<SCINOT>> [[C74KQextFP]] <<SCINOT>> [[C74KQextMeasGradient]] <<FLOAT>> |
| 59 | After completing the Q0 measurement sequence for Cavity 7, record the values of Q0 at 10 MV/m, 14.0 MV/m and at Emaxop. Upload the Q0 measurement file. If Q0 at 10.0 MV/m is less than 8.0E9, write an NCR. If QextFPC for either cavity does not meet the requirement of 2.0E7 +/- 20%, write an NCR. | [[C74KQoTech]] <<SRF>> [[C74KQoMeasTime]] <<TIMESTAMP>> [[C74KQoTen]] <<SCINOT>> [[C74KQoFourteen]] <<SCINOT>> [[C74KQoEmaxop]] <<SCINOT>> [[C74KQoFile]] <<FILEUPLOAD>> |
| 60 | Record the Pressure Sensitivity data collected for Cavity 7 during the Q0 measurement sequence. | [[C74KPressSensitivity]] <<FLOAT>> [[C74KPressSensitivityFile]] <<FILEUPLOAD>> |
| 61 | Record any information about the Q0 performance at 4K of Cavity 7 in the Comment block. | [[C74KQoTestComments]] <<COMMENT>> |

| Cavity 8 Q0 at 4K |  |  |
| --- | --- | --- |
| Step No | Instructions | Data Inputs |
| 62 | Record QextFPC and QextFP for Cavity 7 and the gradient at which measurement was made (in MV/m). If QextFPC for either cavity does not meet the requirement of 2.0E7 +/- 20%, write an NCR. | [[C84KQextTech]] <<SRF>> [[C84KQextMeasTime]] <<TIMESTAMP>> [[C84KQextFPC]] <<SCINOT>> [[C84KQextFP]] <<SCINOT>> [[C84KQextMeasGradient]] <<FLOAT>> |
| 63 | After completing the Q0 measurement sequence for Cavity 8, record the values of Q0 at 7 MV/m, 8.0 MV/m and at Emaxop. Upload the Q0 measurement file. If Q0 at 7.0 MV/m is less than 3.0E9, write an NCR. If QextFPC for either cavity does not meet the requirement of 2.0E7 +/- 20%, write an NCR. | [[C84KQoTech]] <<SRF>> [[C84KQoMeasTime]] <<TIMESTAMP>> [[C84KQoSeven]] <<SCINOT>> [[C84KQoEight]] <<SCINOT>> [[C84KQoEmaxop]] <<SCINOT>> [[C84KQoFile]] <<FILEUPLOAD>> |
| 64 | Record the Pressure Sensitivity data collected for Cavity 8 during the Q0 measurement sequence. | [[C84KPressSensitivity]] <<FLOAT>> [[C84KPressSensitivityFile]] <<FILEUPLOAD>> |
| 65 | Record any information about the Q0 performance of Cavity 8 in the Comment block. | [[C84KQoTestComments]] <<COMMENT>> |

| 2K Pump Down |  |  |
| --- | --- | --- |
| 66 | Record date and time at which pump-down to 2K was initiated. Record time at which 2K fill is complete. (Defined as the time the liquid level measured at return end of the cryomodule reaches the set point.) | [[NBSNPumpDownOperator]] <<SRF>> [[NBSNPumpDownStartTime]] <<TIMESTAMP>> [[NBSNLiqLvlFillCompleteTime]] <<TIMESTAMP>> |
| 67 | Verify that the guard vacuum is being actively pumped. | [[GuardVacVerifyTech]] <<SRF>> [[GuardVacVerifyTime]] <<TIMESTAMP>> |
| 68 | Note any problems associated with the pump down and fill. | [[CMCoolDownProblems]] <<COMMENT>> |
| 69 | Upload logfiles generated by Labview. Logfiles should cover the entire period from start of cool down to end of pump down and fill. | [[CMCoolDownFiles]] <<FILEUPLOAD>> |

| Cavity 7 Q0 at 2K |  |  |
| --- | --- | --- |
| Step No | Instructions | Data Inputs |
| 70 | Record QextFPC and QextFP for Cavity 7 and the gradient at which measurement was made (in MV/m). If QextFPC for either cavity does not meet the requirement of 2.0E7 +/- 20%, write an NCR. | [[C72KQextTech]] <<SRF>> [[C72KQextMeasTime]] <<TIMESTAMP>> [[C72KQextFPC]] <<SCINOT>> [[C72KQextFP]] <<SCINOT>> [[C72KQextMeasGradient]] <<FLOAT>> |
| 71 | After completing the Q0 measurement sequence for Cavity 7, record the values of Q0 at 10.0 MV/m, 19.0 MV/m and Emaxop. Upload the Q0 measurement file. If Q0 at 10.0 MV/m is less than 2.0E10, write an NCR. | [[C72KQoTech]] <<SRF>> [[C72KQoMeasTime]] <<TIMESTAMP>> [[C72KQoTen]] <<SCINOT>> [[C72KQoNineteen]] <<SCINOT>> [[C72KQoEmaxop]] <<SCINOT>> [[C72KQoFile]] <<FILEUPLOAD>> |
| 72 | Record the Pressure Sensitivity data collected for Cavity 7 during the Q0 measurement sequence. | [[C72KPressSensitivity]] <<FLOAT>> [[C72KPressSensitivityFile]] <<FILEUPLOAD>> |
| 73 | Record any information about the performance of Cavity 7 in the Comment block. | [[C72KQoTestComments]] <<COMMENT>> |

| Cavity 8 Q0 at 2K |  |  |
| --- | --- | --- |
| Step No | Instructions | Data Inputs |
| 74 | Record QextFPC and QextFP for Cavity 7 and the gradient at which measurement was made (in MV/m). If QextFPC for either cavity does not meet the requirement of 2.0E7 +/- 20%, write an NCR. | [[C82KQextTech]] <<SRF>> [[C82KQextMeasTime]] <<TIMESTAMP>> [[C82KQextFPC]] <<SCINOT>> [[C82KQextFP]] <<SCINOT>> [[C82KQextMeasGradient]] <<FLOAT>> |
| 75 | After completing the Q0 measurement sequence for Cavity 8, record the values of Q0 at 7 MV/m, 8.0 MV/m and at Emaxop. Upload the Q0 measurement file. If Q0 at 7.0 MV/m is less than 3.0E9, write an NCR. | [[C82KQoTech]] <<SRF>> [[C82KQoMeasTime]] <<TIMESTAMP>> [[C82KQoSeven]] <<SCINOT>> [[C82KQoEight]] <<SCINOT>> [[C82KQoEmaxop]] <<SCINOT>> [[C82KQoFile]] <<FILEUPLOAD>> |
| 76 | Record the Pressure Sensitivity data collected for Cavity 8 during the Q0 measurement sequence. | [[C82KPressSensitivity]] <<FLOAT>> [[C82KPressSensitivityFile]] <<FILEUPLOAD>> |
| 77 | Record any information about the Q0 performance of Cavity 8 in the Comment block. | [[C82KQoTestComments]] <<COMMENT>> |

| Tuner Tests |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Step No | Instructions |  |  |  | Data Inputs |  |  |
| 78 | Complete the +/- 500 kHz Mechanical Tuner Range and +/- 2 kHz Loop Tests for cavities 1-8. Check off the completion of each step and enter the requested information. Record the frequency prior to moving the tuners. Complete the range test and loop test for each cavity. Record requested data in the tables in steps 41 and 42. If either tuner fails to meet the +/- 500 kHz range, write an NCR. If either tuner has hysteresis greater than 100 Hz, write an NCR. If either tuner has a resolution greater than 2 Hz / Step, write an NCR. |  |  |  | [[StepperTech]] <<SRF>> [[StepperTestTime]] <<TIMESTAMP>> [[StepperComments]] <<COMMENT>> |  |  |
| Cavity | Operator | Time of Test | Start Frequency (MHz) | Min Frequency (MHz) | Low Frequency Limit Activated? | Max Frequency (MHz) | High Frequency Limit Activated? |
| 7 | [[C7TunerOperator]] <<SRF>> | [[C7TunerTime]] <<TIMESTAMP>> | [[C7StartFreq]] <<FLOAT>> | [[C7MinFreq]] <<FLOAT>> | [[C7LowLimit]] <<YESNO>> | [[C7MaxFreq]] <<FLOAT>> | [[C7HiLimit]] <<YESNO>> |
| 8 | [[C8TunerOperator]] <<SRF>> | [[C8TunerTime]] <<TIMESTAMP>> | [[C8StartFreq]] <<FLOAT>> | [[C8MinFreq]] <<FLOAT>> | [[C8LowLimit]] <<YESNO>> | [[C8MaxFreq]] <<FLOAT>> | [[C8HiLimit]] <<YESNO>> |

| Step No | Instructions |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| 79 | Record the requested information from the Mechanical Tuner Tests in the table below. Upload the post-processed Excel files. |  |  |  |  |
| Cavity | Tuner Range (kHz) | Hz / Steps | Hysteresis (Hz) | Final Tuned Frequency (MHz) | Tuner File Uploads |
| 7 | [[C7TunerRange]] <<FLOAT>> | [[C7TunerHzPerStep]] <<FLOAT>> | [[C7TunerHysteresis]] <<FLOAT>> | [[C7TunerFinalFreq]] <<FLOAT>> | [[C7TunerFiles]] <<FILEUPLOAD>> |
| 8 | [[C8TunerRange]] <<FLOAT>> | [[C8TunerHzPerStep]] <<FLOAT>> | [[C8TunerHysteresis]] <<FLOAT>> | [[C8TunerFinalFreq]] <<FLOAT>> | [[C8TunerFiles]] <<FILEUPLOAD>> |
| 80 | Use the Comment block to document any problems with the tuner tests. |  |  | [[TunerTestComments]] <<COMMENT>> |  |

| Cavity 7 Gradient and Field Emission at 2K |  |  |
| --- | --- | --- |
| Step No | Instructions | Data Inputs |
| 81 | Record QextFPC and QextFP for Cavity 7 and the gradient at which measurement was made (in MV/m). If QextFPC for either cavity does not meet the requirement of 2.0E7 +/- 20%, write an NCR. | [[C72KQextTech]] <<SRF>> [[C72KQextMeasTime]] <<TIMESTAMP>> [[C72KQextFPC]] <<SCINOT>> [[C72KQextFP]] <<SCINOT>> [[C72KQextMeasGradient]] <<FLOAT>> |
| 82 | Record the Maximum gradient (Emax) for Cavity 7 and the gradient limiting condition. | [[C72KEmaxTech]] <<SRF>> [[C72KEmaxMeasTime]] <<TIMESTAMP>> [[C72KEmax]] <<FLOAT>> (MV/m) [[C72KEmaxLimit]] {{Admin Limit,Quench,FE related,Arc Fault,Window Temp Fault,BL Vacuum Fault,WG Vacuum Fault,RF Power,Heat Load,End Group Quench}} <<SELECT>> |
| 83 | Record the gradient at which a successful One Hour Run was completed for Cavity 7. Upload spreadsheet containing data on the One Hour Run. If the One Hour Run Gradient is lower than 19.0 MV/m, write an NCR. | [[C72KOneHourRunTech]] <<SRF>> [[C72KOneHourRunTime]] <<TIMESTAMP>> [[C72KEmaxop]] <<FLOAT>> [[C72KOneHourRunFile]] <<FILEUPLOAD>> |
| 84 | Record Static Lorentz value for Cavity 7 calculated from data collected during the Field Emission Measurement gradient ramping. | [[C72KStaticLorentz]] <<FLOAT>> [[C72KStaticLorentzFile]] <<FILEUPLOAD>> |
| 85 | Record the Field Emission Onset gradient for Cavity 7. Upload the file containing Field Emission data. | [[C72KFETech]] <<SRF>> [[C72KFEMeasTime]] <<TIMESTAMP>> [[C72KFEOnset]] <<FLOAT>> (MV/m) [[C72KFEMaxDoseRate]] <<FLOAT>> (R/hr) [[C72KFEFile]] <<FILEUPLOAD>> |
| 86 | Record any information about the performance of Cavity 7 in the Comment block. | [[C72KEmaxTestComments]] <<COMMENT>> |

| Cavity 8 Gradient and Field Emission at 4K |  |  |
| --- | --- | --- |
| Step No | Instructions | Data Inputs |
| 87 | Record QextFPC and QextFP for Cavity 8 and the gradient at which measurement was made (in MV/m). If QextFPC for either cavity does not meet the requirement of 2.0E7 +/- 20%, write an NCR. | [[C82KQextTech]] <<SRF>> [[C82KQextMeasTime]] <<TIMESTAMP>> [[C82KQextFPC]] <<SCINOT>> [[C82KQextFP]] <<SCINOT>> [[C82KQextMeasGradient]] <<FLOAT>> |
| 88 | Record the Maximum gradient (Emax) for Cavity 8 and the gradient limiting condition. | [[C82KEmaxTech]] <<SRF>> [[C82KEmaxMeasTime]] <<TIMESTAMP>> [[C82KEmax]] <<FLOAT>> (MV/m) [[C82KEmaxLimit]] {{Admin Limit,Quench,FE related,Arc Fault,Window Temp Fault,BL Vacuum Fault,WG Vacuum Fault,RF Power,Heat Load,End Group Quench}} <<SELECT>> |
| 89 | Record the gradient at which a successful One Hour Run was completed for Cavity 8. Upload spreadsheet containing data on the One Hour Run. If the One Hour Run Gradient is lower than 8.0 MV/m, write an NCR. | [[C82KOneHourRunTech]] <<SRF>> [[C82KOneHourRunTime]] <<TIMESTAMP>> [[C82KEmaxop]] <<FLOAT>> [[C82KOneHourRunFile]] <<FILEUPLOAD>> |
| 90 | Record Static Lorentz value for Cavity 8 calculated from data collected during the Field Emission Measurement gradient ramping. | [[C82KStaticLorentz]] <<FLOAT>> [[C82KStaticLorentzFile]] <<FILEUPLOAD>> |
| 91 | Record the Field Emission Onset gradient for Cavity 8. Upload the file containing Field Emission data. | [[C82KFETech]] <<SRF>> [[C82KFEMeasTime]] <<TIMESTAMP>> [[C82KFEOnset]] <<FLOAT>> (MV/m) [[C82KFEMaxDoseRate]] <<FLOAT>> (R/hr) [[C82KFEFile]] <<FILEUPLOAD>> |
| 92 | Record any information about the performance of Cavity 8 in the Comment block | [[C82KEmaxTestComments]] <<COMMENT>> |

| 30K Cryocycle 2 |  |  |
| --- | --- | --- |
| Step No | Instructions | Data Inputs |
| 93 | Lower Liquid Level to 20 % Record time at which JT valve is closed. Record Power Level for HeVessel Heater Record time at which liquid level reaches 20%. | [[NBSN30K2Operator]] <<SRF>> [[NBSN30K2StartTime]] <<TIMESTAMP>> [[NBSN30K2BoilOffHeaterLvl]] <<FLOAT>> [[NBSN30K2Lvl20Time]] << TIMESTAMP >> |
| 94 | Cryomodule Warm Up Heaters must be turned off when temperatures reach 25K. Record Heater Setting for warm up portion. Record time when liquid level reaches approximately 0%. Record time when all cavity temperature sensor read 25K or higher or when any single cavity temperature sensor reaches 45K. | [[NBSN30K2WarmUpHtrSetting]] <<FLOAT>> [[NBSN30K2ZeroLvlTime]] <<TIMESTAMP>> [[NBSN30K225KTime]] <<TIMESTAMP>> (MV/m) |
| 95 | Record any observations about the boil off and warm up portions of the cryocycle in the comment block. Attach any files relevant to Warm up portion of Cryocycle. | [[NBSN30K2WarmUpComments]] <<COMMENT>> [[NBSN30K2LogFiles]] <<FILEUPLOAD>> |
| 96 | Cryomodule Cool Down Record time at which JT valve is opened. Record time at which any cavity temperature reaches 4K. Record time at which liquid level reaches 90%. | [[NBSN30K2CoolDownStartTime]] <<TIMESTAMP>> [[NBSN30K2CoolDown4KTime]] <<TIMESTAMP>> [[NBSN30K290LvlTime]] <<FLOAT>> |
| 97 | Record any observations about the boil off and warm up portions of the cryocycle in the comment block. Attach any files relevant to Cool Down portion of Cryocycle. | [[NBSN30K2CoolDownComments]] <<COMMENT>> [[NBSN30K2CoolDownFiles]] <<FILEUPLOAD>> |

| Cavity 7 Q0 at 2K Part 2 |  |  |
| --- | --- | --- |
| Step No | Instructions | Data Inputs |
| 98 | Record QextFPC and QextFP for Cavity 7 and the gradient at which measurement was made (in MV/m). If QextFPC for either cavity does not meet the requirement of 2.0E7 +/- 20%, write an NCR.. | [[C72K2QextTech]] <<SRF>> [[C72K2QextMeasTime]] <<TIMESTAMP>> [[C72K2QextFPC]] <<SCINOT>> [[C72K2QextFP]] <<SCINOT>> [[C72K2QextMeasGradient]] <<FLOAT>> |
| 99 | After completing the Q0 measurement sequence for Cavity 7, record the values of Q0 at 10.0 MV/m, 19.0 MV/m and Emaxop. Upload the Q0 measurement file. If Q0 at 10.0 MV/m is less than 4.0E9, write an NCR. | [[C72K2QoTech]] <<SRF>> [[C72K2QoMeasTime]] <<TIMESTAMP>> [[C72K2QoTen]] <<SCINOT>> [[C72K2QoNineteen]] <<SCINOT>> [[C72K2QoEmaxop]] <<SCINOT>> [[C72K2QoFile]] <<FILEUPLOAD>> |
| 100 | Record any information about the performance of Cavity 7 in the Comment block. | [[C72K2QoTestComments]] <<COMMENT>> |

| Cavity 8 Q0 at 2K Part 2 |  |  |
| --- | --- | --- |
| Step No | Instructions | Data Inputs |
| 101 | Record QextFPC and QextFP for Cavity 8 and the gradient at which measurement was made (in MV/m). If QextFPC for either cavity does not meet the requirement of 2.0E7 +/- 20%, write an NCR. | [[C82K2QextTech]] <<SRF>> [[C82K2QextMeasTime]] <<TIMESTAMP>> [[C82K2QextFPC]] <<SCINOT>> [[C82K2QextFP]] <<SCINOT>> [[C82K2QextMeasGradient]] <<FLOAT>> |
| 102 | After completing the Q0 measurement sequence for Cavity 8, record the values of Q0 at 7 MV/m, 8.0 MV/m and at Emaxop. Upload the Q0 measurement file. If Q0 at 7.0 MV/m is less than 3.0E9, write an NCR. | [[C82K2QoTech]] <<SRF>> [[C82K2QoMeasTime]] <<TIMESTAMP>> [[C82K2QoSeven]] <<SCINOT>> [[C82K2QoEight]] <<SCINOT>> [[C82K2QoEmaxop]] <<SCINOT>> [[C82K2QoFile]] <<FILEUPLOAD>> |
| 103 | Record any information about the Q0 performance of Cavity 8 in the Comment block. | [[C82K2QoTestComments]] <<COMMENT>> |

| Microphonics |  |  |
| --- | --- | --- |
| Step No | Instructions | Data Inputs |
| 104 | Upload any microphonics data files below: If microphonics detuning is greater than 30 Hz, generate an NCR. |  |
| Cavity | Operator / Time | File Upload |
| 7 |  | [[C7MicrophonicsTech]] <<SRF>> [[C7MicrophonicsMeasTime]] <<TIMESTAMP>> [[C7MicrophonicsFile]] <<FILEUPLOAD>> |
| 8 |  | [[C8MicrophonicsTech]] <<SRF>> [[C8MicrophonicsMeasTime]] <<TIMESTAMP>> [[C8MicrophonicsFile]] <<FILEUPLOAD>> |
| 105 | Enter any comments relevant to microphonics measurements in the Comment block. | [[MicrophonicsComments]] <<COMMENT>> |

| Step No | Instructions | Data Inputs |
| --- | --- | --- |
| 106 | Record the Average Static Heat Load (in Watts) to the primary (2K) helium circuit. Enter any requested information to the right. | [[PrimaryCktStaticHeatLoad]] <<FLOAT>> (W) [[PrimaryCktStaticHeatLoadFile]] <<FILEUPLOAD>> |
| 107 | Measure the Shield Static Heat Load. Calculate using data acquired at least 1 week after pump down to 2K is completed. Enter the Static Heat Load (in Watts) in the appropriate box. | [[ShieldCktStaticHeatLoad]] <<FLOAT>> (W) [[ShieldCktStaticHeatLoadFile]] <<FILEUPLOAD>> |

| Step No | Instructions | Data Inputs |
| --- | --- | --- |
| 108 | Detune all eight cavities prior to the start of warm up. Use comment block to list any issues. | [[DetuningOperator]] <<SRF>> [[DetuningTime]] <<TIMESTAMP>> [[C7Detuned]] <<CHECKBOX>> [[C8Detuned]] <<CHECKBOX>> [[DetuningComments]] <<COMMENT>> |
| 109 | Record the beamline vacuum, all waveguide vacuums and the insulating vacuum (in torr) prior to beginning warm up. This should be done when HPRF has been off for at least two hours. | [[FinalVacInspector]] <<SRF>> [[FinalVacInspectTime]] <<TIMESTAMP>> [[FinalBLvac]] <<SCINOT>> [[FinalC7WGVac]] <<SCINOT>> [[FinalC8WGVac]] <<SCINOT>> [[FinalInsVac]] <<SCINOT>> |
| 110 | Begin the cryomodule warm up procedure. Record the start time for the warm up. | [[WarmUpOperator]] <<SRF>> [[WarmUpStartTime]] <<TIMESTAMP>> |
| 111 | Record the start time for the U-tube removal procedure. | [[UTubeTech]] <<SRF>> [[UTubeRemoveStartTime]] <<TIMESTAMP>> |
| 112 | Use the comment box to list any problems associated with the warm up. | [[NBSNCMWarmUpComment]] <<COMMENT>> |
| 113 | Upload the logfiles containing cryomodule warm up data. | [[NBSNCMWarmupFiles]] <<FILEUPLOAD>> |