| Traveler Title | All metal right angle valve leak check traveler for L2HE BLA |  |  |  |
| --- | --- | --- | --- | --- |
| Traveler Abstract | Leak check traveler |  |  |  |
| Traveler ID | L2HE-CLNRM-BLAVV-LEAK |  |  |  |
| Traveler Revision | R1 |  |  |  |
| Traveler Author | A. Grabowski |  |  |  |
| Traveler Date | 18-Jan-24 |  |  |  |
| NCR Informative Emails | D.Forehand, |  |  |  |
| NCR Dispositioners | A. Grabowski, G. Cheng |  |  |  |
| D3 Emails | D.Forehand, A. Grabowski, G. Cheng |  |  |  |
| Approval Names | A.Grabowski | D. Forehand | M. Bevins |  |
| Approval Signatures |  |  |  |  |
| Approval Dates |  |  |  |  |
| Approval Title | Author | Reviewer | Project Manager |  |

| References | List and Hyperlink all documents related to this traveler. This includes, but is not limited to: safety (THAs, SOPs, etc), drawings, procedures, and facility related documents. |  |  |  |
| --- | --- | --- | --- | --- |

| Revision Note |  |
| --- | --- |
| R1 | Initial release of this Traveler. |

| Step No | Instructions | Data Input |
| --- | --- | --- |
| 1 | Incoming leak check. The L2HE project uses the VAT all metal right angle vacuum valve in many different areas. BLA assembly will require one of these valves. Because the valve in the BLA assembly will remain in the unit permanently, the leak tightness is critical. Therefore, we leak check the valve before the assembly. - This leak test will be performed by Inficon leak detector in the tuning room. - Wear gloves while handling vacuum valves. - Connect vacuum pump as shown in the figure. - Check the valve and confirm the valve is closed. Use the valve as received. - Start pump down for leak check. If pump down rate is too slow, tighten the valve by a quarter turn and record the torque value. - Start leak test. - Leak check by spraying all joints as shown in figure. Enter leak check date and name. Select the valve serial number Is the unit leak tight? Upload the RGA data Please provide comment if any. | [[BLAVV_DATE]] <<TIMESTAMP>> [[BLAVV_Operator]] <<SRF>> [[BLAVVSN]] <<BLAVVSN>> [[Leakcheck_Pass]] <<YESNO>> [[BLAVV_RGA]] <<FILEUPLOAD>> [[BLAVV_Comment]] <<COMMENT>> |
| 2 | Review by responsible engineer. After review select YES or NO. If valve is YES send to chemistry for metal valve cleaning and baking, chemroom will send into pass thru ready for use. If valve is NO contact supervisior for next step. | [[Reviewer]] <<YESNO>> [[HP1]] {{Grabowski}}<<HOLDPOINT>> |