## Sheet1
NCR SEQ NUM | NCR OPEN DATE | NCR CLOSE DATE | TRAVELER ID | TRAV SEQ NUM | PART DESCRIPTION | SN | CATEGORY | DESCRIBE COMMENT | INITIAL DISPOSITION | INITIAL COMMENT | FINAL DISPOSITION | FINAL COMMENT | Notes
1 | Mar-09-2021 | Mar-16-2021 | L2HE-INSP-FPCWG | 1 |  | 004 | Scratches | Waveguide coupler flange (Dwg # SA-375-600-90) has multiple scratches present on flange. Please see attached. Multiple scratched areas checked w/profilometer. The worst area found was Rz (peak to valley) measured 294 u inch and RA (surface finish) measured 49 u inch. | Return To Vendor | Complete remainder of inspection traveler before returning to vendor
2 | May-04-2021 | Oct-06-2021 | L2HE-VTA-CAV-VTRF | 1 | SINGLE CELL CAVITY | 0331 | VTARFPerformance | Q0(20.8) is 2.4e10, slightly lower than spec which is 2.5e10 | Use As Is
3 | May-05-2021 | Oct-06-2021 | L2HE-VTA-CAV-VTRF | 2 | UNKN | R009 | VTARFPerformance | Cavity did not pass due to FE up to 1R/hr. | Use As Is
4 | May-18-2021 | Aug-10-2021 | L2HE-VTA-CAV-VTRF | 3 | UNKN | 0040 | VTARFPerformance | Cavity frequency was low at 1299.975 MHz | Hold/Re-measure | re-test | Reject
5 | May-18-2021 | Aug-10-2021 | L2HE-VTA-CAV-VTRF | 3 | UNKN | 0040 | N/A | Operator error except for two decay measurements the test was done at the 8pi/9 mode. For those measurements, Qo=2.9e10, Qfpc=1.94e10,Qfp=3.4e11 E = 2.9 MV/m | Reject | re-rinse/re-test did not meet requirements
7 | May-19-2021 | Aug-17-2021 | L2HE-INSP-FPFT | 73 | PICKUP FEEDTHROUGH | 2061 | Stains | Probe has stain | Hold/Modify | Send to chemistry, see if the stain can be removed. | Use As Is | Passed visual inspection after stain removal. Okay to use.
8 | May-19-2021 | Aug-17-2021 | L2HE-INSP-FPFT | 74 |  | 2062 | Scratches | Scratches down outside diameter. | Use As Is | The scratches are not significant. Pictures are captured and okay to proceed to use.
9 | May-25-2021 | Jun-25-2021 | L2HE-INSP-HMFT | 158 |  | 2134 | Dings Roughness | HEH2134 has multiple dings down Antenna. | Use As Is | Surface finish poor but field in this area is not very strong either. Visual findings can be used to relate to cavity performance later on.
10 | May-25-2021 | Jul-28-2021 | L2HE-INSP-HMFT | 167 |  | 2143 | OTHER | HEH2143 probe antenna has metallic foreign material on face. | Hold/Modify | See if chemistry will be able to remove the foreign material or residue. | Use As Is
11 | May-26-2021 | Aug-10-2021 | L2HE-VTA-CAV-VTRF | 4 | PRODUCTION CAVITY ASSEMBLY | R009 | N/A | Qext2 is 2.9200e+11, out of spec which is 0.75-2.5 e12. | Use As Is
12 | May-26-2021 | Aug-10-2021 | L2HE-VTA-CAV-VTRF | 4 |  | HE-R009 | N/A | QextHOMb is 1.11e+11, out of spec which is >=2.7e11. Also, PHOMB at 21MVm was 3.52 W, out of spec which is ? 1.7 W | Hold/Re-measure | retune HOM | Use As Is
13 | May-27-2021 | Jun-25-2021 | L2HE-INSP-FPFT | 90 |  | 2078 | Dings Roughness | Probe has multiple dings down along stem. | Use As Is
14 | May-27-2021 | Aug-17-2021 | L2HE-INSP-FPFT | 94 |  | 2082 | OTHER | Surface seal area has small imperfection. | Hold/Modify | Send to machine shop to remove the irregularity on seal path. | Use As Is | Passed visual inspection after sealing area was polished. Okay to proceed to use.
15 | May-27-2021 | Aug-17-2021 | L2HE-INSP-FPFT | 95 |  | 2083 | OTHER | Probe has gouge on O.D | Use As Is | The gouge does not appear significant. Pictures captured as reference, and okay to proceed to use.
16 | May-27-2021 | Jun-25-2021 | L2HE-INSP-FPFT | 97 |  | 2085 | Roughness | Probe O.D all the way around has rough finish. | Use As Is
17 | Jun-02-2021 | Sep-09-2021 | L2HE-INSP-FPFT | 155 | PICKUP FEEDTHROUGH | 2143 | OTHER | Probe has line of rough imperfections. | Use As Is | Imperfection is not significant. Probe pictures are captured and available for reference if there is future issue.
18 | Jun-02-2021 | Sep-09-2021 | L2HE-INSP-FPFT | 163 |  | 2151 | Roughness | Seal surface area has rough finish. | Use As Is | Seal surface is measured with profilometer, average Ra is 12 uin, within drawing requirement which is 64 uin (1.6 um).
19 | Jun-03-2021 | Jun-25-2021 | L2HE-INSP-FPFT | 169 |  | 2157 | Roughness | Probe has rough finish | Use As Is
20 | Jun-03-2021 | Jun-25-2021 | L2HE-INSP-FPFT | 170 |  | 2158 | Roughness | Probe has rough finish | Use As Is
21 | Jun-03-2021 | Sep-09-2021 | L2HE-INSP-HMFT | 259 |  | 2235 | Roughness | Seal surface area has rough finish. Probe 6mm radius has rough finish. | Use As Is | Seal surface measured with profilometer, average Ra is 22 uin, within drawing requirement which is 32 uin (0.8 um).
22 | Jun-03-2021 | Sep-09-2021 | L2HE-INSP-HMFT | 262 | HOM FEEDTHROUGH | 2238 | Roughness | Seal surface area is not smooth. Probe radius has rough finish. | Use As Is | Seal surface measured with profilometer, average Ra 21 uin, within drawing requirement which is 32 uin (0.8 um). Probe roughness is okay to use as is. Probe pictures are captured as reference if there is issue later on.
23 | Jun-04-2021 | Dec-07-2021 | L2HE-INSP-CAV | 7 | PRODUCTION CAVITY ASSEMBLY | 0284 | OTHER | Visual inspection of cavity bellows on SN CAV-284 found (3) minor dents in between bellow convolutions. See pictures in attach box. | Reject | Cavity was shipped to FNAL and is no longer & JLab.
24 | Jun-08-2021 | Sep-24-2021 | L2HE-INSP-HMFT | 73 |  | 2049 | Dimensional | Flange seal surface to probe tip distance measured 29.50 mm, slightly out of spec which is 29.51-29.61 mm. | Hold/Re-measure | Hold for now, further evaluation will be done. | Use As Is | Per communication with SLAC, the antenna being 0.01 mm short is not likely to cause a problem in this case.
26 | Jun-15-2021 | Sep-07-2021 | L2HE-INSP-FPFT | 60 |  | 2048 | OTHER | Probe is loose. | Reject | Set aside and do not use for production.
27 | Jun-18-2021 | Sep-09-2021 | L2HE-INSP-FPFT | 150 | PICKUP FEEDTHROUGH | 2138 | Roughness | Seal surface finish is not smooth | Use As Is | Seal surface measured with profilometer, average Ra 12 uin, within drawing requirement which is 64 uin (1.6 um).
28 | Jun-24-2021 | Jun-25-2021 | L2HE-INSP-CAV | 11 | PRODUCTION CAVITY ASSEMBLY | 0145 | OTHER | Visual inspection of bellow convolutions found (2) small dents in bottom of convolution. | Use As Is |  | Use As Is
29 | Jun-24-2021 | Jun-25-2021 | L2HE-INSP-CAV | 13 | PRODUCTION CAVITY ASSEMBLY | 0118 | OTHER | Visual inspection of bellow convolution only,found (1) small dent on bottom of bellow convolution. | Use As Is
30 | Jul-06-2021 | Aug-11-2021 | L2HE-CLNRM-FPCC-INSP | 1 | UNKN | CPI-FPCC-HE-001 | Dings | found small ding in bellows of CPI-001 coupler | Return To Vendor
31 | Jul-07-2021 | Jul-07-2021 | L2HE-INSP-CAV | 16 | PRODUCTION CAVITY ASSEMBLY | 0342 | OTHER | Testing Only! | Hold/Modify | TESTING ONLY | Conforming
32 | Jul-12-2021 | Aug-10-2021 | L2HE-CMA-UCM-INSP | 1 | UPPER COLD MASS ASSEMBLY | RFCHE-UC-002 | OTHER | Drawing part # F10074420 Vertical support lines twisted. Please see attached. |  |  | Use As Is | Checked with Fermi. Rods can be used as is. See attached transcript
33 | Jul-13-2021 | Aug-10-2021 | L2HE-CMA-UCM-INSP | 1 | UPPER COLD MASS ASSEMBLY | RFCHE-UC-002 | OTHER | Copper clamp indium foil applied to wrong half. | Hold/Modify | 1. Check/disassemble all 16 copper clamps. 2. Apply indium foil to the long halves on the clamps. No need to remove the indium on the short half. 3. Re-assemble clamps, torque three times with 8-hr intervals. | Conforming
34 | Jul-13-2021 | Aug-10-2021 | L2HE-CMA-UCM-INSP | 1 | UPPER COLD MASS ASSEMBLY | RFCHE-UC-002 | OTHER | Cavity hangar has sharp edges. | Hold/Modify | Deburr sharp edges | Conforming
35 | Jul-13-2021 | Aug-10-2021 | L2HE-CMA-UCM-INSP | 1 | UPPER COLD MASS ASSEMBLY | RFCHE-UC-002 | OTHER | During the inspection, we came across bolts that had suspicious head markings.We had M16 and 1/4-20 hardware that had "vvv" on the head instead of "THE". The proper authorities were notified. |  |  | Use As Is | Vendor supplied material certs per JLAB request, see attached. Certs look good. JLAB also run metal analysis on a sample screw and the meterial is found to be 316 SS. JLAB consulted Fermi on this issue. Fermi's comments are "Graded fasteners are required to have specific markings, much like lifting hardware. Non-graded fasteners which are required to follow some standard (e.g. ASTM) will also have markings, and it’s common industry practice to have markings. That being said, not all fasteners are required to have markings", "we didn’t specify any graded fasteners...nor invoke any fastener codes for the coldmass upper assembly".
36 | Jul-13-2021 | Aug-11-2021 | L2HE-CWI-VV | 1 | VESSEL AFTER MACHING WELDMENT | RFCHE-VV-002 | N/A | The fillet weld located on weld#C1 have been reduced in size via some machining operation. The welds are not complaint with the drawing, specifications or the welding code. It is my intention to reject the fillet weld portion of this flange weld joint that also includes a CJP groove weld. However, I found the groove weld portion of this same flange weld joint to be in compliance. |  |  | Use As Is | After consulting Fermi on 8/11, decision has been made to accept VV002 welds as is.
37 | Jul-13-2021 | Oct-06-2021 | L2HE-VTA-CAV-VTRF | 8 | PRODUCTION CAVITY ASSEMBLY | 0331 | N/A | Cavity quench at 21.8 MV/m | Hold/Re-measure | We will look into re-rinse/re-test for this cavity | Use As Is
38 | Jul-31-2021 | Jul-31-2021 | L2HE-VTA-CAV-VTRF | 9 | PRODUCTION CAVITY ASSEMBLY | HE-331 | N/A | FE onset 5.5 MV/m, FE loading and 1R/hr at 14 MV/m | Hold/Re-measure | re-rinse https://bit.ly/3jbEoiG | Reject | Wait for HPR evaluation Re-test
39 | Aug-02-2021 | Aug-10-2021 | L2HE-CMA-FPCW-INSP | 1 | FPC WARM PART | CPI-FPCW-HE-001 | Scratches | Found scratch on copper plating of DN40 flange |  |  | Return To Vendor
40 | Aug-02-2021 | Aug-13-2021 | L2HE-CMA-FPCW-INSP | 2 | FPC WARM PART | CPI-FPCW-HE-002 | Scratches | Found scratch on copper plating of DN40 flange |  |  | Return To Vendor
41 | Aug-02-2021 | Aug-13-2021 | L2HE-CMA-FPCW-INSP | 3 | FPC WARM PART | CPI-FPCW-HE-007 | Scratches | Found scratches on copper plating of DN40 flange |  |  | Return To Vendor
42 | Aug-02-2021 | Aug-13-2021 | L2HE-CMA-FPCW-INSP | 5 | FPC WARM PART | CPI-FPCW-HE-005 | Plating Scratches | Found scratch on copper plating of DN40 flange |  |  | Return To Vendor
43 | Aug-02-2021 | Aug-13-2021 | L2HE-CMA-FPCW-INSP | 6 | FPC WARM PART | CPI-FPCW-HE-006 | Mechanical | Push rod bellows distorted and slightly twisted. Push rod was difficult to remove and reinstall. The M6 antenna vented bolt was loose |  |  | Return To Vendor
44 | Aug-02-2021 | Aug-13-2021 | L2HE-CMA-FPCW-INSP | 7 | FPC WARM PART | CPI-FPCW-HE-009 | Scratches | Face if DN100 flange scratched, small scratch on outer conductor copper plating. |  |  | Use As Is
45 | Aug-02-2021 | Aug-13-2021 | L2HE-CMA-FPCW-INSP | 8 | FPC WARM PART | CPI-FPCW-HE-010 | Mechanical | Push rod bellows very bent and twisted, some convolutions were pinched. Push rod was not engaged. |  |  | Return To Vendor
46 | Aug-02-2021 | Aug-13-2021 | L2HE-CMA-FPCW-INSP | 8 | FPC WARM PART | CPI-FPCW-HE-010 | Plating | RF contact area of inner conductor copper was pitted and copper plating had defects |  |  | Return To Vendor
47 | Aug-10-2021 | Sep-15-2021 | L2HE-INSP-IMAG | 24 | MAGNETIC SHIELD CAVITY ASSEMBLY | RFCHE-MS-028 | Dings | Part I.D.# F10048728 Rev.A Shield, LCLS-11 2nd layer center section A. has minor ding about 5/16" diameter and .010 deep. Part I.D.# F10048731 Shield, LCLS-11 2nd layer coupler side A. Has (2) sets of imperfections (1)is about 6" long the 2nd is about 4" long . | Hold/Re-measure | SN 28 is selected to carry out on-axis field mapping to see if the shielding performance is compromised by the imperfections | Use As Is | The shield's performance is not compromised hence it can be used as is
48 | Aug-10-2021 | Sep-15-2021 | L2HE-INSP-IMAG | 26 | MAGNETIC SHIELD CAVITY ASSEMBLY | RFCHE-MS-026 | Dings | Part I.D.# F10048728 Rev.A Shield, LCLS-11 2nd layer center section A. Has light ding 3/8" diameter and .010 deep. Part I.D.# F10048731 Rev.A Shield,LCLS-11 2nd layer coupler side A. Has 6" long imperfection. marked in red. | Hold/Re-measure | SN 26 is selected to carry out on-axis field mapping to see if the shielding performance is compromised by the imperfections | Use As Is | Shelding performance is not compromised so the shield can be used as is
49 | Aug-10-2021 | Sep-24-2021 | L2HE-INSP-HMFT | 186 | HOM FEEDTHROUGH | 2161 | Dimensional | Flange seal surface to probe tip is 29.5 mm, slightly out of spec which is 29.51-29.61 mm | Hold/Re-measure | Hold for now. Further evaluation is needed. | Use As Is | Per communication with SLAC, the antennas being 0.01 mm short is not likely to cause a problem in this case.
50 | Aug-12-2021 | Aug-31-2021 | L2HE-INSP-IMAG | 32 | MAGNETIC SHIELD CAVITY ASSEMBLY | RFCHE-MS-046 | OTHER | Damage noted during visual inspection. Part ID#s F10048802 & F10048803 has a cut through the MAG Cavity turner end. See attached photos. Normal shield shown. |  |  | Conforming | This shield, SN 46, is determined to be a cavity 6 magnetic shield that suppose to have the cut
51 | Aug-13-2021 | Aug-13-2021 | L2HE-CMA-FPCW-INSP | 8 | FPC WARM PART | CPI-FPCW-HE-010 | Mechanical | One of the RF spring had come undone from itself |  |  | Return To Vendor
52 | Aug-13-2021 | Aug-31-2021 | L2HE-CMA-FPCW-INSP | 5 | FPC WARM PART | CPI-FPCW-HE-005 | OTHER | Kapton sleeve missing covering the ceramic. | Hold/Modify | Install spare kapton sleeve prior to storage | Return To Vendor
53 | Aug-13-2021 | Aug-31-2021 | L2HE-CMA-FPCW-INSP | 6 | FPC WARM PART | CPI-FPCW-HE-006 | OTHER | Kapton sleeve missing covering the ceramic. | Hold/Modify | Install spare kapton sleeve prior to storage | Return To Vendor
54 | Aug-13-2021 | Aug-31-2021 | L2HE-CMA-FPCW-INSP | 7 | FPC WARM PART | CPI-FPCW-HE-009 | OTHER | Kapton sleeve missing covering ceramic. | Hold/Modify | Install spare kapton sleeve prior to storage | Return To Vendor
55 | Aug-13-2021 | Aug-31-2021 | L2HE-CMA-FPCW-INSP | 8 | FPC WARM PART | CPI-FPCW-HE-010 | OTHER | Kapton sleeve missing covering the ceramic. | Hold/Modify | Install spare kapton sleeve prior to storage | Return To Vendor
56 | Aug-19-2021 | Sep-24-2024 | L2HE-VTA-CAV-VTRF | 10 | PRODUCTION CAVITY ASSEMBLY | HE-040 | N/A | The lock frequency was below specification. It was 1299.966 MHz. |  |  | Reject | 2021-08-24 CTB decision: put cavity aside, no further treatment/testing
57 | Aug-27-2021 | Oct-19-2021 | L2HE-CMA-VV-INSP | 2 | VESSEL AFTER MACHING WELDMENT | RFCHE-VV-004 | OTHER | Paint present inside of groove of flange and on face of flanges. Drawing number from picture identifies flange. Please see attached. | Hold/Modify | Having discussed with cryomodule assembly team lead, we will review the visual inspection findings and make corrections as needed. Vendor and Fermi are noted of the issues for improvements on future vessels. | Conforming | Repair work has been done. We will keep track of this VV to see if leak would develop at final assembly level.
58 | Aug-27-2021 | Oct-19-2021 | L2HE-CMA-VV-INSP | 2 | VESSEL AFTER MACHING WELDMENT | RFCHE-VV-004 | Scratches | Scratches and pitting present on multiple flanges. Please see attached pictures. Drawing number from pictures identify flange with issue. | Hold/Modify | Having discussed with cryomodule assembly team lead, we will review the visual inspection findings and make corrections as needed. Vendor and Fermi are noted of the issues for improvements on future vessels. | Conforming | Repair work has been done. We will keep track of this VV to see if leak would develop at final assembly level.
59 | Sep-02-2021 | Sep-03-2021 | L2HE-VTA-CAV-VTRF | 12 | PRODUCTION CAVITY ASSEMBLY | HE-295 | N/A | Cavity lock frequency 50 kHz below target range |  |  | Use As Is
60 | Sep-02-2021 | Sep-03-2021 | L2HE-VTA-CAV-VTRF | 12 | PRODUCTION CAVITY ASSEMBLY | HE-295 | N/A | Qext 1 slightly above allowable range, Qext2 slightly below allowable range (per Traveler criteria) |  |  | Use As Is
61 | Sep-08-2021 | Sep-17-2021 | L2HE-INSP-BLBU | 5 | WELDMENT BELLOWS - US END | 007 | Dimensional | Dim2 & Dim9 are OOT. See attached IR. Dim9 is multiplied (x2) to get the 96.0 dimension. |  |  | Use As Is | OOT values are for bellows only, are relatively small, and will not effect the flange features.
62 | Sep-09-2021 | Sep-17-2021 | L2HE-INSP-BLBU | 1 | WELDMENT BELLOWS - US END | 001 | Dimensional | Dim2, Dim3, & Dim9 are OOT. Dim9 is multiplied (x2) to get the 96.0 dimension. |  |  | Use As Is | OOT values are for bellows only, are relatively small, and will not effect the flange features.
63 | Sep-09-2021 | Sep-17-2021 | L2HE-INSP-BLBU | 4 | WELDMENT BELLOWS - US END | 004 | Dimensional | Dim2 is OOT. See attached report. |  |  | Use As Is | OOT values are for bellows only, are relatively small, and will not effect the flange features.
64 | Sep-09-2021 | Sep-17-2021 | L2HE-INSP-BLBU | 6 | WELDMENT BELLOWS - US END | 010 | Dimensional | Dim2 & Dim3 are OOT. See attached IR. |  |  | Use As Is | OOT values are for bellows only, are relatively small, and will not effect the flange features.
65 | Sep-14-2021 | Sep-16-2021 | L2HE-VTA-CAV-VTRF | 14 |  | 0128 | N/A | Cavity HE-CAV128 Qext2 was low during the decay measurement. |  |  | Use As Is
66 | Sep-15-2021 | Sep-21-2021 | L2HE-CWI-VV | 2 | VESSEL AFTER MACHING WELDMENT | RFCHE-VV-004 | OTHER | The internal CJP groove weld on Weld Joint #C1 is acceptable (*see weld-map). However, the external flange intermittent fillet welds were reduced in leg size. That reduction in fillet weld size is below an acceptable size in accordance with the drawings and specifications. |  |  | Use As Is | Having consulted Fermi Lab SOTR, this weld NCR is deemed non-critical
67 | Sep-16-2021 | Sep-20-2021 | L2HE-CMA-UCM-INSP | 2 | UPPER COLD MASS ASSEMBLY | RFCHE-UC-004 | OTHER | F10074420 Line D vertical supports twisted/bent. Please see attached picture. |  |  | Use As Is | On UC002, similar issues are found on Line D supports. Fermi and the vendor are both noted. Fermi advises to use as is. The same approach is adopted for UC004
68 | Sep-16-2021 | Sep-21-2021 | L2HE-CMA-UCM-INSP | 2 | UPPER COLD MASS ASSEMBLY | RFCHE-UC-004 | OTHER | Cold mass assembly is missing two F10043365 5k clamps. | Hold/Re-measure | Suspect that cryomodule assembly group took a couple 5K clamps off to work on re-application of indium foils, let's double check | Conforming
69 | Sep-21-2021 | Nov-03-2021 | L2HE-CMA-UCM-INSP | 2 | UPPER COLD MASS ASSEMBLY | RFCHE-UC-004 | OTHER | Indium foil applied to wrong half. | Hold/Modify | 1. Check/disassemble all 16 copper clamps. 2. Apply indium foil to the long halves on the clamps. No need to remove the indium on the short half. 3. Re-assemble clamps, torque three times with 8-hr intervals. | Conforming | ready for use
70 | Sep-27-2021 | Oct-19-2021 | L2HE-CMA-VV-INSP | 1 | VESSEL AFTER MACHING WELDMENT | RFCHE-VV-002 | Dings Roughness Scratches | Visual inspection performed on 17-Sep-2021. All photos on the server: M:\cmm\Pictures\Inspection Pictures\L2HE\Vacuum vessel\VV002\Receiving Inspection | Hold/Modify | Having discussed with cryomodule assembly team lead, we will review the visual inspection findings and make corrections as needed. Vendor and Fermi are noted of the issues for improvements on future vessels. | Conforming | Repair work has been done. We will keep track of this VV to see if leak would develop at final assembly level.
71 | Oct-17-2021 | Oct-18-2021 | L2HE-VTA-CAV-VTRF | 16 |  | HE-331 | N/A | Pi mode frequency 1300.1297 MHz is lower than min frequency specifications (1300.150MHz). |  |  | Use As Is | The acceptance criteria has been updated. This will be reflected in the future traveler.
72 | Oct-20-2021 | Mar-01-2022 | L2HE-INSP-CFL | 1 | CURRENT LEADS CF FLANGE ASSEMBLY | 0043 | Dimensional Instrumentation OTHER | All CL Flange lead wires are 12 inches, not 15 inches. According to FNAL Drawing F10022642-Rev G, under Table 1, Item 10, part number FC0062680 should be 15 inches long. Every CL Flange showed up with 12 inch long lead wires instead of the proper 15 inch long lead wires. | Hold/Modify | Fermi Lab will supply 15 inches long lead wires to replace all 12 inches long wires in the CL flange assemblies we received directly from the vendor. J lab technical staff will inspect the new wires and replace the shorter lead wires to qualify the part with serial no #0043 for further use in the corresponding CM assembly. | Conforming | The rework results confirm that the replacement lead wires meet the dimensional specification.
73 | Oct-25-2021 | Dec-21-2021 | L2HE-INSP-BLBU | 1 | WELDMENT BELLOWS - US END | 001 | Dings Scratches | Slotted flange seal area has several small dings and light scratches on the bottom. | Hold/Modify | Deliver to machine shop for polishing of seal surfaces. | Conforming | Seal surfaces have been polished.
74 | Oct-25-2021 | Feb-17-2022 | L2HE-INSP-BLBU | 2 | WELDMENT BELLOWS - US END | 002 | Dings Scratches | Slotted flange has light scratches and small dings on bottom of seal surface. | Hold/Modify | Deliver to machine shop for polishing of seal surfaces. | Conforming | Scratches have been polished, seal surface is acceptable.
75 | Oct-25-2021 | Feb-17-2022 | L2HE-INSP-BLBU | 3 | WELDMENT BELLOWS - US END | 003 | Dings Scratches | Slotted flange has several light scratches and (2) small dings. | Hold/Modify | Deliver to machine shop for polishing of seal surfaces. | Conforming | Scratches have been polished, seal surface is acceptable.
76 | Oct-25-2021 | Dec-21-2021 | L2HE-INSP-BLBU | 4 | WELDMENT BELLOWS - US END | 004 | Scratches | Slotted flange has several light scratches on bottom of seal surface area. | Hold/Modify | Deliver to machine shop for polishing of seal surfaces. | Conforming | Seal surfaces have been polished.
77 | Nov-02-2021 | Nov-04-2021 | L2HE-INSP-BLBS | 1 | WELDMENT BELLOWS SHORT | 002 | Dimensional | Please see attached for dimensional non-conforming. |  |  | Use As Is | OOT values are for bellows only, are relatively small, and will not effect the flange features.
78 | Nov-03-2021 | Nov-04-2021 | L2HE-INSP-BLBP | 3 | WELDMENT BELLOWS PRCM | 001 | Dimensional | Please see attached for dimensions out of tolerance. |  |  | Use As Is | OOT values are for bellows only, are relatively small, and will not effect the flange features.
79 | Nov-05-2021 | Nov-05-2021 | L2HE-INSP-BLBS | 2 | WELDMENT BELLOWS SHORT | 014 | Dimensional | Please see attached for dimensional non-conforming. |  |  | Use As Is | OOT values are relatively small, and will not effect the flange features.
80 | Nov-08-2021 | Nov-09-2021 | L2HE-VTA-CAV-VTRF | 22 | LCLS-II PRODUCTION CAVITY ASSEMBLY | 0351 | N/A | Qext1 was 8.78E+9, out of spec (1.1e10? |  |  | Use As Is | If I averaged the test data correctly, we are actually at an Qext1 of 9.08e9. That is slightly too low wrt spec but still acceptable.
81 | Nov-08-2021 | Nov-09-2021 | L2HE-VTA-CAV-VTRF | 22 | LCLS-II PRODUCTION CAVITY ASSEMBLY | 0351 | N/A | There was radiation from 22 MV/m until quench. | Hold/Re-measure | This cavity will be remeasured as part of the test of dewar 5. Furthermore, it should be re-rinsed & re-tested wrt the occurrence of radiation. | Reject
82 | Nov-08-2021 | Nov-09-2021 | L2HE-VTA-CAV-VTRF | 23 | PRODUCTION CAVITY ASSEMBLY | HE-331 | N/A | Qext1 was 2.56E+10, out of spec (1.1e10? |  |  | Reject | The deviation of Qext1 is known. Cavity is rejected for other reasons too.
83 | Nov-08-2021 | Nov-09-2021 | L2HE-VTA-CAV-VTRF | 23 | PRODUCTION CAVITY ASSEMBLY | HE-331 | N/A | Qo at 20.8+/-0.1 MV/m was 2.04e10, out of spec (>/= 2.5 e10). FE onset at 18 MV/m. |  |  | Reject | Cavity is rejected & put on hold.
84 | Nov-08-2021 | Nov-09-2021 | L2HE-VTA-CAV-VTRF | 24 | PRODUCTION CAVITY ASSEMBLY | HE-295 | N/A | Qext2 was 5.01E+11, out of spec (>/= 7 e11) | Hold/Re-measure | Cavity will be retested as part of dewar 5 test | Reject | This value was better in earlier tests. The cavity will be retested at a different position in dewar 5 as part of the qualification.
85 | Nov-09-2021 |  | L2HE-CMA-VV-INSP | 3 | VESSEL AFTER MACHING WELDMENT | RFCHE-VV-006 | Dings Residues Roughness Scratches ShippingDamage | NCR generated for Non-Conformances on the attached F10026610 VV-006 QC Task List. All items QC Task List that are "RED" are Non-Conforming and will need resolution. All photos can be found on server: M:\cmm\Pictures\Inspection Pictures\L2HE\Vacuum vessel\VV006 | Hold/Modify | Having discussed with cryomodule assembly team lead, CMA will review the visual inspection findings and polish surfaces with imperfections. The damaged tube will be rounded when this VV is moved to the test lab to be ready for use on a cryomodule.
86 | Nov-10-2021 | Nov-12-2021 | L2HE-VTA-CAV-VTRF | 25 | PRODUCTION CAVITY ASSEMBLY | HE-295 | N/A | Qext2 was 5.8e11, out of spec (>/= 7 e11 ) |  |  | Use As Is | Accepted by CTB
87 | Nov-10-2021 | Nov-12-2021 | L2HE-VTA-CAV-VTRF | 26 | LCLS-II PRODUCTION CAVITY ASSEMBLY | 0351 | N/A | Qext1 was 7.2e9, out of spec (1.1e10? | Hold/Re-measure | cavity needs to be retested, this deviation can be accepted | Use As Is
88 | Nov-10-2021 | Sep-24-2024 | L2HE-VTA-CAV-VTRF | 26 | LCLS-II PRODUCTION CAVITY ASSEMBLY | 0351 | N/A | FE onset 21.8 MV/m |  |  | Return To Vendor | 2024-09-24 sent to FNAL after big HPR for test in VTS
88 | Nov-10-2021 | Sep-24-2024 | L2HE-VTA-CAV-VTRF | 26 | LCLS-II PRODUCTION CAVITY ASSEMBLY | 0351 | N/A | FE onset 21.8 MV/m |  |  | Return To Vendor | 2024-09-24 sent to FNAL after big HPR for test in VTS
89 | Nov-10-2021 | Nov-12-2021 | L2HE-VTA-CAV-VTRF | 27 | PRODUCTION CAVITY ASSEMBLY | HE-331 | N/A | Q0(20.8) was 1.97e10, out of spec (>/= 2.5 e10) |  |  | Reject | cavity is put on hold by CTB
90 | Nov-10-2021 | Nov-12-2021 | L2HE-VTA-CAV-VTRF | 27 | PRODUCTION CAVITY ASSEMBLY | HE-331 | N/A | FE onset 14 MV/m |  |  | Reject | cavity is put on hold by CTB
91 | Nov-11-2021 | Mar-04-2024 | L2HE-CMA-UCM-INSP | 3 | LCLS-II UPPER COLD MASS ASSEMBLY | RFCHE-UC-006 | OTHER | Vertical support line D twisted. Please see attached. |  |  | Use As Is | Similar bending in rods has been identified previously. Fermi's recommendation is to use them as is. Please refer to the attached email for details.
92 | Nov-15-2021 | Nov-18-2021 | L2HE-INSP-CAV | 17 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R016 | Dimensional | Cavity center line distance to header pipe out of tolerance 0.1 mm. Nominal 300.0 mm. Measured 298.4 mm |  |  | Use As Is | We can accept this small deviation of .1 mm
93 | Nov-15-2021 | Apr-26-2022 | L2HE-INSP-CAV | 17 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R016 | OTHER | After de-mag cycles Tee ss joint and bellow restraint bracket (fpc side) exceed allowable amount. SS joint is 0.9 Gauss above limit. Bellow restraint bracket is .07 Gauss above limit. | Hold/Re-measure | Cavity will proceed to clean room and vertical test. If results confirm significant deterioration of Q, another degaussing at the larger degausser should be conducted. | Use As Is | Cavity tested with good Q
94 | Nov-17-2021 | Nov-18-2021 | L2HE-INSP-CAV | 18 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R017 | Dimensional | Cavity center line distance to header pipe out of tolerance +0.1 mm. DIM1: Nominal 300.0 +0.5/-1.5 mm. Measured 300.6 mm |  |  | Use As Is | We can accept this small deviation of .1 mm.
95 | Nov-17-2021 | Apr-26-2022 | L2HE-INSP-CAV | 18 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R017 | Dimensional OTHER | Traveler Step6: Magnetic Hygiene Control. Multiple attempts were made to demagnetize cavity. 5 of 7 OOT areas were corrected. (2)Max Allowable dimensions OOT: NCR95 generated. MA3 As Found= 2.2 Gauss, Final= 2.2 Gauss MA7 As Found= 4.5 Gauss, Final= 2.0 Gauss, 1 stud shown in photo. 5 of 7 OOT areas were corrected. | Hold/Re-measure | Cavity will proceed to clean room and vertical test. If results confirm significant deterioration of Q, another degaussing at the larger degausser should be conducted. Furthermore, we might consider taking the individual bolt out and degauss it directly. | Use As Is | Cavity tested with good Q
96 | Nov-22-2021 | Dec-13-2021 | L2HE-INSP-CAV | 19 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R013 | Dings | Please see attached pictures of pits/dings on bellows. |  |  | Return To Vendor | Vendor will replace bellows
97 | Nov-23-2021 | Dec-13-2021 | L2HE-INSP-CAV | 19 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R013 | OTHER | Long fill line transition joint exceeds limit for magnetic hygiene check. The area has been demagnetized multiple times. Initial measurement 1.1 Gauss, Final measurement is 0.8 Gauss | Hold/Re-measure | Which degausser has been used? The Maurer one? Suggest to proceed with RF measurement etc. and see if this causes an issue. | Return To Vendor | Cavity was/is returned to vendor for bellows repair. After return, a new survey of magnetization has to be conducted.
98 | Dec-03-2021 | Dec-08-2021 | L2HE-INV-IMAG | 28 | MAGNETIC SHIELD CAVITY ASSEMBLY | RFCHE-MS-092 | ShippingDamage | See traveler for photos |  |  | Return To Vendor | Having communicated with Fermilab and the vendor, vendor asks the entire shipment to be returned.
99 | Dec-06-2021 | Mar-15-2022 | L2HE-VTA-CAV-VTRF | 28 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R016 | N/A | Unable to reach 23 MV/m, FE onset 12 MV/m | Hold/Modify | Cavity has to be re-rinsed and re-tested | Use As Is | cleaned up after re-rinse
100 | Dec-06-2021 | Mar-15-2022 | L2HE-VTA-CAV-VTRF | 28 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R016 | N/A | Qext2 is low, 2.2e11 | Hold/Modify | Cavity has to be re-rinsed and re-tested Will check on Q_ext again. | Reject | This cavity will be send to FNAL for a retest. If the cavity is going to be used in a string, the pickup antenna has to be replaced.
101 | Dec-06-2021 | Dec-08-2021 | L2HE-INSP-IMAG | 36 | MAGNETIC SHIELD CAVITY ASSEMBLY | RFCHE-MS-092 | ShippingDamage | Receiving visual inspection of two pallets, 4 boxes on each, of Magnetic Shield Assemblies(MAG) performed. Both pallets showed damage to the upper boxes. Only SN-092 MAG assembly was damaged. See attached photos. See attached photos. |  |  | Return To Vendor | Having communicated with Fermilab and the vendor, vendor asks the entire shipment to be returned.
102 | Dec-07-2021 | Jan-06-2022 | L2HE-INSP-QUAD | 1 | MAGNET PACKAGE ASSEMBLY | SPQA201 | InspectionFailure | At section 3.2.4, recording of the Q factor the spec/requirement is .3-.4. The actual value at time of inspection is .264 | Hold/Re-measure | Retest and record the q factor. Attach digital picture showing the displayed value by the equipemnt. | Use As Is | Our electrical inspection results indicate that the measured value of the Q factor at 100 Hz (0.264 (from test #1) and 0.265 (from test #2)) for the SPQA 201 magnet is marginally below the lower bound (0.3) of the anticipated value. Note that these limits were set according to the data, rounded off to one decimal place, obtained from the series of measurements at the manufacturer site and further at Fermi Lab. Since the room temperature resistance of the coils and magnet assembly and also the Q factor at 1 kHz fall within the suggested range, the observed discrepancy in the Q value at 100 Hz can be attributed to a statistical error in measurements. Nonetheless, it is expected to have no measurable impact on magnet performance at the actual operating conditions.
103 | Dec-07-2021 | Jun-28-2023 | L2HE-CMA-UCM-INSP | 4 | LCLS-II UPPER COLD MASS ASSEMBLY | RFCHE-UC-008 | Dings | Knife edge damaged on 2.2K conflat. Please see attached. |  |  | Use As Is | Leak check has been done
104 | Dec-08-2021 | Aug-29-2023 | L2HE-CMA-VV-INSP | 4 | VESSEL LCLS-II AFTER MACHING WELDMENT | RFCHE-VV-008 | Dings OTHER Roughness Scratches | NCR generated for Non-Conformances on the attached F10026610 VV-008 QC Task List. All items QC Task List that are "RED" are Non-Conforming and will need resolution. All photos can be found on server: M:\cmm\Pictures\Inspection Pictures\L2HE\Vacuum vessel\VV008\Receiving Inspection | Hold/Modify | Visual inspection was done previously. VV 008 is moved into Test Lab to get ready for installation onto J1.3-24 cryomodule. Issues discovered during visual inspection are marked on the VV008 and annotated in the QC spreadsheet. Use these as guidelines and make necessary polishing, cleaning | Conforming | Necessary polishing of the flanges are done.
105 | Dec-08-2021 | Mar-15-2022 | L2HE-VTA-CAV-VTRF | 29 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R017 | RFTestFailure | FE limited at 7 MV/m | Hold/Modify | Cavity will have to be re-rinsed & retested | Use As Is | cleaned up after re-rinse
106 | Dec-09-2021 | Dec-09-2021 | L2HE-INSP-CAV | 22 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R012 | Dings | Dents present on flange blank (valve end of cavity). Please see pictures attached. |  |  | Use As Is | Parts are not permanent. Probably reused parts from L2
107 | Dec-09-2021 | Dec-17-2021 | L2HE-INSP-CAV | 21 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R022 | Dimensional | Long fill line-SS joint: MaxAllowable6 was not able to be demagnetized. Started at 12.68 Gauss; after several attempts to demagnetize, final reading is 8.64 Gauss. All other areas are within tolerance. | Hold/Modify | Is it possible to take that to the large Maurer degausser in the North Annex? Please try to degauss there. | Conforming | Good work again!
108 | Dec-13-2021 | Dec-17-2021 | L2HE-INSP-CAV | 22 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R012 | Dimensional | 1) Long fill line-transition joint: MaxAllowable5 was not able to be demagnetized. Started at 3.35 Gauss; after several attempts to demagnetize, final reading is 0.99 Gauss. 2) Long fill line-SS joint: MaxAllowable6 was not able to be demagnetized. Started at 8.56 Gauss; after several attempts to demagnetize, final reading is 6.6 Gauss. 3) Beamline coupler flange Stud: MaxAllowable7; 1 stud was not able to be demagnetized. Started at 2.95 Gauss; after several attempts to demagnetize, final reading is 2.76 Gauss. 4) All other areas are within tolerance. | Hold/Modify | Noted -- will wait for the test results. Another degauss with the Maurer device might be necessary. | Conforming | Good work!
109 | Dec-20-2021 | Dec-20-2021 | L2HE-CLNRM-FPCC-INSP | 2 | FPC COLD PART | CPI-FPCC-HE-007 | Residues | Excessive aluminum residue from the seal on the cavity flange |  |  | Return To Vendor | Return to vendor to clean sealing surface on cavity flange
110 | Dec-20-2021 | Dec-20-2021 | L2HE-CLNRM-FPCC-INSP | 2 |  | CPI-FPCC-HE-008 | Scratches | Three scratches on copper |  |  | Return To Vendor
111 | Jan-05-2022 | Jan-14-2022 | L2HE-INSP-CAV | 25 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R029 | Dings | Damage present on bellows. Please see attached pictures. |  |  | Return To Vendor | Cavity will be returned to cavity for warranty repair (+HE-CAVR032)
112 | Jan-05-2022 | Feb-20-2024 | L2HE-CWI-VV | 3 | VESSEL AFTER MACHING WELDMENT | RFCHE-VV-006 | N/A | The internal CJP groove weld on Weld Joint #C1 is acceptable (*see weld-map). However, the external flange intermittent fillet welds were reduced in leg size. That reduction in fillet weld size is below an acceptable size in accordance with the drawings and specifications. |  |  | Use As Is | VV004 has similar issue on the stitched welds. Fermi informed JLAB that the stitched welds are not needed for structural strength.
113 | Jan-06-2022 | Jan-14-2022 | L2HE-INSP-CAV | 26 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R032 | Dings | Please see attached pictures. Damage present on bellows. |  |  | Return To Vendor | Cavity will be returned to cavity for warranty repair (+HE-CAVR029)
114 | Jan-10-2022 | Jan-11-2022 | L2HE-CMA-FPCW-INSP | 11 | FPC WARM PART | CPI-FPCW-HE-013 | InspectionFailure | E-probe fitting is damaged |  |  | Return To Vendor
115 | Jan-13-2022 | Jan-13-2022 | L2HE-CMA-FPCW-INSP | 4 | FPC WARM PART | CPI-FPCW-HE-008 | Plating | Blisters on the ID of outer bellows. Large stain on the OD of the inner bellows |  |  | Use As Is | Defects were found to be smaller than the 1mm max size limit
116 | Jan-13-2022 | Jul-27-2022 | L2HE-INSP-CLMPSQ | 1 | SPLIT QUAD CLAMP KIT | 01 | OTHER | Missing hardware. F10042265 item 4 is missing 4 washers. Please see picture attached. | Hold/Modify | Complete the inspection and provide an itemized list and count of the missing hardware. TR will report the total count of missing hardware to Fermi Lab procurement team. They will contact the vendor and arrange to ship the missing components to J Lab. We will carry out the rework and verify that the specifications are met. | Conforming | Missing washers were purchased and installed. The assembly is now complete.
117 | Jan-13-2022 | Feb-28-2022 | L2HE-INSP-CLMPSQ | 1 | SPLIT QUAD CLAMP KIT | 01 | OTHER | Components are dirty. Please see attached. | Hold/Modify | Gently wipe off the dirt on the parts using a lint-free and scratch-free cloth/tissue saturated with isopropyl alcohol or ethanol. Take digital pictures before and after cleaning the parts, assign filename consistent with the naming format mentioned in the original inspection travel, and attach the electronic files to this document. Verify that parts are clean. Report any further discrepancies found during the rework. | Conforming | The parts are clean and therefore qualify for further use in the respective cryomodule assembly.
119 | Jan-13-2022 | Jan-18-2022 | L2HE-CLNRM-FPCC-INSP | 9 | FPC COLD PART | CPI-FPCC-HE-017 | OTHER | Unable to remove stiffening plates due to the top two 3m bolts being stripped on both FPC bellows |  |  | Return To Vendor
120 | Jan-13-2022 | Sep-19-2024 | L2HE-VTA-CAV-VTRF | 31 |  | HE-R017 | RFTestFailure | Cavity unable to reach required gradient 23 MV/m | Hold/Modify | We will put this cavity on hold for the time being. It might be selected for a string in combination with strong performing cavities at a later point in time. | Use As Is | 02-Jul-2024 CTB Decision: Accepted by CTB as marginal and to be used in string J31. See RoD No. LCLSII-HE-1.2-PM-1156 02-Feb-2022 Cavity was put on hold by CTB. Previous Final Disposition was: Reject, corrected 16-Sep-2024.
121 | Jan-14-2022 | Jul-27-2022 | L2HE-INSP-CLMPSQ | 2 | SPLIT QUAD CLAMP KIT | 02 | OTHER | F10042265 item 4 is missing four washers. Please see attached picture. | Hold/Modify | Complete the inspection, provide an itemized list and count of the missing hardware. TR will report the total count of missing hardware to Fermi Lab procurement team. They will contact the vendor and arrange to ship the missing components to J Lab. We will carry out the rework and verify that the specifications are met. | Conforming | Missing washers were purchased and installed. The assembly is now complete.
122 | Jan-14-2022 | Feb-28-2022 | L2HE-INSP-CLMPSQ | 2 | SPLIT QUAD CLAMP KIT | 02 | OTHER | Components are dirty. Please see attached. | Hold/Modify | Gently wipe off the dirt on the parts using a lint-free and scratch-free cloth/tissue saturated with isopropyl alcohol or ethanol. Take digital pictures before and after cleaning the parts, assign filename consistent with the naming format mentioned in the original inspection travel, and attach the electronic files to this document. Verify that parts are clean. Report any further discrepancies found during the rework. | Conforming | The parts are clean and therefore qualify for further use in the respective cryomodule assembly.
123 | Jan-14-2022 | Feb-15-2022 | L2HE-INV-BLXD | 9 | WELDMENT EXTENSION-DS | 018 | ShippingDamage | Box appears to be crushed |  |  | Conforming | Upon inspection, spool extension appears fine
124 | Jan-14-2022 | Feb-15-2022 | L2HE-INV-BLXD | 10 | WELDMENT EXTENSION-DS | 019 | ShippingDamage | Box appears to be crushed. |  |  | Conforming | Upon inspection, spool extension appears fine
125 | Jan-14-2022 | Feb-15-2022 | L2HE-INV-BLXD | 11 | WELDMENT EXTENSION-DS | 020 | ShippingDamage | Box appears to be crushed. |  |  | Conforming | Upon inspection, spool extension appears fine
126 | Jan-14-2022 | Feb-15-2022 | L2HE-INV-BLXD | 12 | WELDMENT EXTENSION-DS | 021 | ShippingDamage | Box appears to be crushed. |  |  | Conforming | Upon inspection, spool extension appears fine
127 | Jan-18-2022 | Feb-28-2022 | L2HE-INSP-CLMP50K | 1 | 50K CLAMP ASSEMBLY FULL WRAP KIT | 01 | OTHER | Please see attached. Components are dirty. | Hold/Modify | Gently wipe off the dirt on the parts using a lint-free and scratch-free cloth/tissue saturated with isopropyl alcohol or ethanol. Take digital pictures before and after cleaning the parts, assign filename consistent with the naming format mentioned in the original inspection travel, and attach the electronic files to this document. Verify that parts are clean. Report any further discrepancies found during the rework. | Conforming | The parts are clean and therefore qualify for further use in the respective cryomodule assembly.
128 | Jan-18-2022 | Feb-28-2022 | L2HE-INSP-CLMP50K | 2 | 50K CLAMP ASSEMBLY FULL WRAP KIT | 02 | OTHER | Components are dirty. Please see attached. | Hold/Modify | Gently wipe off the dirt on the parts using a lint-free and scratch-free cloth/tissue saturated with isopropyl alcohol or ethanol. Take digital pictures before and after cleaning the parts, assign filename consistent with the naming format mentioned in the original inspection travel, and attach the electronic files to this document. Verify that parts are clean. Report any further discrepancies found during the rework. | Conforming | The parts are clean and therefore qualify for further use in the respective cryomodule assembly.
129 | Jan-19-2022 | Feb-28-2022 | L2HE-INSP-CLMP05K | 2 | 5K CLAMP ASSEMBLY KIT | 01 | OTHER | Components are dirty. Please see attached. | Hold/Modify | Gently wipe off the dirt on the parts using a lint-free and scratch-free cloth/tissue saturated with isopropyl alcohol or ethanol. Take digital pictures before and after cleaning the parts, assign filename consistent with the naming format mentioned in the original inspection travel, and attach the electronic files to this document. Verify that parts are clean. Report any further discrepancies found during the rework. | Conforming | The parts are clean and therefore qualify for further use in the respective cryomodule assembly.
130 | Jan-19-2022 | Feb-02-2022 | L2HE-VTA-CAV-VTRF | 32 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R016 | RFTestFailure | Failed to reach 23 MV/m, quench limited to 21 MV/m. | Hold/Modify | Put cavity on hold (compare HE-CAVR017) Waiting for CTB input. | Use As Is | Cavity on hold for use in string in combination with high gradient cavities.
131 | Jan-19-2022 | Feb-28-2022 | L2HE-INSP-CLMP05K | 3 | 5K CLAMP ASSEMBLY KIT | 02 | OTHER | Please see attached. Components are dirty. | Hold/Modify | Gently wipe off the dirt on the parts using a lint-free and scratch-free cloth/tissue saturated with isopropyl alcohol or ethanol. Take digital pictures before and after cleaning the parts, assign filename consistent with the naming format mentioned in the original inspection travel, and attach the electronic files to this document. Verify that parts are clean. Report any further discrepancies found during the rework. | Conforming | The parts are clean and therefore qualify for further use in the respective cryomodule assembly.
132 | Jan-19-2022 | Jul-27-2022 | L2HE-INSP-CLMP02K | 1 | 2K CLAMP ASSEMBLY KIT | 01 | OTHER | Drawing # F10040841 item 11 (split lock washers) missing. Please see attached pictures. | Hold/Modify | Complete the inspection, provide an itemized list and count of the missing hardware. TR will report the total count of missing hardware to Fermi Lab procurement team. They will contact the vendor and arrange to ship the missing components to J Lab. We will carry out the rework and verify that the specifications are met. | Conforming | Missing washers were purchased and installed. The assembly is now complete.
133 | Jan-19-2022 | Feb-28-2022 | L2HE-INSP-CLMP02K | 1 | 2K CLAMP ASSEMBLY KIT | 01 | OTHER | Component very dirty. Please see attached. | Hold/Modify | Gently wipe-off the dirt on the parts using a lint-free and scratch-free cloth/tissue saturated with isopropyl alcohol or ethanol. Take digital pictures before and after cleaning the parts, assign filename consistent with the naming format mentioned in the original inspection travel, and attach the electronic files to this document. Verify that parts are clean. Report any further discrepancies found during the rework. | Conforming | The cleanliness issue has been resolved.
134 | Jan-19-2022 | Jul-27-2022 | L2HE-INSP-CLMP02K | 3 | 2K CLAMP ASSEMBLY KIT | 02 | OTHER | Drawing # F10040841 item 11 (split lock washers) missing. | Hold/Modify | Complete the inspection, provide an itemized list, and total count of the missing hardware per part. TR will report the total count of missing hardware to the Fermi Lab procurement team. They will contact the vendor and arrange to ship the missing components to J Lab. We will carry out the rework and verify that the specifications are met. | Conforming | Missing washers were purchased and installed. The assembly is now complete.
135 | Jan-19-2022 | Feb-28-2022 | L2HE-INSP-CLMP02K | 3 | 2K CLAMP ASSEMBLY KIT | 02 | OTHER | Components dirty. Pictures attached. | Hold/Modify | Gently wipe off the dirt on the parts using a lint-free and scratch-free cloth/tissue saturated with isopropyl alcohol or ethanol. Take digital pictures before and after cleaning the parts, assign filename consistent with the naming format mentioned in the original inspection travel, and attach the electronic files to this document. Verify that parts are clean. Report any further discrepancies found during the rework. | Conforming | The parts are clean and therefore qualify for further use in the respective cryomodule assembly.
136 | Jan-20-2022 | Mar-01-2022 | L2HE-INSP-CFL | 2 | CURRENT LEADS CF FLANGE ASSEMBLY | 0044 | Dimensional Instrumentation OTHER | All (6) CL Flange lead wires are too short. All (6) pcs of item 15 on dwg F10022642-Rev G are ~12 inches long, not 15.0 +/-0.1 inches, as specified on F10078481-Rev A, WIRE LEAD WITH LUG, SPQA FLANGE. | Hold/Modify | Fermi Lab will supply 15 inches long lead wires to replace all 12 inches long wires in the CL flange assemblies we received directly from the vendor. J lab technical staff will inspect the new wires and replace the shorter lead wires to qualify the part with serial no #0044 for further use in the corresponding CM assembly. | Conforming | The rework results confirm that the replacement lead wires meet the dimensional specification.
137 | Jan-20-2022 | Mar-01-2022 | L2HE-INSP-CFL | 3 | CURRENT LEADS CF FLANGE ASSEMBLY | 0045 | Dimensional Instrumentation OTHER | All (6) CL Flange lead wires are too short. All (6) pcs of item 15 on dwg F10022642-Rev G are ~12 inches long, not 15.0 +/-0.1 inches, as specified on F10078481-Rev A, WIRE LEAD WITH LUG, SPQA FLANGE. | Hold/Modify | Fermi Lab will supply 15 inches long lead wires to replace all 12 inches long wires in the CL flange assemblies we received directly from the vendor. J lab technical staff will inspect the new wires and replace the shorter lead wires to qualify the part with serial no #0045 for further use in the corresponding CM assembly. | Conforming | The rework results confirm that the replacement lead wires meet the dimensional specification.
138 | Jan-20-2022 | Mar-01-2022 | L2HE-INSP-CFL | 4 | CURRENT LEADS CF FLANGE ASSEMBLY | 0046 | Dimensional Instrumentation OTHER | All (6) CL Flange lead wires are too short. All (6) pcs of item 15 on dwg F10022642-Rev G are ~12 inches long, not 15.0 +/-0.1 inches, as specified on F10078481-Rev A, WIRE LEAD WITH LUG, SPQA FLANGE. | Hold/Modify | Fermi Lab will supply 15 inches long lead wires to replace all 12 inches long wires in the CL flange assemblies we received directly from the vendor. J lab technical staff will inspect the new wires and replace the shorter lead wires to qualify the part with serial no #0046 for further use in the corresponding CM assembly. | Conforming | The rework results confirm that the replacement lead wires meet the dimensional specification.
139 | Jan-20-2022 | Mar-01-2022 | L2HE-INSP-CFL | 5 | CURRENT LEADS CF FLANGE ASSEMBLY | 0047 | Dimensional Instrumentation OTHER | All (6) CL Flange lead wires are too short. All (6) pcs of item 15 on dwg F10022642-Rev G are ~12 inches long, not 15.0 +/-0.1 inches, as specified on F10078481-Rev A, WIRE LEAD WITH LUG, SPQA FLANGE. | Hold/Modify | Fermi Lab will supply 15 inches long lead wires to replace all 12 inches long wires in the CL flange assemblies we received directly from the vendor. J lab technical staff will inspect the new wires and replace the shorter lead wires to qualify the part with serial no #0047 for further use in the corresponding CM assembly. | Conforming | The rework results confirm that the replacement lead wires meet the dimensional specification.
140 | Jan-20-2022 | Mar-01-2022 | L2HE-INSP-CFL | 6 | CURRENT LEADS CF FLANGE ASSEMBLY | 0048 | Dimensional Instrumentation OTHER | All (6) CL Flange lead wires are too short. All (6) pcs of item 15 on dwg F10022642-Rev G are ~12 inches long, not 15.0 +/-0.1 inches, as specified on F10078481-Rev A, WIRE LEAD WITH LUG, SPQA FLANGE. | Hold/Modify | Fermi Lab will supply 15 inches long lead wires to replace all 12 inches long wires in the CL flange assemblies we received directly from the vendor. J lab technical staff will inspect the new wires and replace the shorter lead wires to qualify the part with serial no #0048 for further use in the corresponding CM assembly. | Conforming | The rework results confirm that the replacement lead wires meet the dimensional specification.
141 | Jan-20-2022 | Mar-01-2022 | L2HE-INSP-CFL | 7 | CURRENT LEADS CF FLANGE ASSEMBLY | 0049 | Dimensional Instrumentation OTHER | All (6) CL Flange lead wires are too short. All (6) pcs of item 15 on dwg F10022642-Rev G are ~12 inches long, not 15.0 +/-0.1 inches, as specified on F10078481-Rev A, WIRE LEAD WITH LUG, SPQA FLANGE. | Hold/Modify | Fermi Lab will supply 15 inches long lead wires to replace all 12 inches long wires in the CL flange assemblies we received directly from the vendor. J lab technical staff will inspect the new wires and replace the shorter lead wires to qualify the part with serial no #0049 for further use in the corresponding CM assembly. | Conforming | The rework results confirm that the replacement lead wires meet the dimensional specification.
142 | Jan-20-2022 | Mar-01-2022 | L2HE-INSP-CFL | 8 | CURRENT LEADS CF FLANGE ASSEMBLY | 0050 | Dimensional Instrumentation OTHER | All (6) CL Flange lead wires are too short. All (6) pcs of item 15 on dwg F10022642-Rev G are ~12 inches long, not 15.0 +/-0.1 inches, as specified on F10078481-Rev A, WIRE LEAD WITH LUG, SPQA FLANGE. | Hold/Modify | Fermi Lab will supply 15 inches long lead wires to replace all 12 inches long wires in the CL flange assemblies we received directly from the vendor. J lab technical staff will inspect the new wires and replace the shorter lead wires to qualify the part with serial no #0050 for further use in the corresponding CM assembly. | Conforming | The rework results confirm that the replacement lead wires meet the dimensional specification.
143 | Jan-20-2022 | Mar-01-2022 | L2HE-INSP-CFL | 9 | CURRENT LEADS CF FLANGE ASSEMBLY | 0051 | Dimensional Instrumentation OTHER | All (6) CL Flange lead wires are too short. All (6) pcs of item 15 on dwg F10022642-Rev G are ~12 inches long, not 15.0 +/-0.1 inches, as specified on F10078481-Rev A, WIRE LEAD WITH LUG, SPQA FLANGE. | Hold/Modify | Fermi Lab will supply 15 inches long lead wires to replace all 12 inches long wires in the CL flange assemblies we received directly from the vendor. J lab technical staff will inspect the new wires and replace the shorter lead wires to qualify the part with serial no #0051 for further use in the corresponding CM assembly. | Conforming | The rework results confirm that the replacement lead wires meet the dimensional specification.
144 | Jan-20-2022 | Mar-01-2022 | L2HE-INSP-CFL | 10 | CURRENT LEADS CF FLANGE ASSEMBLY | 0052 | Dimensional Instrumentation OTHER | All (6) CL Flange lead wires are too short. All (6) pcs of item 15 on dwg F10022642-Rev G are ~12 inches long, not 15.0 +/-0.1 inches, as specified on F10078481-Rev A, WIRE LEAD WITH LUG, SPQA FLANGE. | Hold/Modify | Fermi Lab will supply 15 inches long lead wires to replace all 12 inches long wires in the CL flange assemblies we received directly from the vendor. J lab technical staff will inspect the new wires and replace the shorter lead wires to qualify the part with serial no #0052 for further use in the corresponding CM assembly. | Conforming | The rework results confirm that the replacement lead wires meet the dimensional specification.
145 | Jan-20-2022 | Jan-26-2022 | L2HE-INSP-CFL | 11 | CURRENT LEADS CF FLANGE ASSEMBLY | 0053 | Dimensional Instrumentation OTHER | All (6) CL Flange lead wires are too short. All (6) pcs of item 15 on dwg F10022642-Rev G are ~12 inches long, not 15.0 +/-0.1 inches, as specified on F10078481-Rev A, WIRE LEAD WITH LUG, SPQA FLANGE. | Hold/Modify | Fermi Lab will supply 15 inches long lead wires to replace all 12 inches long wires in the CL flange assemblies we received directly from the vendor. J lab technical staff will inspect the new wires and replace the shorter lead wires to qualify the part with serial no #0053 for further use in the corresponding CM assembly. | Return To Vendor | This assembly did not pass the vacuum leak test (Ref - NCR#147). It will be returned to the vendor. Therefore, further rework on the lead wires is not required.
147 | Jan-26-2022 | Jan-26-2022 | L2HE-INSP-CFL | 11 | CURRENT LEADS CF FLANGE ASSEMBLY | 0053 | LeakTestFailure | THE LEAK SHOWS FAULTY FEED THROUGH PIN |  |  | Return To Vendor | The observed vacuum leak disqualifies the part from being used in the CM assembly. Fermi Lab led this procurement. So, I notified the non-conformance issue to the vendor through the FNAL procurement staff and established a path forward. Action plan -J Lab will send the item back to the vendor for a replacement. We will receive the new assembly directly from the vendor. We will inspect and qualify it.
148 | Jan-31-2022 | Feb-01-2022 | L2HE-VTA-CAV-VTRF | 33 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R022 | RFTestFailure | Qext1=1.96e10 a bit high Qext2=2.0e11 low |  |  | Use As Is | Use as is -- CTB decision 2022-02-1
149 | Jan-31-2022 | Feb-01-2022 | L2HE-VTA-CAV-VTRF | 34 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R023 | RFTestFailure | Qo at 20.8 MV/m low (2.45e10 on second run) |  |  | Use As Is | Use as is -- CTB decision 2022-02-1
150 | Jan-31-2022 | Feb-01-2022 | L2HE-VTA-CAV-VTRF | 35 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R011 | RFTestFailure | Max gradient of 22.67 MV/m |  |  | Use As Is | Use as is -- CTB decision 2022-02-1 Cavity will be put on hold to be combined with better cavities in a string (open CTB decision)
151 | Feb-07-2022 | Jul-27-2022 | L2HE-INSP-CLMPSQ | 3 | SPLIT QUAD CLAMP KIT | 03 | OTHER | Component is missing hardware. DWG # F10042265 Item 4 is missing four washers. Please see picture attached. | Hold/Modify | Complete the inspection, provide an itemized list and count of the missing hardware. TR will report the total count of missing hardware to Fermi Lab procurement team. They will contact the vendor and arrange to ship the missing components to J Lab. We will carry out the rework and verify that the specifications are met. | Conforming | Missing washers were purchased and installed. The assembly is now complete.
152 | Feb-07-2022 | Feb-28-2022 | L2HE-INSP-CLMPSQ | 3 | SPLIT QUAD CLAMP KIT | 03 | OTHER | Component is dirty. Please see attached. | Hold/Modify | Gently wipe off the dirt on the parts using a lint-free and scratch-free cloth/tissue saturated with isopropyl alcohol or ethanol. Take digital pictures before and after cleaning the parts, assign filename consistent with the naming format mentioned in the original inspection travel, and attach the electronic files to this document. Verify that parts are clean. Report any further discrepancies found during the rework. | Conforming | The parts are clean and therefore qualify for further use in the respective cryomodule assembly.
153 | Feb-08-2022 | Jul-27-2022 | L2HE-INSP-CLMP02K | 4 | 2K CLAMP ASSEMBLY KIT | 03 | OTHER | Component is missing hardware. Item 11 is missing 8 split lock washers and item 12 is missing 2 belleville washers Please see attached. | Hold/Modify | Complete the inspection, provide an itemized list, and total count of the missing hardware per part. TR will report the total count of missing hardware to the Fermi Lab procurement team. They will contact the vendor and arrange to ship the missing components to J Lab. We will carry out the rework and verify that the specifications are met. | Conforming | Missing washers were purchased and installed. The assembly is now complete.
154 | Feb-08-2022 | Feb-28-2022 | L2HE-INSP-CLMP02K | 4 | 2K CLAMP ASSEMBLY KIT | 03 | OTHER | Please see attached pictures, component is dirty. | Hold/Modify | Gently wipe off the dirt on the parts using a lint-free and scratch-free cloth/tissue saturated with isopropyl alcohol or ethanol. Take digital pictures before and after cleaning the parts, assign filename consistent with the naming format mentioned in the original inspection travel, and attach the electronic files to this document. Verify that parts are clean. Report any further discrepancies found during the rework. | Conforming | The parts are clean and therefore qualify for further use in the respective cryomodule assembly.
155 | Feb-08-2022 | Feb-28-2022 | L2HE-INSP-CLMP50K | 3 | 50K CLAMP ASSEMBLY FULL WRAP KIT | 03 | OTHER | Component is dirty. Please see attached. | Hold/Modify | Gently wipe off the dirt on the parts using a lint-free and scratch-free cloth/tissue saturated with isopropyl alcohol or ethanol. Take digital pictures before and after cleaning the parts, assign filename consistent with the naming format mentioned in the original inspection travel, and attach the electronic files to this document. Verify that parts are clean. Report any further discrepancies found during the rework. | Conforming | The parts are clean and therefore qualify for further use in the respective cryomodule assembly.
156 | Feb-08-2022 | Feb-28-2022 | L2HE-INSP-CLMP05K | 4 | 5K CLAMP ASSEMBLY KIT | 03 | OTHER | Component is dirty. Please see attached. | Hold/Modify | Gently wipe off the dirt on the parts using a lint-free and scratch-free cloth/tissue saturated with isopropyl alcohol or ethanol. Take digital pictures before and after cleaning the parts, assign filename consistent with the naming format mentioned in the original inspection travel, and attach the electronic files to this document. Verify that parts are clean. Report any further discrepancies found during the rework. | Conforming | The parts are clean and therefore qualify for further use in the respective cryomodule assembly.
157 | Feb-09-2022 | Feb-16-2022 | L2HE-VTA-CAV-VTRF | 36 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R012 | N/A | Eacc low: quench at 22.49 MV/m Qext1 high: 2.07e10 |  |  | Use As Is | Use as is -- CTB decision 2022-02-15 Cavity will be put on hold to be combined with better cavities in a string (open CTB decision)
158 | Feb-11-2022 | Feb-11-2022 | L2HE-CMA-FPFT-LEAK | 149 | PICKUP FEEDTHROUGH | 2115 | LeakTestFailure | Failed leak check. | Hold/Re-measure | Hold from production. Send back to SLAC. | Reject | This feedthrough is shipped to SLAC.
159 | Feb-11-2022 | Feb-11-2022 | L2HE-CMA-HMFT-LEAK | 281 | HOM FEEDTHROUGH | 2239 | LeakTestFailure | Failed leak check | Hold/Re-measure | Hold from production. Send back to SLAC. | Reject | This feedthrough is shipped to SLAC.
160 | Feb-11-2022 | Feb-11-2022 | L2HE-INSP-HMFT | 330 | HOM FEEDTHROUGH | 2306 | Dimensional InspectionFailure | Probe perpendicularity 1.26 degree, out of spec which is 1.0 deg max. | Hold/Re-measure | Hold from production. Send back to SLAC for further evaluation. | Reject | This feedthrough is shipped back to SLAC for further evaluation.
161 | Feb-16-2022 | Feb-17-2022 | L2HE-INSP-CAV | 29 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R035 | Dimensional | Dimensional non-conforming. 2-phase tee parallelism out of tolerance 0.4 mm |  |  | Use As Is | Should not cause an issue -- checked back with John F
162 | Feb-16-2022 | Feb-17-2022 | L2HE-INSP-CAV | 30 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R040 | Dimensional | DIM2 OOT: Should be //(parallelism) 0.0 +1.0/-.0. Is; 1.1 |  |  | Use As Is | Should not cause an issue -- checked back with John F
163 | Feb-23-2022 | Feb-23-2022 | L2HE-INSP-BLBP | 8 | WELDMENT BELLOWS PRCM | 077 | Dimensional | Please see attached for dimensional non-conforming. |  |  | Use As Is | OOT values are relatively small, and will not effect the flange features.
164 | Feb-24-2022 | Feb-25-2022 | L2HE-INSP-BLBP | 9 | WELDMENT BELLOWS PRCM | 089 | Dimensional | Dimensional non-conforming. Bellow ID nominal 78.00 mm. Bellow ID measured 79.16 mm. Flange pocket distance nominal 3.10 mm. Flange pocket distance measured 2.97 mm |  |  | Use As Is | Very small deviation that will not impact sealing performance
165 | Feb-25-2022 | Feb-25-2022 | L2HE-VTA-CAV-VTRF | 42 | PRODUCTION CAVITY ASSEMBLY | HE-295 | N/A | Qe2 is below specified range, 5.54e+11 |  |  | Use As Is | Noted in the database -- use as is.
166 | Feb-25-2022 | Feb-25-2022 | L2HE-VTA-CAV-VTRF | 40 | LCLS-II PRODUCTION CAVITY ASSEMBLY | 0351 | N/A | Qe1 is below specified range, 9.62e+9 |  |  | Use As Is | Results still ok -- antenna will be replaced during string assembly.
168 | Mar-01-2022 | Mar-15-2022 | L2HE-CLNRM-FPCC-INSP | 13 | FPC COLD PART | CPI-FPCC-HE-023 | Dimensional Parallelism | bellows were received in this condition. removed braces to see if it would straighten on its own but nothing happened |  |  | Return To Vendor | Both CPI-FPCC-HE-022 and CPI-FPCC-HE-023 are to be returned to the vendor on the same shipping frame. RMA 8594
169 | Mar-02-2022 | Mar-02-2022 | L2HE-INSP-CAV | 31 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R039 | Dimensional | 3 dimensions are Out of Tolerance (OOT). See attached report, or Traveler Page:4. |  |  | Use As Is | Should be ok for string integration.
170 | Mar-03-2022 | Mar-09-2022 | L2HE-CMA-FPCW-INSP | 16 | FPC WARM PART | CPI-FPCW-HE-018 | ShippingDamage | Both inner and outer bag had rips in them |  |  | Use As Is | This will not affect the operation of the FPC. The vendor has been informed and will take steps to ensure that the bags remain intact.
171 | Mar-03-2022 | Mar-09-2022 | L2HE-CMA-FPCW-INSP | 18 | FPC WARM PART | CPI-FPCW-HE-020 | ShippingDamage | Both inner and outer bags had holes in them. Holes may have been caused by excessive suction force from evacuated internal region. |  |  | Use As Is | This will not affect the operation of the FPC. The vendor has been informed and will take steps to ensure that the bags remain intact.
172 | Mar-07-2022 | Mar-08-2022 | L2HE-VTA-CAV-VTRF | 43 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R040 | N/A | Qe1 (2.2e+10) and Qe2 (6.7e+12) are out of tolerance |  |  | Use As Is | Accepted by CTB on 2022-03-08
173 | Mar-08-2022 | Mar-08-2022 | L2HE-INSP-BLBP | 10 | WELDMENT BELLOWS PRCM | 074 | Plating | Lots of deep pits present on multiple areas on copper plating. Pictures attached. |  |  | Return To Vendor | Return to Nomura for re-plating
174 | Mar-08-2022 | May-17-2022 | L2HE-INSP-BLBP | 15 | WELDMENT BELLOWS PRCM | 003 | Plating | Unusual rough/uneven copper plating present. Please see attached pictures. | Hold/Modify | Send to Chemistry for cleaning. | Conforming | Cleaned by chemistry
175 | Mar-08-2022 | May-31-2022 | L2HE-VTA-CAV-VTRF | 44 | 1.3GHZ LCLS-II RF CAVITY ASSY (BARE) | 362 | RFTestFailure | Cavity quenched at 22 MV/m in VTRF, also Qext1=2.2e10. | Hold/Modify | Wait for further test results of next two nine-cells | Reject | This cavity is rejected for further work and will be delivered to SLAC as is.
176 | Mar-09-2022 | Mar-11-2022 | L2HE-INSP-BLBP | 16 | WELDMENT BELLOWS PRCM | 011 | Dings | Seal groove outside diameter of rotatable flange damaged. Please see attached. |  |  | Return To Vendor | Damaged seal groove prevents seal ring from falling into place freely. Return to vendor Nomura.
177 | Mar-09-2022 | Mar-11-2022 | L2HE-INSP-BLBP | 16 | WELDMENT BELLOWS PRCM | 011 | Scratches Stains | Scratch/stain present in non-rotatable flange seal groove. Pictures attached. |  |  | Return To Vendor | Damaged seal groove reported by NCR 176 prevents seal ring from sitting in groove.
178 | Mar-14-2022 | May-17-2022 | L2HE-INSP-BLBP | 20 | WELDMENT BELLOWS PRCM | 029 | Scratches | Scratches in seal area of rotatable flange. Please see attached. | Hold/Modify | Hand polish seal surface | Conforming | Seal surface polished by machine shop
179 | Mar-25-2022 | Mar-28-2022 | L2HE-VTA-CAV-VTRF | 45 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R012 | N/A | Qext1 was 2.22e10, out of spec which is 1.1 e10? |  |  | Use As Is | This can be accepted as the FPC will be exchanged. See NCR 157
180 | Mar-25-2022 | Mar-28-2022 | L2HE-VTA-CAV-VTRF | 46 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R022 | N/A | FE onset at 4.8 MV/m during final run. |  |  | Reject | Cavity will be re-rinsed & re-tested
181 | Mar-25-2022 | May-31-2022 | L2HE-VTA-CAV-VTRF | 47 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R035 | N/A | Qext2 was 8.88e12, out of spec which is 7.5 e11 |  |  | Use As Is | Accepted by CTB 2022-03-29
182 | Mar-25-2022 | Mar-28-2022 | L2HE-VTA-CAV-VTRF | 45 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R012 | N/A | Emax was 22.5 MV/m, lower than spec which is >/= 23 MV/m. |  |  | Use As Is | CTB decision 2022-02-15 Cavity will be put on hold to be combined with better cavities in a string (open CTB decision) See NCR 157
183 | Mar-30-2022 | Apr-06-2022 | L2HE-INSP-BLBP | 17 | WELDMENT BELLOWS PRCM | 024 | Dimensional | Dimensional non-conforming. PDF attached. Length measured 107.0 mm Bellows ID measured 79.2 mm Flange depth measured 2.97 mm (Nominal 3.10 mm) |  |  | Use As Is | Length and ID measurements do not effect the functionality of the bellows, the OOT values are relatively low.
184 | Mar-30-2022 | May-17-2022 | L2HE-INSP-BLBP | 14 | WELDMENT BELLOWS PRCM | 002 | Plating | Excessive spots/stains larger than 1/8 inch inside convolutions of copper plating. Please see pictures attached. | Hold/Modify | Send to chemistry for cleaning | Conforming | Cleaned by chemistry
185 | Mar-31-2022 | Mar-31-2022 | L2HE-INSP-CAV | 35 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R057 | Dings | Two small dings/imperfections located on bellows. Please see attached. |  |  | Use As Is | Dings seem acceptable, cavity will be shipped to Fermi with a corresponding note.
186 | Apr-07-2022 | Apr-08-2022 | L2HE-VTA-CAV-VTRF | 48 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R039 | N/A | Qext1 was 2.04e10, out of spec (1.1 e10? |  |  | Use As Is | This deviation is acceptable, given the Q result and the fact that the high-Q antenna is being replaced with the FPC.
187 | Apr-08-2022 | Feb-28-2023 | L2HE-VTA-CAV-VTRF | 49 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R042 | N/A | Radiation onset at 16 MV/m, out of spec (FE free to the peak gradient). | Hold/Re-measure | re-rinse & re-test | Conforming | as of test on 2023-02-01 conforming
188 | Apr-08-2022 | May-27-2023 | L2HE-VTA-CAV-VTRF | 50 | PRODUCTION CAVITY ASSEMBLY | HE-295 | N/A | Final run Q0(20.8) was 2.2300e+10, out of spec which is >/= 2.5 e10 | Hold/Re-measure | We will re-test this cavity on the next D5 run, as a thermal cycle was not feasible. There are indications that the FE problem cleaned up during the first ramp up. | Reject | Cavity rejected due to performance
189 | Apr-08-2022 | May-31-2022 | L2HE-VTA-CAV-VTRF | 51 | 1.3GHZ LCLS-II RF CAVITY ASSY (BARE) | 360 | N/A | Cavity unable to reach 23 MV/m due to quench without FE. Qe1 also slightly high. | Hold/Modify | Wait for further test results of next two nine-cells | Reject | This cavity is rejected for further work and will be delivered to SLAC as is.
190 | Apr-10-2022 | Apr-11-2022 | L2HE-INSP-CAV | 36 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R043 | Dimensional | Dimension2 OOT. Should be: //0.0 to 1.0= 1.1; OOT by .1. See attached report. |  |  | Use As Is
191 | Apr-15-2022 | May-16-2022 | L2HE-CLNRM-FPCC-INSP | 5 | FPC COLD PART | CPI-FPCC-HE-014 | N/A | When removing cap from FPC coupler i broke the ceramic isolator on coupler 014 |  | placed coupler back on box and sent back for shipping | Return To Vendor | CPI-FPCC-HE-014 will be shipped back to the vendor installed on the same shipping stand as CPI-FPCC-HE-019
192 | Apr-15-2022 | May-16-2022 | L2HE-CLNRM-FPCC-INSP | 12 | FPC COLD PART | CPI-FPCC-HE-019 | OTHER | when removing cap from FPC coupler 019 i broke the ceramic isolator on top |  |  | Return To Vendor | CPI-FPCC-HE-019 will be shipped back to the vendor installed on the same shipping stand as CPI-FPCC-HE-014
193 | Apr-18-2022 | Jun-09-2022 | L2HE-INSP-FWM | 1 | FARADAY WINDOW MANIFOLD WELDMENT | 010 | Dimensional | Please see attached for dimensional non-conforming. | Hold/Modify | Determine if scratches in window can be removed and remeasure window thickness. | Conforming | Port Window has been repaired - scratches and thickness were removed by Machine Shop.
194 | Apr-21-2022 | Apr-21-2022 | L2HE-VTA-CAV-VTRF | 54 | PRODUCTION CAVITY ASSEMBLY | HE-295 | N/A | Qext2 was 5.89e11, out of spec which is 7.5 e11 |  |  | Use As Is | Accepted by CTB after earlier tests.
195 | Apr-21-2022 | Sep-19-2024 | L2HE-VTA-CAV-VTRF | 54 | PRODUCTION CAVITY ASSEMBLY | HE-295 | N/A | Final power rise Q0(20.8) 2.23e10, out of spec which is >/= 2.5 e10 | Hold/Modify | HPR (bottom flange drop) and re-test please document on page 3 | Use As Is | Cavity put in storage 2024-02 returned to TL for requalification 2024-03-01 qualified, non-conforming Qe2 accepted by CTB, see NCR 194
196 | Apr-22-2022 | May-02-2022 | L2HE-VTA-CAV-VTRF | 52 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R043 | N/A | Qext1 was 2.15e10, out of spec which is 1.1e10? |  |  | Use As Is | Accepted by CTB
197 | Apr-22-2022 | Jun-26-2022 | L2HE-VTA-CAV-VTRF | 52 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R043 | N/A | Qext2 was 1.12e13, out of spec which is 7.5e11 | Hold/Modify | Please replace pickup probe after disassembly | Conforming | Pick-up probe is now replaced Test on 2022-06-22 shows Qe2 in spec
198 | Apr-25-2022 | Apr-26-2022 | L2HE-INSP-CAV | 38 |  | HE-R056 | Dimensional | CMM inspection completed. See attached CMM report. DIM3 OOT. Should be: 0.5°= .7°; OOT by .2° |  |  | Use As Is | Deviation acceptable, cavity shall be used in a rear/downstream position, preferably position 8 in next string.
199 | Apr-25-2022 | Apr-26-2022 | L2HE-INSP-CAV | 39 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R059 | Dings | Burst Disc has small indentation (ding) on the seal surface. The ding was discovered when I uncrated the cavity from the shipping container and performed visual inspection. |  |  | Use As Is
200 | Apr-25-2022 | May-31-2022 | L2HE-VTA-CAV-VTRF | 55 |  | 359 | RFTestFailure | a)Cavity quench at 22.3 MV/m b)Second frequency manifest at 19 MV/m, more or less at the same time with radiation onset. |  |  | Reject | This cavity is rejected for further work and will be delivered to SLAC as is.
201 | Apr-27-2022 | May-02-2022 | L2HE-INSP-BLBS | 3 | WELDMENT BELLOWS SHORT | 004 | Plating | Please see attached for copper plating issue on Non-Rotatable flange. Plating is not consistent, bare spot present. |  |  | Use As Is | Copper plating does not need to go over the lip.
202 | Apr-27-2022 | Aug-04-2022 | L2HE-INSP-BLBS | 4 | WELDMENT BELLOWS SHORT | 005 | Scratches | Scratches through seal area present. Picture attached. | Hold/Modify | Send to machine shop for seal surface polishing | Conforming | Seal surfaces have been polished and are now acceptable.
203 | Apr-28-2022 | May-16-2022 | L2HE-CLNRM-FPCC-INSP | 17 | FPC COLD PART | CPI-FPCC-HE-038 | Dings Scratches | chipped ceramic and digs/scratches inside top covering. felt no resistance when separating the top covering from FPC body but found debris once separated |  |  | Return To Vendor | Part is to be kept on shipping stand in the cleanroom and returned to vendor once the partner unit (CPI-FPCC-HE-037) is installed on a string
204 | May-02-2022 | Sep-19-2024 | L2HE-VTA-CAV-VTRF | 57 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R022 | RFTestFailure | Qo at Emax (22.5 MV/m) 1.3e10 (smaller than specification 2.5e10). Cavity quenching from 19 MV/m. Multiple Pi modes over 19 MV/m. | Hold/Modify | What the operator meant to say is: "Q0 at 20.8 MV\m below spec (1.8e10 vs. 2.5e10), cavity started quenching and showing FE around 19 MV/m" (The occurrence of a 7/9 pi mode is noted in step 11, the Qext1 which is above spec should be indicated in a separate NCR.) The cavity will receive a full disassembly and 4 pass HPR. Please document the treatment on page 3 of this NCR. | Conforming | CTB 2022-06-07 rejected Cavity returned to vendor for reprocessing. Passed vertical test on 2024-03-08 as HE-CAVR022_RP
205 | May-11-2022 | May-12-2022 | L2HE-INSP-CAV | 42 |  | HE-R048 | Dings | Two small blemishes on bellows. Please see pictures attached. |  |  | Use As Is | discussed in CTB meeting on 2022-05-20 Feedback from Chuck -- damage should not limit functionality
206 | May-12-2022 | May-13-2022 | L2HE-VTA-CAV-VTRF | 59 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R060 | N/A | High: Qext1 = 2.03e10 |  |  | Use As Is | CTB Meeting of 2022-05-10 These Qe1's are acceptable. Spec to be updated.
207 | May-16-2022 | May-16-2022 | L2HE-CLNRM-FPCC-INSP | 6 | FPC COLD PART | CPI-FPCC-HE-011 | Scratches | Ceramic has large black mark where it was hit with the tophat. There is a possible crack at the edge closest to the flange. This was discovered during the inspection of the J22 string at WS2 |  |  | Return To Vendor | The FPC must be removed from the string and returned to the vendor for rework
208 | May-16-2022 | May-16-2022 | L2HE-CLNRM-FPCC-INSP | 6 | FPC COLD PART | CPI-FPCC-HE-013 | N/A | The ceramic is chipped. The damage was discovered during inspection of J22 FPCs at WS2 |  |  | Return To Vendor | The FPC must be removed from the string and sent back to the vendor for rework.
209 | May-16-2022 | May-16-2022 | L2HE-CLNRM-FPCC-INSP | 10 | FPC COLD PART | CPI-FPCC-HE-020 | N/A | Ceramic is cracked. This was discovered during the inspection of J22 at WS2 |  |  | Return To Vendor | The FPC must be removed from the string and returned to the vendor for rework
210 | May-26-2022 | May-31-2022 | L2HE-INSP-CAV | 44 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R062 | Residues | Coupler sapphire has two black dots on it. It did not clean up using 40psi compressed air. See attached picture for location. Unable to capture in photo. |  |  | Use As Is | Checked the stains, no electric effect noted by RF test. Should not cause any issue.
211 | May-31-2022 | May-31-2022 | L2HE-VTA-CAV-VTRF | 39 | 1.3GHZ LCLS-II RF CAVITY ASSY (BARE) | 361 | RFTestFailure | Cavity quenched at 17.7 MV/m in VTRF (NCR added retroactively) |  |  | Reject | This cavity is rejected for further work and will be delivered to SLAC as is.
212 | Jun-02-2022 | Apr-17-2025 | L2HE-CMA-VV-INSP | 5 | VESSEL LCLS-II AFTER MACHING WELDMENT | RFCHE-VV-019 | Dings Residues Scratches | Visual inspection performed on 1-Jun-2022. Scratches, surface imperfections and paint over spray on several flange/sealing surfaces. All photos on the server: M:\cmm\Pictures\Inspection Pictures\L2HE\Vacuum vessel\VV019\Receiving Inspection | Hold/Modify | Use the inspection results as a guideline, polish surfaces when feasible | Conforming | NCR reopened to change final disposition from use as is to conforming. Discrepancies repaired, see page 3, component is now conforming.
213 | Jun-03-2022 | Jun-16-2022 | L2HE-INSP-FWM | 2 | LCLS-II CM FARADAY PORT WINDOW | 01 | Dings Roughness Scratches | Scratches and dings noted on both sides of the window area. Surface finish exceeds 16u/in. See attached photos. | Hold/Modify | Return to vendor for repair. Will be re-inspected after repair. | Conforming | Vendor has corrected the surface finish and the window thickness is within drawing requirements.
214 | Jun-03-2022 | Jun-14-2022 | L2HE-INSP-FWM | 2 | LCLS-II CM FARADAY PORT WINDOW | 01 | Dimensional | Machined undercut area noted. Not machined per drawing. See attachment. |  |  | Conforming | After discussion with vendor and review of the drawing, this undercut / counterbore is acceptable since it is present in the stock blank flange that the drawing allows to be the starting material for manufacturing the port window. The traveler will be revised to note that this feature is acceptable in the final product.
215 | Jun-03-2022 | Jun-15-2022 | L2HE-INSP-FWM | 2 | LCLS-II CM FARADAY PORT WINDOW | 01 | OTHER | Not stamped/engraved per Dwg:"F10040909 - 1MM THICK" | Hold/Modify | Return to vendor for repair. Will be re-inspected after repair. | Conforming | Port Window SN01 is now etched in accordance with the drawing requirements.
216 | Jun-03-2022 | Jun-06-2022 | L2HE-INSP-FWM | 2 | FARADAY WINDOW MANIFOLD WELDMENT | 01 | Dimensional | (1)dimension is OOT; 2.80 +/- 0.040" = 2.750", OOT by .010". See attached report. |  |  | Use As Is | The vendor has been informed of dimensional out-of-tolerance for future weldments. Current oot acceptable as-is for use.
217 | Jun-06-2022 | Jun-07-2022 | L2HE-CLNRM-FPCC-INSP | 18 | FPC COLD PART | CPI-FPCC-HE-039 | N/A | after cleaning of first bag removed the bag and found two holes in second bag cleaned the bag and put cleanroom tape on bag |  |  | Return To Vendor | Assembly will be returned to vendor due to bellows damage
218 | Jun-06-2022 | Jun-06-2022 | L2HE-CLNRM-FPCC-INSP | 18 | FPC COLD PART | CPI-FPCC-HE-039 | N/A | FCP bellow bent, removed stiffening plates to see if there was tension causing the deformity but the bend stayed the same |  |  | Return To Vendor
219 | Jun-06-2022 | Jun-06-2022 | L2HE-CLNRM-FPCC-INSP | 18 | FPC COLD PART | CPI-FPCC-HE-040 | N/A | FPC bellow bent/twisted, removed stiffening plates to see if there was tension but nothing changed |  |  | Return To Vendor
220 | Jun-06-2022 | Jun-06-2022 | L2HE-CLNRM-FPCC-INSP | 18 | FPC COLD PART | CPI-FPCC-HE-039 | N/A | missing parts |  |  | Return To Vendor | This unit is being returned to the vendor due to bellows damage
221 | Jun-06-2022 | Jun-06-2022 | L2HE-CLNRM-FPCC-INSP | 18 | FPC COLD PART | CPI-FPCC-HE-040 | N/A | missing parts |  |  | Return To Vendor | This unit is being returned to the vendor due to bellows damage
222 | Jun-06-2022 | Jun-09-2022 | L2HE-VTA-CAV-VTRF | 63 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R048 | N/A | Qe1 slightly high, 2.6e10 |  |  | Use As Is | CTB has accepted this before In update of the spec to lift the upper boundary to 3.1e10 is in the making
223 | Jun-06-2022 | Sep-25-2023 | L2HE-VTA-CAV-VTRF | 64 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R062 | N/A | FE onset 19.5 MV/m. Qe1 also slightly low | Hold/Modify | Cavity should receive a "big" HPR (full disassembly) Please document on page 3 of this NCR. | Conforming | Cavity qualified on 2023-07-28
229 | Jun-08-2022 | Jun-08-2022 | L2HE-CLNRM-FPCC-INSP | 21 | FPC COLD PART | CPI-FPCC-HE-053 | OTHER | Small hole in one corner of inner bag |  |  | Use As Is | Particulate count will be verified at later inspection step
230 | Jun-08-2022 | Jun-08-2022 | L2HE-CLNRM-FPCC-INSP | 21 | FPC COLD PART | CPI-FPCC-HE-054 | OTHER | Small hole on one corner of the inner bag |  |  | Use As Is | Particulate count will be verified at later inspection step
231 | Jun-08-2022 | Jun-13-2022 | L2HE-CLNRM-FPCC-INSP | 22 | FPC COLD PART | CPI-FPCC-HE-055 | OTHER | Small hole on protruding point of inner bag |  |  | Use As Is | Particulate count will be verified at later inspection step
232 | Jun-08-2022 | Jun-08-2022 | L2HE-CLNRM-FPCC-INSP | 22 | FPC COLD PART | CPI-FPCC-HE-056 | OTHER | Small hole on protruding point of inner bag |  |  | Use As Is | Particulate count will be verified at later inspection step
233 | Jun-13-2022 | Mar-09-2023 | L2HE-INV-CAV | 39 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R072 | OTHER | RED light blinking on the accelerometer |  |  | Conforming | Cavity has passed all inspections and is fully qualified. No further action required.
234 | Jun-13-2022 | Mar-09-2023 | L2HE-INV-CAV | 40 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R073 | OTHER | RED light blinking on the accelerometer |  |  | Conforming | Cavity has passed all inspections and is fully qualified. No further action required.
235 | Jun-13-2022 | Mar-10-2023 | L2HE-INV-CAV | 41 |  | HE-R065 | OTHER | RED light blinking on the accelerometer |  |  | Conforming | Cavity passed all incoming inspections and RFIN. No further action required.
236 | Jun-13-2022 | Mar-09-2023 | L2HE-INV-CAV | 42 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R066 | OTHER | RED light blinking on the accelerometer |  |  | Conforming | Cavity has passed all inspections and is fully qualified. No further action required.
237 | Jun-13-2022 | Mar-10-2023 | L2HE-INV-CAV | 43 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R054 | OTHER | RED light blinking on the accelerometer |  |  | Conforming | Cavity passed all incoming inspections and RFIN. No further action required.
238 | Jun-13-2022 | Mar-09-2023 | L2HE-INV-CAV | 44 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R069 | OTHER | RED light blinking on the accelerometer |  |  | Conforming | Cavity has passed all inspections and is fully qualified. No further action required.
239 | Jun-15-2022 | Jun-28-2022 | L2HE-INSP-CAV | 45 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R054 | Dings | There are (3) small dings on the outer convolutions of the bellows. See attached photos. |  |  | Use As Is | CTB decision 2022-06-28
240 | Jun-16-2022 | Mar-09-2023 | L2HE-INV-CAV | 45 | SINGLE CELL LCLS-II CAVITY | SC-28-T129 | OTHER | VM Max exceeded threshold, 6.97g / 3.00g max. |  |  | Conforming | Cavity has passed all inspections and tests. No further action required.
241 | Jun-16-2022 | Mar-09-2023 | L2HE-INV-CAV | 46 | SINGLE CELL LCLS-II CAVITY | SC-31-T132 | OTHER | VM Max exceeded threshold, 6.97g / 3.00g max. |  |  | Conforming | Cavity has passed all inspections and tests. No further action required.
242 | Jun-16-2022 | Mar-09-2023 | L2HE-INV-CAV | 47 | SINGLE CELL LCLS-II CAVITY | SC-29-T130 | OTHER | VM Max exceeded threshold, 8.99g / 3.00g max. |  |  | Conforming | Cavity has passed all inspections and tests. No further action required.
243 | Jun-16-2022 | Mar-09-2023 | L2HE-INV-CAV | 48 | SINGLE CELL LCLS-II CAVITY | SC-30-T131 | OTHER | VM Max exceeded threshold, 8.99g / 3.00g max. |  |  | Conforming | Cavity has passed all inspections and tests. No further action required.
244 | Jun-16-2022 | Jun-16-2022 | L2HE-CLNRM-FPCC-INSP | 15 | FPC COLD PART | CPI-FPCC-HE-001 | Stains | Black marks were found on the ceramic |  |  | Return To Vendor | Return to vendor to get ceramic cleaned and re-coated
245 | Jun-17-2022 | Jun-17-2022 | L2HE-CLNRM-FPCC-INSP | 17 | FPC COLD PART | CPI-FPCC-HE-037 | Stains | Dark marks were found on the ceramic during inspection prior to string assembly |  |  | Return To Vendor
246 | Jun-17-2022 | Jun-17-2022 | L2HE-CLNRM-FPCC-INSP | 15 | FPC COLD PART | CPI-FPCC-HE-008 | N/A | Black marks were found on the ceramic during inspection before installation on the string |  |  | Return To Vendor
247 | Jun-17-2022 | Jun-17-2022 | L2HE-CLNRM-FPCC-INSP | 5 | FPC COLD PART | CPI-FPCC-HE-015 | N/A | Black marks were found on the ceramic during inspection at WS2 |  |  | Return To Vendor
248 | Jun-17-2022 | Jun-17-2022 | L2HE-CLNRM-FPCC-INSP | 6 | FPC COLD PART | CPI-FPCC-HE-013 | N/A | A chip on the ceramic was found during inspection at WS2 |  |  | Return To Vendor
249 | Jun-17-2022 | Jun-17-2022 | L2HE-CLNRM-FPCC-INSP | 14 | FPC COLD PART | CPI-FPCC-HE-017 | Stains | Dark marks were found on the ceramic during inspection in the cleanroom. Unit was taken off the J1.3-22 string. |  |  | Return To Vendor
250 | Jun-17-2022 | Jun-17-2022 | L2HE-CLNRM-FPCC-INSP | 12 | FPC COLD PART | CPI-FPCC-HE-018 | N/A | Dark marks were found on the ceramic in the cleanroom. Unit was taken off the J1.3-22 string |  |  | Return To Vendor
251 | Jun-17-2022 | Jun-17-2022 | L2HE-CLNRM-FPCC-INSP | 14 | FPC COLD PART | CPI-FPCC-HE-016 | N/A | Dark marks were found on the ceramic during inspection at WS2. The unit was removed from the J1.3-22 string |  |  | Return To Vendor | RTV to get the ceramic sandblasted and re-coated
252 | Jun-17-2022 | Jun-17-2022 | L2HE-CLNRM-FPCC-INSP | 10 | FPC COLD PART | CPI-FPCC-HE-021 | N/A | Dark marks were found on the ceramic during inspection at WS2. The unit was taken off the J1.3-22 string |  |  | Return To Vendor | RTV to get the ceramic snadblasted and recoated
253 | Jun-17-2022 | Jun-28-2022 | L2HE-INSP-CAV | 47 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R069 | Dings | There is one small ding on the outer convolution of the bellow. See attached photos. These dings are not acceptable I.A.W. the attached document. |  |  | Use As Is | CTB decision 2022-06-28
254 | Jun-23-2022 | Jun-29-2022 | L2HE-INSP-CAV | 49 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R073 | Dimensional | Parallelism of 2 phase Tee out of tolerance 0.3 mm |  |  | Use As Is | Can be compensated in transfer line
255 | Jun-23-2022 | Mar-10-2023 | L2HE-INV-CAV | 49 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R067 | OTHER | Accelerometer recorded 7.61g / 3.00g Max |  |  | Conforming | Cavity passed all incoming inspection and is fully qualified. No further action required.
256 | Jun-24-2022 | Jun-29-2022 | L2HE-INSP-CAV | 48 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R065 | Dimensional | DIM2: Parallelism of 2-phase TEE is OOT. See attached report. |  |  | Use As Is | Can be compensated in transfer line | JOHN V - Was it compensated for?
257 | Jun-27-2022 | Jun-29-2022 | L2HE-VTA-CAV-VTRF | 66 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R043 | N/A | Qext1 was 2.09e10, out of spec which is 1.1~1.9 e10 |  |  | Use As Is | Qext 1 spec has been updated in latest revision of traveler (1.1e10 to 3.1e10)
258 | Jun-27-2022 | Sep-25-2023 | L2HE-VTA-CAV-VTRF | 66 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R043 | N/A | Emax was below 23 MV/m | Hold/Re-measure | Cavity needs a full disassembly & rinse and re-test | Reject | CTB decision 2023-09-12 put in storage
259 | Jun-27-2022 | Sep-25-2023 | L2HE-VTA-CAV-VTRF | 66 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R043 | N/A | FE onset at 18 MV/m during final power rise. | Hold/Modify | Cavity needs a full disassembly & rinse and re-test please track HPR on page 3 of this traveler | Reject | CTB decision 2023-09-12 put in storage
260 | Jun-27-2022 | Dec-03-2024 | L2HE-VTA-CAV-VTRF | 67 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R037 | N/A | Qext2 was 1.19e11, out of spec which is 7.5 e11~2.5 e12 | Hold/Modify | Please replace Pick-up probe AND high-Q antenna in clean room and track progress on page 3 of this NCR | Conforming | Cavity passes after repair.
261 | Jun-27-2022 | Sep-19-2024 | L2HE-VTA-CAV-VTRF | 68 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R054 | N/A | Emax was 14.9 MV/m, out of spec which is >=23MV/m. |  |  | Conforming | Cavity rejected by CTB 2022-06-28 Cavity returned to vendor for reprocessing. Passed vertical test on 2024-03-08 as HE-CAVR054_RP
262 | Jun-29-2022 | Jul-05-2022 | L2HE-CLNRM-FPCC-INSP | 24 | FPC COLD PART | CPI-FPCC-HE-022 | OTHER | Hole in middle of bag, went through to inner bag as well. Hole was on back side of bag, opposite side of the label stating the pair number. Hole is approximately 1in x 1in if not smaller. Tech(s) noticed hole when doing initial clean, taped over hole on inner bag before sending into the clean room. |  |  | Use As Is | External particulate count later in the inspection process will determine whether this unit is usable
263 | Jun-29-2022 | Jul-05-2022 | L2HE-CLNRM-FPCC-INSP | 24 | FPC COLD PART | CPI-FPCC-HE-023 | OTHER | Hole in middle of bag, went through to inner bag as well. Hole was on back side of bag, opposite side of the label stating the pair number. Hole is approximately 1in x 1in if not smaller. Tech(s) noticed hole when doing initial clean, taped over hole on inner bag before sending into the clean room. |  |  | Use As Is | External particulate count later in the inspection process will determine whether this unit is usable
264 | Jul-05-2022 | Jul-05-2022 | L2HE-CLNRM-FPCC-INSP | 25 | FPC COLD PART | CPI-FPCC-HE-059 | OTHER | Small hole on inner bag only |  |  | Use As Is | Cleanliness will be verified by particulate count at a later point in the inspection process
265 | Jul-05-2022 | Jul-05-2022 | L2HE-CLNRM-FPCC-INSP | 25 | FPC COLD PART | CPI-FPCC-HE-060 | OTHER | Small hole on inner bag only |  |  | Use As Is | Cleanliness will be verified by particulate count at a later point in the inspection process
266 | Jul-07-2022 | Sep-12-2023 | L2HE-VTA-CAV-VTRF | 73 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R065 | N/A | Cavity produced measurable radiation brought on by mode mixing induced multipacting at maximum of 0.033mem/hr at 28.25 MV/m | Hold/Modify | re-rinse & re-test CTB 2022-07-19 | Conforming | Met spec in last test on 2023-09-12
267 | Jul-08-2022 | Dec-19-2022 | L2HE-VTA-CAV-VTRF | 74 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R069 | N/A | Cavity mode mixing started multipacting with radiation above 23 MV/m and quench. A second clean run was done to 23 MV/m without multipacting, but Qo was below spec. due to previous quench. | Hold/Re-measure | re-rinse & re-test CTB 2022-07-19 | Conforming | Cavity successfully tested on 2022-07-20
268 | Jul-10-2022 | Jul-11-2022 | L2HE-INSP-FWM | 3 | FARADAY WINDOW MANIFOLD WELDMENT | 07 | Dimensional | F10045762-SN-07: (1) OOT; 2.80"±.040" = 2.723 |  |  | Conforming | Dimensional out-of-tolerance NCR for this dimension (distance from beamline center to reducing tee flange) not required per new revision of the INSP-FWM traveler.
269 | Jul-11-2022 | Aug-11-2022 | L2HE-INSP-QUAD | 2 | LCLS-II MAGNET PACKAGE ASSEMBLY | 203 | InspectionFailure | Will be a NCR submitted for section 3.2.4, 100Hz Q-factor is slightly out of range. Spec requirement is .3-.4. At time of inspection reading of Q-FACTOR was .263 | Hold/Re-measure | Retest and record the Q factor at 100 Hz for thequadrupole winding. | Use As Is | Our electrical inspection results indicate that the measured value of the Q factor at 100 Hz (0.266 (from test #1) and 0.263 (from test #2)) for the SPQA 203 magnet is marginally below the lower bound (0.3) of the anticipated value. Note that these limits were set according to the data, rounded off to one decimal place, obtained from measurements at the manufacturer site and Fermi Lab. Since the room temperature resistance of the coils and magnet assembly and the Q factor at 1 kHz fall within the suggested range, the observed discrepancy in the Q value at 100 Hz can be attributed to a statistical error in measurements. Nonetheless, it is expected to have no measurable impact on magnet performance at the actual operating conditions.
270 | Jul-11-2022 | Aug-05-2022 | L2HE-CLNRM-FPCC-INSP | 23 | FPC COLD PART | CPI-FPCC-HE-057 | N/A | HE057 has Brown discoloration around ceramic surface |  |  | Use As Is | SMEs from partner labs were consulted and the decision was to use this part as is.
271 | Jul-11-2022 | Jul-26-2022 | L2HE-CLNRM-FPCC-INSP | 24 | FPC COLD PART | CPI-FPCC-HE-023 | Scratches | FPC023-BLACK SCUFF'S ON CERAMIC SURFACE |  |  | Return To Vendor
272 | Jul-13-2022 | Aug-11-2022 | L2HE-INSP-QUAD | 3 | LCLS-II MAGNET PACKAGE ASSEMBLY | 204 | InspectionFailure | section 3.2.4, 100Hz Q-factor is slightly out of range. Spec requirement is .3-.4. At time of inspection reading of Q-FACTOR was .263 | Hold/Re-measure | Retest and record the Q factor at 100 Hz for the quadrupole winding. | Use As Is | Our electrical inspection results indicate that the measured value of the Q factor at 100 Hz (0.263 (from test #1) and 0.264 (from test #2)) for the SPQA 204 magnet is marginally below the lower bound (0.3) of the anticipated value. Note that these limits were set according to the data, rounded off to one decimal place, obtained from measurements at the manufacturer site and Fermi Lab. Since the room temperature resistance of the coils and magnet assembly and the Q factor at 1 kHz fall within the suggested range, the observed discrepancy in the Q value at 100 Hz can be attributed to a statistical error in measurements. Nonetheless, it is expected to have no measurable impact on magnet performance at the actual operating conditions.
273 | Jul-15-2022 | Aug-11-2022 | L2HE-INSP-BPM | 4 | BPM HOUSING W/ CU PLATING | 104 | Plating | Some areas of plating located on edges look thinner or perhaps deteriorated. Other questionable areas of copper appear to be just shinier than other spots. Please see attached. Cross-hatch tape test shows no sign of flaking/peeling. |  |  | Use As Is | Spoke with FNAL and SLAC, they okayed this for use. There is no flaking and scratches are only present on feedthrough flange.
274 | Jul-15-2022 | Aug-11-2022 | L2HE-CLNRM-FPCC-INSP | 26 | FPC COLD PART | CPI-FPCC-HE-039 | N/A | not fully assembled. RF connection pin assembly missing from both HEFPC039 and HEFPC040 | Hold/Modify | Replace screw and spring-pin once they are shipped from the vendor | Conforming | The FPC now has the screw installed
275 | Jul-15-2022 | Jul-20-2022 | L2HE-VTA-CAV-VTRF | 76 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R073 | N/A | Qext2 was 3.95e12, out of spec which is 7.5 e11 ~ 2.5 e12 |  |  | Use As Is | Accepted by CTB 2022-07-19
276 | Jul-15-2022 | Dec-19-2022 | L2HE-VTA-CAV-VTRF | 77 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R068 | N/A | Final power rise, FE onset at 12 MV/m. | Hold/Modify | re-rinse & re-test CTB 2022-07-19 | Reject | HOM hook scratched -- CTB 2022-12-20
277 | Jul-21-2022 | Mar-09-2023 | L2HE-INV-CAV | 53 | SINGLE CELL LCLS-II CAVITY | SC-32-T133 | OTHER | VM Max (8.19g) exceeds the Program Alert Max (3.00g) |  |  | Conforming | Cavity has passed all inspections and tests. No further action required.
278 | Jul-21-2022 | Mar-09-2023 | L2HE-INV-CAV | 54 | SINGLE CELL LCLS-II CAVITY | SC-34-T135 | OTHER | VM Max (8.19g) exceeds the Program Alert Max (3.00g) |  |  | Conforming | Cavity has passed all inspections and tests. No further action required.
279 | Jul-21-2022 | Mar-09-2023 | L2HE-INV-CAV | 55 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R070 | OTHER | VM Max (6.35g) exceeds the Program Alert Max (3.00g) |  |  | Conforming | Cavity has passed all inspections and tests. No further action required.
280 | Jul-21-2022 | Sep-27-2022 | L2HE-CMA-UCM-INSP | 5 | LCLS-II UPPER COLD MASS ASSEMBLY | RFCHE-UC-019 | Dings | 5K flange damaged. Please see attached. |  |  | Use As Is | Having exchanged with CMA lead, we will decide whether to polish the 5K flange when the cryomodule is at the final testing site. | GARY - Need final resolution as final disposition, this comment is more for initial disposition.
281 | Jul-22-2022 | Jul-26-2022 | L2HE-INSP-BLBS | 5 | WELDMENT BELLOWS SHORT | 001 | Dings Scratches | Blemish/scuff mark on bellows exterior. Length of feature is roughly 3/4 inch. After closer look with 4x loop is appears very minor. Please see attached. |  |  | Use As Is | Scuffs on outside of convolutions do not effect functionality of bellows as long as they are not dented.
282 | Jul-26-2022 | Mar-28-2023 | L2HE-CLNRM-FPCC-INSP | 18 | FPC COLD PART | CPI-FPCC-HE-040 | Scratches | When probe protective cover being removed for installation onto cavity technician scrapped probe tip of FPC #40 | Hold/Re-measure |  | Use As Is | The attached PDF shows copper plating defects on SNS-PPU FPCs which are similar to those found here. The PPU FPCs were RF tested with these defects and no issues were found. As such, no issues are expected on this FPC.
283 | Jul-29-2022 | Jul-29-2022 | L2HE-INSP-CAV | 55 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R081 | OTHER | Field Probe Pin appears to have some plastic shavings in the Field Probe (HEp2101) pin receptacle area. See attached. |  |  | Use As Is
284 | Jul-29-2022 | Sep-25-2023 | L2HE-VTA-CAV-VTRF | 80 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R074 | N/A | There were some mild FE at high gradient above 24 MV/m | Hold/Modify | HPR and re-test CTB 2022-08-02 replace high-Q antenna | Reject | CTB 2023-08-01 -- sent to FNAL for HF rinse
285 | Jul-29-2022 | Aug-18-2022 | L2HE-VTA-CAV-VTRF | 81 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R067 | N/A | There was a little radiation just above FE threshold at high gradient close to quenching. | Hold/Modify | HPR and re-test | Conforming | Thanks Peter & cleanroom/chem room staff!
286 | Jul-29-2022 | Aug-03-2022 | L2HE-INSP-CAV | 55 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R081 | Dimensional | Dimension 3 (ANGL6) on CMM report is Out of Tolerance (OOT). See attached report. |  |  | Use As Is
287 | Aug-01-2022 | Aug-01-2022 | L2HE-CLNRM-FPCC-INSP | 27 | FPC COLD PART | CPI-FPCC-HE-001 | InspectionFailure Scratches | HE_FPCC_001 has erosion/scratches on the RF probe |  |  | Return To Vendor
288 | Aug-01-2022 | May-27-2023 | L2HE-VTA-CAV-VTRF | 82 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R080 | N/A | Final power rise, FE onset 12 MV/m | Hold/Modify | HPR and re-test | Conforming | Cavity meets spec after last test
290 | Aug-01-2022 | Aug-01-2022 | L2HE-CLNRM-FPCC-INSP | 27 | FPC COLD PART | CPI-FPCC-HE-008 | InspectionFailure Scratches | HE_FPC_008 has a scratch the goes 360* around the probe |  |  | Return To Vendor
291 | Aug-02-2022 | Aug-02-2022 | L2HE-CWI-MCSHLD | 1 | BARATRON TRANSDUCER 0-1000 | 023247253 | OTHER | 80 50K Shields were inspected and 64 of had a groove welds that were undersized and were rejected and returned to the vendor for rework. |  | 84 50K Shields were inspected and 64 of had a groove welds that were undersized and were rejected and returned to the vendor for rework. | Return To Vendor | Having communicated with Fermi Lab and the shield vendor, it is decided that 64 out of the 80 shields at JLAB should be returned for repair and 16 shields to keep at JLAB for in-house weld repair. See attached spreadsheets for the sets to RTV and keep. A weld repair sketch is also attached for information. Also refer to JLAB Shipping Authorization Form #134987 onshipping of the 64 shields.
292 | Aug-05-2022 | Feb-20-2024 | L2HE-CLNRM-FPCC-INSP | 31 | FPC COLD PART | CPI-FPCC-HE-021 | Scratches | FPC SN 021-Scratch at probe tip on side. Also small gray stains on shaft | Hold/Modify |  | Reject | Hold until further notice. Clean/repair scratches only if the unit is required for a string
293 | Aug-07-2022 | Aug-08-2022 | L2HE-INSP-CAV | 58 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R082 | OTHER | Water was noticed inside HOM Probe B - long pin area. See attached photos. |  |  | Use As Is | This is acceptable. Cavity was stuck in customs/DC for about 18 days. RF incoming should show if there is a deviation/leak.
294 | Aug-07-2022 | Aug-08-2022 | L2HE-INSP-CAV | 58 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R082 | Dimensional | HE-R082 CMM inspection performed. (1) DIM2 is OOT. See attached report. |  |  | Use As Is | Should not affect string/module assembly.
295 | Aug-08-2022 | Mar-21-2023 | L2HE-INSP-QUAD | 4 | LCLS-II MAGNET PACKAGE ASSEMBLY | 205 | InspectionFailure | There will be a NCR on section 3.2.4, Q-factor is slightly out of tolerance .262 @100Hz. Spec is .3-.4 | Hold/Re-measure | Retest and record the Q factor at 100 Hz for the quadrupole winding. | Use As Is | Our electrical inspection results indicate that the measured value of the Q factor at 100 Hz (0.262 (from test #1) and 0.266 (from test #2)) for the SPQA 205 magnet is marginally below the lower bound (0.3) of the anticipated value. Note that these limits were set according to the data, rounded off to one decimal place, obtained from measurements at the manufacturer site and Fermi Lab. Since the room temperature resistance of the coils and magnet assembly and the Q factor at 1 kHz fall within the suggested range, the observed discrepancy in the Q value at 100 Hz can be attributed to a statistical error in measurements. Nonetheless, it is expected to have no measurable impact on magnet performance at the actual operating conditions.
296 | Aug-08-2022 | Aug-12-2022 | L2HE-INSP-BLXD | 4 | WELDMENT EXTENSION-DS | 006 | Oxidation Plating | (1)Dark spot >1/8". See attached photos. | Hold/Modify | Send to chemistry for cleaning | Conforming | Plating is now clean
297 | Aug-08-2022 | Aug-12-2022 | L2HE-INSP-BLXD | 4 | WELDMENT EXTENSION-DS | 006 | Oxidation Plating | Oxidation shown on plating. See attached photos. | Hold/Modify | Send to chemistry for cleaning | Conforming | Plating is now clean
298 | Aug-09-2022 | Aug-12-2022 | L2HE-INSP-BLXD | 5 | WELDMENT EXTENSION-DS | 005 | Plating | Dark areas shown on plating >1/8" diameter. See attached photos. | Hold/Modify | Bring to chemistry for cleaning | Conforming | Plating is now clean
299 | Aug-09-2022 | Aug-12-2022 | L2HE-INSP-BLXD | 5 | WELDMENT EXTENSION-DS | 005 | Plating Scratches | Scratches noticed on plating in one area on rotating flange end. See attached photo. |  |  | Use As Is | Visually inspected, could not find scratch. If still present, it's superficial and not through the plating.
300 | Aug-09-2022 | Sep-09-2022 | L2HE-INSP-BLBS | 6 | WELDMENT BELLOWS SHORT | 006 | Plating | Copper plating has excessive spots/stains present. Please see attached. | Hold/Modify | Send to chemistry for cleaning | Return To Vendor | Return to vendor for replating
301 | Aug-11-2022 | Mar-21-2023 | L2HE-INSP-QUAD | 5 | LCLS-II MAGNET PACKAGE ASSEMBLY | 206 | InspectionFailure | Section 3.2.4 Q-Factor @100Hz out of traveler specification. | Hold/Re-measure | Retest and record the Q factor at 100 Hz for the quadrupole winding. | Use As Is | Our electrical inspection results indicate that the measured value of the Q factor at 100 Hz (0.261 (from test #1) and 0.264 (from test #2)) for the SPQA 206 magnet is marginally below the lower bound (0.3) of the anticipated value. Note that these limits were set according to the data, rounded off to one decimal place, obtained from measurements at the manufacturer site and Fermi Lab. Since the room temperature resistance of the coils and magnet assembly and the Q factor at 1 kHz fall within the suggested range, the observed discrepancy in the Q value at 100 Hz can be attributed to a statistical error in measurements. Nonetheless, it is expected to have no measurable impact on magnet performance at the actual operating conditions.
302 | Aug-11-2022 | Aug-15-2022 | L2HE-INSP-BLBP | 25 | WELDMENT BELLOWS PRCM | 034 | Dimensional | PDF attached. Dimensional non-conforming. Length out of tolerance 1.12 mm and bellow ID out of tolerance 0.9mm |  |  | Use As Is | OOT values will not effect functionality of bellows
303 | Aug-12-2022 | Apr-29-2025 | L2HE-INSP-CAV | 59 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R086 | OTHER | Tapped holes damaged on bellows braces (both sides of cavity). Holes failed go/nogo thread gage check (5/16 -18). Please see pictures attached. | Hold/Modify | Could you (or someone from John F's group) please try to re-tap the holes? | Conforming | Many thanks for the rework, Earl!
304 | Aug-15-2022 | Aug-16-2022 | L2HE-INSP-CAV | 60 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R088 | Dings | Dings/pits present on bellows. Please see attached pictures. |  |  | Use As Is | CTB decision on 2022-08-16 use as is
305 | Aug-15-2022 | Oct-13-2022 | L2HE-INSP-BLBP | 25 | WELDMENT BELLOWS PRCM | 034 | Oxidation Plating | (1)dark spot on plating = 1/8"x1/4" shape. See attached photos. | Hold/Modify | Send to chemistry for cleaning | Use As Is | Small discoloration, likely not oxidation, no flaking observed.
306 | Aug-15-2022 | Aug-16-2022 | L2HE-INSP-CAV | 60 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R088 | Dimensional | Please see attached PDF. |  |  | Use As Is | Can be compensated during module assembly.
307 | Aug-18-2022 | Oct-11-2022 | L2HE-INSP-BLBP | 27 | WELDMENT BELLOWS PRCM | 037 | Scratches | Scratches thru seal of rotatable flange. Pictures attached. | Hold/Modify | Send to machine shop for seal surface polishing | Conforming | Seal surfaces conforming after polishing
308 | Aug-18-2022 | Sep-13-2022 | L2HE-INSP-BLBP | 28 | WELDMENT BELLOWS PRCM | 035 | Dings | Protrusion: from ding on rotating flange end, interior plating. |  |  | Return To Vendor | Return to vendor for replating due to blistering
309 | Aug-18-2022 | Sep-09-2022 | L2HE-INSP-BLBP | 28 | WELDMENT BELLOWS PRCM | 035 | Plating | Blistering noted on rotating flange end. |  |  | Return To Vendor | Return to vendor for replating due to blisters
310 | Aug-18-2022 | Sep-13-2022 | L2HE-INSP-BLBP | 28 | WELDMENT BELLOWS PRCM | 035 | Scratches | Scratches noted in seal area on rotating flange. |  |  | Return To Vendor | Return to vendor for replating due to blistering
311 | Aug-18-2022 | Sep-13-2022 | L2HE-INSP-BLBP | 28 | WELDMENT BELLOWS PRCM | 035 | Plating Scratches | Scratches noted in plating longer than 1/8". |  |  | Return To Vendor | Return to vendor for replating due to blistering
312 | Aug-18-2022 | Oct-11-2022 | L2HE-INSP-BLBP | 29 | WELDMENT BELLOWS PRCM | 038 | Scratches | Scratches thru seal area of rotatable flange. Picture attached. | Hold/Modify | Send to machine shop for polishing | Conforming | Seal surfaces conforming after polishing
313 | Aug-20-2022 | Aug-22-2022 | L2HE-INSP-CAV | 62 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R085 | Dings | Bellows have three areas of damage (clamp pin side). 1) Ding/nick on OD convolution. 2) Side wall indentation, straight edge shows about .01" gap. 3) Ding/nick on OD convolution. All photos on server: M:\cmm\Pictures\Inspection Pictures\L2HE\Dressed Cavities\085 |  |  | Use As Is | These damages had been discussed during our visit to RI and deemed acceptable by the CTB
314 | Aug-24-2022 | Aug-24-2022 | L2HE-CLNRM-FPCC-INSP | 32 | FPC COLD PART | CPI-FPCC-HE-061 | OTHER | Large hole on inner bag in bottom corner |  |  | Use As Is | Cleanliness will be determined during particulate count
315 | Aug-24-2022 | Aug-24-2022 | L2HE-CLNRM-FPCC-INSP | 32 | FPC COLD PART | CPI-FPCC-HE-062 | OTHER | Large hole on inner bag in bottom corner |  |  | Use As Is | Cleanliness will be determined during particulate count
316 | Aug-24-2022 | Oct-04-2022 | L2HE-INSP-CAV | 63 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R084 | Dings | Burst Disc has a scratch and indentation on the disc surface. | Hold/Re-measure | Temporarily "use as is" -- we have to see what the leak check tells us about this. | Conforming | Leak check on CAVR084 was ok (traveler # 23 L2HE-CLNRM-CAV-TSTD5)
317 | Aug-25-2022 | Aug-25-2022 | L2HE-CLNRM-FPCC-INSP | 33 | FPC COLD PART | CPI-FPCC-HE-063 | N/A | Large hole in bottom corner of inner bag |  |  | Use As Is | Cleanliness will be check during the particulate count
318 | Aug-25-2022 | Aug-30-2022 | L2HE-VTA-CAV-VTRF | 88 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R037 | N/A | After quench, cavity failed to pass Q0 spec at 20.8MV/m, 2.3e10 |  |  | Conforming | Cavity meets spec. (Since no FE or MP was present, the cavity qualified in the first ramp up.)
319 | Aug-25-2022 | Aug-30-2022 | L2HE-VTA-CAV-VTRF | 89 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R086 | N/A | Cavity quenches before 23 MV/m, passband quench gradients recorded. |  |  | Use As Is | CTB decision 2022-08-30 Combine in string with stronger cavities (See CAVR011/12)
320 | Aug-29-2022 | Dec-06-2022 | L2HE-CLNRM-FPCC-INSP | 34 | FPC COLD PART | CPI-FPCC-HE-065 | Oxidation | Upon removing the cap discoloration was observed with a purple spot on the copper above the ceramic near the RF pin, and general purple discoloration around the copper base. |  |  | Use As Is | Discoloration does not affect FPC performance
321 | Aug-29-2022 | Sep-06-2022 | L2HE-CLNRM-FPCC-INSP | 33 | FPC COLD PART | CPI-FPCC-HE-063 | OTHER Stains | Tech J Casiano The copper around the top side looks like oxidation on top and bottom of the ceramic tower |  |  | Return To Vendor
322 | Sep-01-2022 | Dec-19-2022 | L2HE-VTA-CAV-VTRF | 91 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R084 | N/A | rad onset at 24 MV/m brought on with strong mode mixing, max rad 57mrad/hr at 30 MV/m. | Hold/Modify | HPR and re-test | Conforming | Successful test on 2022-10-07
323 | Sep-01-2022 | Jul-11-2023 | L2HE-VTA-CAV-VTRF | 92 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R085 | N/A | radiation onset at 12.9 MV/m. max rad before quench at 23.5 MV/m, 173mrem/hr. | Hold/Modify | HPR and re-test | Conforming | Test on 2022-06-05
324 | Sep-01-2022 | Jul-11-2023 | L2HE-VTA-CAV-VTRF | 93 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R087 | N/A | Rad onset at 14.5 MV/m and ran to 1003mrem/hr at 27.5 MV/m, and set off Neutron Alarm during pulse processing. | Hold/Modify | HPR and re-test | Conforming | Test on 2023-06-01
325 | Sep-13-2022 | Sep-14-2022 | L2HE-INSP-CAV | 66 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R095 | Dimensional | Please see attached PDF for dimensional non-conforming. Parallelism and perpendicularity are out of tolerance on page 1. |  |  | Use As Is | Should not cause an issue -- checked back with John F
326 | Sep-15-2022 | Sep-27-2022 | L2HE-CLNRM-FPCC-INSP | 34 | FPC COLD PART | CPI-FPCC-HE-065 | Scratches Stains | FPC 65 has discoloration around it, it looks more orange than normal and has two small scratches around the antenna |  |  | Use As Is | The issues will not affect the operation of the FPC
327 | Sep-22-2022 | Sep-23-2022 | L2HE-INSP-CAV | 69 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R091 | Dimensional | Unable to demagnetize HE-R091 in the Tee-transition joint; Should be: .65 Gauss, Actual= 2.4 Gauss. | Hold/Modify | Gary wants to check in on you (tomorrow) and move the cavity to the North annex to use the Maurer degausser on the T. | Use As Is | Proceed with vertical test. Many thanks to Earl, Aaron and Gary for the effort!
328 | Sep-26-2022 | Sep-27-2022 | L2HE-INSP-CAV | 70 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R098 | OTHER | One 5/16-18 tapped hole in Lever Tuner Adapter Plate (FP side) was not tapped correctly. The Hex bolt threads are damaged. The Coupler and FP downstream bolts are removed for CMM inspection. Coupler threads are ok. | Hold/Modify | Please try to re-tap the thread and replace the screw If no quick fix is possible, we have to discuss with John F. | Conforming | Many thanks Earl!
329 | Sep-26-2022 | Sep-26-2022 | L2HE-CLNRM-FPCC-INSP | 35 | FPC COLD PART | CPI-FPCC-HE-068 | InspectionFailure Stains | Black stains on ceramic. RF Pin hole has foreign object debris. |  |  | Use As Is | Issues will not affect the operation of the FPC
330 | Sep-26-2022 | May-27-2023 | L2HE-VTA-CAV-VTRF | 94 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R095 | N/A | Cavity FE onset at 17.8 MV/m. Peak radiation was 1R/hr | Hold/Modify | re-rinse & re-test (small HPR -- bottom flange drop) CTB 2022-09-28 | Use As Is | CTB 2023-05-23 Cavity accepted as marginal
331 | Sep-26-2022 | Jul-11-2023 | L2HE-VTA-CAV-VTRF | 96 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R093 | N/A | FE onset 24 MV/m | Hold/Modify | re-rinse & re-test (small HPR -- bottom flange drop) CTB 2022-09-28 | Conforming | Cavity qualified at partner lab
332 | Sep-28-2022 | Sep-28-2022 | L2HE-INSP-CAV | 71 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R099 | Dimensional | Dimension DIM2 is OOT: Parallelism of 2-phase TEE. Should Be: 1.0mm max, Is: 1.4mm |  |  | Use As Is | Should not cause an issue -- checked back with John F -- HE-CAVR035
333 | Sep-29-2022 | Oct-03-2022 | L2HE-INSP-FWM | 4 | LCLS-II CM FARADAY PORT WINDOW | 08 | Dimensional | Please see attached for dimensional non-conforming. F10045762 SN04 small conflat flange out of tolerance. |  |  | Use As Is | The location of the small conflat flange is acceptable as is. This dimension will not impact string or cryomodule build. The weldment is removed from the cryomodule after acceptance testing.
334 | Oct-04-2022 | Oct-05-2022 | L2HE-INSP-CAV | 73 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R108 | Dimensional | Please see attached for dimensional non-conforming. |  |  | Use As Is | Should not cause an issue -- checked back with John F
335 | Oct-07-2022 | Mar-23-2023 | L2HE-CLNRM-AMGV-LEAK | 7 | ALL METAL GATE VALVE | 0030 | InspectionFailure Residues Stains | There is a dark residue on the triangle stamp side of gate valve. The other side of the gate valve has like light yellow stains on it. | Hold/Re-measure | The project has decided to use the discolored valves in their as-is condition. This AMGV will receive a standard leak check in a following traveler | Use As Is | The project has decided to use the discolored valves in their as-is condition. This valve will proceed through the AMGV regular leak check procedure as if it was not stained.
336 | Oct-09-2022 | May-19-2023 | L2HE-INSP-BLBU | 7 | WELDMENT BELLOWS - US END | 005 | Dings | Bellows have dings in two areas. Possibly from fixturing. NCR generated. See attached photos. |  |  | Reject | Vendor will not accept back, reject for bellows dings
337 | Oct-09-2022 | May-19-2023 | L2HE-INSP-BLBU | 8 | WELDMENT BELLOWS - US END | 006 | Dings | Bellows have dings in two areas. Possibly from fixturing. See attached photos. |  |  | Reject | Vendor won't accept back, reject for bellows dings
338 | Oct-09-2022 | Dec-08-2022 | L2HE-INSP-BLXD | 6 | WELDMENT EXTENSION-DS | 007 | Scratches | Scratches on Rotating seal area. | Hold/Modify | Send to machine shop for polishing of seal surfaces | Conforming | Seal surfaces now conforming
339 | Oct-09-2022 | Apr-17-2025 | L2HE-INSP-BLXD | 6 | WELDMENT EXTENSION-DS | 007 | Stains | Some dark spots exceed 1/8" on plating. | Hold/Modify | Send to chemistry for cleaning after taking to machine shop for polishing | Conforming | NCR reopened to change from use as is to conforming. Copper plating is clean.
340 | Oct-09-2022 | Dec-08-2022 | L2HE-INSP-BLXD | 7 | WELDMENT EXTENSION-DS | 009 | Scratches | Scratches on Rotating seal area. | Hold/Modify | Please send to the machine shop for polishing of seal surfaces | Conforming | Seal surfaces conforming
341 | Oct-09-2022 | Mar-30-2023 | L2HE-INSP-BLXD | 7 | WELDMENT EXTENSION-DS | 009 | Stains | Some dark areas exceed 1/8" on plating. Marks noted. Plating does not look scratched. | Hold/Modify | Please send to chemistry for cleaning | Conforming | Copper plating is clean
342 | Oct-10-2022 | May-27-2023 | L2HE-VTA-CAV-VTRF | 100 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R099 | RFTestFailure | Multiple mod mixing events, Qo post quench event at 19 MV/m, not radiation free at the end of the VTA tests. | Hold/Modify | After 2022-12-20 test re-rinse & re-test (big HPR -- full disassembly) CTB 2023-01-10 re-rinse & re-test (small HPR -- bottom flange drop) | Conforming | Cavity now qualified after re-treatment (HF rinse) at Argonne
343 | Oct-11-2022 | Oct-13-2022 | L2HE-INSP-BPM | 4 | BPM HOUSING W/ CU PLATING | 104 | Dings | Divot/depression in seal path. | Hold/Modify | Sent to machine shop for polishing | Conforming | Seal surfaces now conforming
344 | Oct-11-2022 | Oct-13-2022 | L2HE-INSP-BPM | 4 | BPM HOUSING W/ CU PLATING | 104 | Dings | Ding/raised edge on flange. Picture attached. | Hold/Modify | Sent to machine shop for smoothing | Conforming | Seal surfaces now conforming
345 | Oct-11-2022 | Oct-12-2022 | L2HE-INSP-FWM | 5 | LCLS-II CM FARADAY PORT WINDOW | 07 | Dimensional | Weldment (F10045762) SN05 dimensional non-conforming. PDF attached. |  |  | Use As Is | Dimensions will not impact cavity string assembly. Faraday Window Manifold is removed after cryomodule testing and is not provided to the customer.
346 | Oct-11-2022 | Oct-28-2022 | L2HE-INSP-BLBP | 30 | WELDMENT BELLOWS PRCM | 075 | Plating | Small blisters noted on internal copper coating. Both flange areas and all convolutions have some pitting. |  |  | Return To Vendor | Blisters on plating require removal and replating
347 | Oct-12-2022 | Oct-28-2022 | L2HE-INSP-BLBS | 7 | WELDMENT BELLOWS SHORT | 012 | Plating Stains | One dark spot exceeds "dark spots over 1/8" diameter" on the plating, 4th convolution from Rotating flange. | Hold/Modify | Please send to Chemistry for cleaning | Conforming | Copper plating was cleaned and is now conforming
348 | Oct-17-2022 | Oct-17-2022 | L2HE-INSP-CAV | 74 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R105 | Scratches | Scratch noted on burst disc on sealing surface. See attached photos. |  |  | Use As Is | We had this before, as long as the RF incoming is ok (and the leak check) this is tolerable.
349 | Oct-20-2022 | Jan-17-2023 | L2HE-INSP-BPM | 5 | BPM HOUSING W/ CU PLATING | 110 | Stains | Dark spots/stains present. Picture attached. | Hold/Modify | Please send to chemistry to cleaning after sending to machine shop for seal surface polishing. | Use As Is | Some small oxidation spots acceptable, has been cleaned by chemistry
350 | Oct-20-2022 | Dec-08-2022 | L2HE-INSP-BPM | 5 | BPM HOUSING W/ CU PLATING | 110 | Scratches | Pictures attached. | Hold/Modify | Please send to machine shop for seal surface polishing, then send to chemistry for plating cleaning. | Conforming | Seal surfaces now conforming
351 | Oct-20-2022 | Sep-24-2024 | L2HE-VTA-CAV-VTRF | 103 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R108 | RFTestFailure | HE-R108 quench limited at 14.3 MV/m |  |  | Conforming | CTB 2022-10-25: Set aside -- can be stored at Bluecrab Cavity returned to vendor for reprocessing. Passed vertical test on 2024-01-26 as HE-CAVR108_RP. Sent to FNAL for use in their string on 2024-03-21.
352 | Oct-24-2022 | Jan-17-2023 | L2HE-INSP-FWM | 8 | LCLS-II CM FARADAY PORT WINDOW | 04 | Scratches | Scratch present on faraday manifold flange. Please see attached. | Hold/Modify | Lap flange to remove scratch and re-inspect flange. | Conforming | Part was repaired and subsequently passed inspection.
353 | Oct-26-2022 | Oct-26-2022 | L2HE-INSP-BLBP | 31 | WELDMENT BELLOWS PRCM | 039 | Dimensional | Four dimensions are nonconforming. See attached report and traveler data for OOT conditions. |  |  | Use As Is | Out of tolerance values are small and are on values that won't effect bellows performance.
354 | Oct-27-2022 | Oct-31-2022 | L2HE-INSP-FWM | 9 | LCLS-II CM FARADAY PORT WINDOW | 06 | Scratches | Scratches on Weldment (F10045762) flange. Scratches on Port Window flange(F10040909). Pictures attached below. |  |  | Return To Vendor | The window thickness is under the specification and has scratches. This window cannot be repaired to meet the specifications and will be sent back to the vendor.
355 | Oct-28-2022 | Oct-31-2022 | L2HE-INSP-FWM | 9 | LCLS-II CM FARADAY PORT WINDOW | 06 | Dimensional | Dimensional non-conforming. Please see attached PDF. |  |  | Return To Vendor | The window thickness is under the specification and has scratches. This window cannot be repaired to meet the specifications and will be sent back to the vendor.
356 | Oct-28-2022 | Oct-31-2022 | L2HE-INSP-FWM | 10 | LCLS-II CM FARADAY PORT WINDOW | 02 | Scratches | Scratches on Port Window (F10040909). Picture attached. |  |  | Use As Is | Window thickness is on the high end of the specification and surface finish is within specification. Scratches are not on a sealing surface. Scratches should not impact form, fit, or function of the window when installed on the cavity string. Window is for cryomodule testing only and will not be delivered to customer.
357 | Nov-04-2022 | Nov-11-2022 | L2HE-INSP-BLBU | 9 | WELDMENT BELLOWS - US END | 002 | Dings OTHER | Inspection of the bellow dings in two rows, 180° apart. One row (1st set of dents in photos) has dents with sharper edges. Damaged resulted from vendor fixturing. |  |  | Use As Is | Tentatively released for production. The L2HE project is conducting further investigations.
358 | Nov-04-2022 | Nov-11-2022 | L2HE-INSP-BLBU | 10 | WELDMENT BELLOWS - US END | 003 | Dings | Inspection of the bellow dings in two rows, 180° apart. One row (1st set of dents in photos) has dents with sharper edges. Damaged resulted from vendor fixturing. |  |  | Use As Is | Tentatively released for production. The L2HE project is conducting further investigations.
359 | Nov-04-2022 | Nov-11-2022 | L2HE-INSP-BLBU | 11 | WELDMENT BELLOWS - US END | 008 | Dings | Inspection of the bellow dings in two rows, 180° apart. One row (1st set of dents in photos) has dents with sharper edges. Damaged resulted from vendor fixturing. |  |  | Use As Is | Tentatively released for production. The L2HE project is conducting further investigations.
360 | Nov-04-2022 | Nov-11-2022 | L2HE-INSP-BLBU | 12 | WELDMENT BELLOWS - US END | 009 | Dings | Inspection of the bellow dings in two rows, 180° apart. One row (1st set of dents in photos) has dents with sharper edges. Damaged resulted from vendor fixturing. |  |  | Use As Is | Tentatively released for production. The L2HE project is conducting further investigations. | MICHELLE O - Need more information - Datamine shows nothing else was done, cleaned and used in string. Was anything discovered in the investigations?
361 | Nov-04-2022 | Nov-11-2022 | L2HE-INSP-BLBU | 13 | WELDMENT BELLOWS - US END | 011 | Dings | Inspection of the bellow dings in two rows, 180° apart. One row (1st set of dents in photos) has dents with sharper edges. Damaged resulted from vendor fixturing. |  |  | Use As Is | Tentatively released for production. The L2HE project is conducting further investigations.
362 | Nov-04-2022 | May-19-2023 | L2HE-INSP-BLBU | 14 | WELDMENT BELLOWS - US END | 012 | Dings | Inspection of the bellow dings in two rows, 180° apart. One row (1st set of dents in photos) has dents with sharper edges. Damaged resulted from vendor fixturing. |  |  | Reject | Vendor won't accept back, reject for pin holes and one of the worst example of dings
363 | Nov-07-2022 | Nov-11-2022 | L2HE-INSP-BLBU | 15 | WELDMENT BELLOWS - US END | 010 | Dings | Inspection of the bellow dings in two rows, 180° apart. One row (1st set of dents in photos) has dents with sharper edges. Damaged resulted from vendor fixturing. |  |  | Use As Is | Tentatively released for production. The L2HE project is conducting further investigations.
364 | Nov-21-2022 | Jul-11-2023 | L2HE-VTA-CAV-VTRF | 109 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R105 | N/A | FE onset at 29 MV/m, 0.47 mR/hr. | Hold/Modify | re-rinse & re-test (small HPR -- bottom flange drop) | Reject | Cavity rejected due to bellows damage
365 | Nov-23-2022 | Nov-30-2022 | L2HE-INSP-BLBP | 34 | WELDMENT BELLOWS PRCM | 042 | Scratches | Non-rotating flange has scratches in seal area "1". See attached photos. | Hold/Modify | Send to machine shop for polishing | Conforming | Seal surfaces now conforming
366 | Nov-30-2022 | Dec-05-2022 | L2HE-CLNRM-FPCC-INSP | 36 | FPC COLD PART | CPI-FPCC-HE-020 | N/A | small tear through both inner and outer bags at valve location. |  |  | Use As Is | Particle counts to be tested later
367 | Dec-02-2022 | Dec-02-2022 | L2HE-INSP-CAV | 78 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R110 | OTHER | No serial # on the Helium Tank. |  |  | Use As Is | Will contact vendor about this.
368 | Dec-07-2022 | Jan-19-2024 | L2HE-VTA-CAV-VTRF | 113 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R113 | N/A | Q0 degraded to below spec at 20.8 MV/m after quenching events during initial power rise. | Hold/Modify | After 2022-12-20 test re-rinse & re-test (small HPR -- bottom flange drop) CTB 2023-01-10 Cavity has to meet performance spec w/o MP events in final ramp up. Please re-test -- CTB 2022-12-06 | Reject | CTB 2023-12-19 send to Fermi/Argonne for HF rinse
369 | Dec-13-2022 | Dec-15-2022 | L2HE-INSP-CAV | 80 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R124 | Dings | (1) indentation on the bellows near HOM tuning arm. See attached photos. |  |  | Use As Is | Thanks for catching this! No material removal, so use as is.
370 | Dec-14-2022 | Dec-15-2022 | L2HE-INSP-CAV | 81 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R120 | Dings | Concave convolution has a ding/kink. This is located in relation to the clamp pin rotation. |  |  | Use As Is | After consulting SLAC -- use as is
371 | Dec-15-2022 | Dec-15-2022 | L2HE-INSP-CAV | 81 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R120 | OTHER | The SN for the FlangeShortSN is not correct. Missing a “1”. See attached pic. Should be: F10047913-A-006 Is: F1004793-A-006 |  |  | Use As Is | Part is functional
372 | Dec-15-2022 | Jan-17-2023 | L2HE-CMA-CM-ASSY1 | 3 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-24 | Vacuum | Leak was discovered in Helium vessel CAVR082 in cavity string J1.3-24 Position 2 during leak check of the 2 phase process piping. leak size determined to be 7e-7 torr l/s | Hold/Modify | Meeting held with Partner Labs. Repair of weld will be performed. Procedure steps attached with comment from Lead engineer, CMA Lead, CAM, and others. FNAL did a similar repair, their documents attached for reference as well. | Conforming | The leak has been fixed. J24 is good to proceed to subsequent assembly work. There will be a pressure test of Lines A, B, G & H at the final assembly level. The repaired weld will eventually experience that pressure test
373 | Dec-16-2022 | Mar-10-2023 | L2HE-CMA-FPCW-INSP | 13 | FPC WARM PART | CPI-FPCW-HE-015 | LeakTestFailure | Leak check showed a 4.8e-9 Torr L/sec located at the sapphire window. | Hold/Modify | replace sapphire window with a spare from coupler in nitrogen cabinet | Use As Is | Window replaced, leak check successful
374 | Dec-16-2022 | Mar-10-2023 | L2HE-CMA-FPCW-INSP | 14 | FPC WARM PART | CPI-FPCW-HE-016 | LeakTestFailure | Leak check showed a 5.2e-10 Torr L/sec leak located at the sapphire window. | Hold/Modify | Switch out windows with ones from couplers in N2 cabinet | Use As Is | window was replaced, leak test successful
375 | Dec-19-2022 | Dec-22-2022 | L2HE-INSP-FWM | 11 | LCLS-II CM FARADAY PORT WINDOW | 011 | Dimensional | Please see attached. |  |  | Reject | Window is not acceptable for use due to being under thickness. This part is under vendor warranty and the vendor will make a new window. The window SN011 does not need to be returned to the vendor per the vendor's response to Tom Huratiak. Analysis of window thickness was performed by Phillip Dobrenz. Email with results of analysis attached to this NCR.
376 | Dec-21-2022 | Oct-07-2024 | L2HE-VTA-CAV-VTRF | 110 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R106 | N/A | FE onset 24 MV/m, unable to process away. After quench, Q0 degraded below spec. | Hold/Modify | re-rinse & re-test (small HPR -- bottom flange drop) CTB 2023-01-10 | Return To Vendor | CTB 2023-05-02 Put cavity on hold for potential repair at vendor 2024-06-03 Went to big HPR for re-test 2024-07-01 RF issue -- cold leak made test impossible 2024-09-24 sent to FNAL after big HPR for test in VTS
377 | Jan-03-2023 | Aug-29-2023 | L2HE-CMA-VV-INSP | 6 | VESSEL LCLS-II AFTER MACHING WELDMENT | RFCHE-VV-018 | Residues Scratches | Visual inspection completed on 08-Dec-2022. All photos are on Server: M:\cmm\Pictures\Inspection Pictures\L2HE\Vacuum vessel\VV018 Scratches and paint over-spray was observed. See attached photos. | Hold/Modify | VV018 is installed onto J1.3-23 cryomodule. Use visual inspection findings as guide, perform necessary polishing/cleaning | Conforming | J1.3-23 CM assembly is complete. VV isuues have been take care of
378 | Jan-05-2023 | Jan-05-2023 | L2HE-INSP-BPM | 5 | BPM HOUSING W/ CU PLATING | 110 | Plating Stains | A few spots were not able to be removed (photo -0599 and -0600). Rough boundary where the copper plating ends was noticed as it snagged on the wipes used, no flaking of Cu was noted(photo -0596). Ultra Fine Scotch-Brite used on seal surfaces (gray-white-white method) to remove minor stains and attempt to smooth rough boundary of copper plating unsuccessfully. BPM blew down over particle counter. Suggestion to inform chemistry and clean room staff to focus on the plating boundary area when performing final blow down before assembly to ensure no fibers snagged from clean room wipes present (if using as-is). |  |  | Use As Is | Clean room will be notified of rough copper boundary, small oxidation spots are acceptable
379 | Jan-17-2023 | Feb-21-2023 | L2HE-CMA-CM-ASSYF | 1 | LCLS-II UPPER COLD MASS ASSEMBLY | RFCHE-UC-019 | Vacuum | During the leak check of the circuit "jumpers" during Cryomodule final assembly activities a leak was discovered in Line F, the aluminum weld behind the bonded joint. Leak was E-4 and found by pressurizing the circuit to 15 psig and applying snoop. See attached video and photos | Hold/Modify | We have decided to repair the leak, perform cold shock, pressure test and leak check after the repair. Technical Representative (Gary Cheng) of the upper cold mass assembly is communicating with FermiLab about interacting with the vendor to understand the root cause of the leak | Conforming | After the successful weld repair and verification test, the Line F is found leak tight
380 | Jan-19-2023 | Feb-15-2023 | L2HE-CMA-UCM-INSP | 7 | LCLS-II UPPER COLD MASS ASSEMBLY | RFCHE-UC-016 | Dings Scratches | Please see attached for damage on knife edge. |  |  | Use As Is | J1.3-24 WS 3 leak check is performed with attention to the marked damages. Fortunately the damages did not cause leak during the test
381 | Jan-19-2023 | Feb-15-2023 | L2HE-CMA-UCM-INSP | 7 | LCLS-II UPPER COLD MASS ASSEMBLY | RFCHE-UC-016 | OTHER | Sharp edges present on hangars. Pictures attached. |  |  | Use As Is | J1.3-24 cavity string is successfully united with UCM 016. The sharp edge was paid attention to during the assembly
382 | Jan-30-2023 | Feb-01-2023 | L2HE-CHEM-CAV-USC | 116 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R086 | Dings | Damage found on bellows, dents. See photos |  |  | Return To Vendor | CTB 2023-01-31 decision: We are going to pull this cavity and prepare it for shipment to the vendor for repair. Thanks for catching this!
383 | Jan-30-2023 | Feb-01-2023 | L2HE-CHEM-CAV-USC | 118 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R105 | Dings | Ding in bellow, see photos |  |  | Return To Vendor | CTB 2023-01-31 decision: We are going to pull this cavity and prepare it for shipment to the vendor for repair. Thanks for catching this!
384 | Jan-30-2023 | Feb-01-2023 | L2HE-CHEM-CAV-USC | 119 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R099 | OTHER | Bellows are skewed out of plane, see photos. No dings or dents found on bellows. |  |  | Use As Is | Please use as it. This type of "unconventional" deformation of the bellows is known to the project. However, it does neither deviate from any given specification nor hinder the functionally. The project will try to provide a better guideline for noting this at the incoming inspection in the future.
385 | Feb-04-2023 | Feb-06-2023 | L2HE-INSP-CAV | 84 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R123 | Dimensional | (1) Dimension OOT: Perpendicularity of main coupler port to helium vessel lugs. Should be: .5° max deviation Is: .6, OOT by .1° |  |  | Use As Is | Acceptable as is for FPC installation.
386 | Feb-09-2023 | Jul-11-2023 | L2HE-VTA-CAV-VTRF | 128 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R124 | RFTestFailure | Cavity failed to lock frequency with SEL on, cables check good with TDR and VNA, suspect superfluid leak. |  |  | Conforming | Test on 2023-03-30
387 | Feb-11-2023 | Feb-27-2023 | L2HE-CWI-UCM | 7 | LCLS-II UPPER COLD MASS ASSEMBLY | RFCHE-UC-018 | InspectionFailure | The pipe weld joint on the aluminum side and the pipe weld on the stainless-steel side of the transition joints have evidence of grinding operation conducted by the vendor’s welder. The aluminum weld joint has one spot located at the 12:00 position. The SS weld joint was totally ground down very close to an unacceptable level. | Hold/Re-measure | Having discussed with project management, we decide to cold shock (3x), pressure test (to 320 psig) and leak check the Line F to confirm whether it remains leak tight after being stressed | Use As Is | The Line F is found to be leak tight after multiple cold shocks and pressure test. It is proved to be leak tight hence the J23 CM assembly work may proceed
388 | Feb-21-2023 | Feb-22-2023 | L2HE-INSP-CAV | 92 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R132 | Scratches | Light scratch on Burst Disc. See attached photos. |  |  | Use As Is | RF incoming does not indicate a leak
389 | Feb-23-2023 | Feb-23-2023 | L2HE-INSP-CAV | 93 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R151 | Scratches | Light scratch on Burst Disc. |  |  | Use As Is
390 | Feb-23-2023 | Mar-27-2023 | L2HE-INSP-CAV | 93 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R151 | OTHER | Per John V.; "When the cavity arrives in CMM, please double-check if the Field Probe SN is GMM-B4787-0266 Please contact the PI so we can initiate a replacement in the clean room". | Hold/Modify | The FP has to be replaced using the N2 purge system in the clean room | Conforming | FP has been replaced
391 | Feb-23-2023 | May-17-2023 | L2HE-VTA-CAV-VTRF | 132 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R123 | VTARFPerformance | Final run, FE onset 8.5 MV/m. Emax did not reach 23 MV/m. | Hold/Modify | re-rinse & re-test (big HPR -- full disassembly) | Conforming | No FE and all requirements met after HPR
392 | Feb-24-2023 | Jul-13-2023 | L2HE-CLNRM-CST-ASSY | 5 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-25 | Scratches | Damage has been found on three cavity bellows, in positions 1, 5 and 6, s/n: HE-CAV-R094, 078, 110. |  |  | Reject | Cavity string rejected for use; string has been rebuilt using different cavities. Damaged cavities will be repaired.
393 | Feb-24-2023 | Jul-18-2023 | L2HE-CHEM-CAV-USC | 128 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R094 | Dings | Damage found on cavity bellows during J25 string assembly. | Hold/Re-measure | As of CTB 2023-02-21 Cavity will be removed from string & clean room for further investigation | Return To Vendor | Will be returned to vendor for repair
394 | Feb-24-2023 | Jul-18-2023 | L2HE-CHEM-CAV-USC | 112 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R078 | Dings | Damage found on cavity bellows during J25 string assembly. | Hold/Re-measure | As of CTB 2023-02-21 Cavity will be removed from string & clean room for further investigation | Return To Vendor | Will be returned to vendor for repair
395 | Feb-24-2023 | Jul-18-2023 | L2HE-CHEM-CAV-USC | 113 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R110 | Dings | Damage found on cavity bellows during J25 string assembly. | Hold/Re-measure | As of CTB 2023-02-21 Cavity will be removed from string & clean room for further investigation | Return To Vendor | Will be returned to vendor for repair
396 | Feb-24-2023 | Jul-18-2023 | L2HE-CHEM-CAV-USC | 73 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R108 | Dings | Damage found on cavity bellows while on cleanroom cavity rack. |  |  | Return To Vendor | Will be returned to vendor for rework on bellows and performance
397 | Feb-27-2023 | Mar-02-2023 | L2HE-CMA-CM-ASSY2 | 3 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-24 | OTHER | This NCR is to document the fact that a helium vessel bellows inspection was done at WS3 on J1.3-24. This was in response too bellows damage seen on other helium vessels in string J1.3-25 and other. | Hold/Re-measure | The helium vessel magnetic shield were partially disassembled to provide access to the lower quadrant of the bellows (where other damage was found) for inspection. See attached photos. | Conforming | In WS2 assembly traveler L2HE-CMA-CM-ASSY1, there is a step to visually inspect the cavity string when rolled out of the cleanroom. We think that step would catch bellows defects before further assembly work is done. The fact that we did not find apparent helium vessel bellows defects by performing finger touch inspection on J24 gives us confidence in believing that J22 & J23 have a good chance in not having helium vessel bellows defect.
398 | Feb-27-2023 | Feb-27-2023 | L2HE-CMA-FPCW-INSP | 2 | FPC WARM PART | CPI-FPCW-HE-002 | LeakTestFailure | Leak was found at the sapphire window. The size of the leak is 8.33E-10 Torr L/sec | Hold/Modify | Replace sapphire window with a spare | Use As Is | Leak has been corrected with replacement window
399 | Mar-10-2023 | Apr-17-2025 | L2HE-INSP-BPM | 6 | BPM HOUSING W/ CU PLATING | 114 | Stains | Large stains exceed 1/8 inch on copper. Pictures attached. | Hold/Modify | Send to chemistry for cleaning | Conforming | NCR reopened to change from use as is to conforming. Copper plating is clean.
400 | Mar-10-2023 | Apr-17-2025 | L2HE-INSP-BPM | 6 | BPM HOUSING W/ CU PLATING | 114 | Delamination | Questionable area on copper plating could be possible bare spot or could be staining. Please see pictures attached. Area is over 1/8 inch in length. | Hold/Modify | Send to chemistry for cleaning | Conforming | NCR reopened to change from use as is to conforming. Copper plating is clean.
401 | Mar-10-2023 | Mar-31-2023 | L2HE-INSP-BPM | 7 | BPM HOUSING W/ CU PLATING | 120 | Stains | Stains/dark spots exceed 1/8 inch. Picture attached. | Hold/Modify | Send to machine shop for seal surface polishing, then send to chemistry for cleaning | Use As Is | Seal surfaces polished and copper plating cleaned
402 | Mar-10-2023 | Mar-31-2023 | L2HE-INSP-BPM | 7 | BPM HOUSING W/ CU PLATING | 120 | Scratches | Scratches in seal area on multiple flanges. Please see attached. | Hold/Modify | Send to machine shop for seal surface polishing, then send to chemistry for cleaning | Use As Is | Seal surfaces polished
403 | Mar-13-2023 | Mar-14-2023 | L2HE-INSP-CAV | 98 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R134 | Dimensional | DIMTOL6 is OOT. See attached report, page 1 of 3. |  |  | Use As Is | The deviation is acceptable.
404 | Mar-21-2023 | May-17-2023 | L2HE-VTA-CAV-VTRF | 134 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R143 | RFTestFailure VTARFPerformance | Cannot find the cavity - lock frequency. Probably something is wrong with RF cable inside Dewar. During cable calibration had: Popen=28.60mW, Pfp=20.76 mW and Pfpc=21.66 mW. Suspect at least one connection with issues near to the cavity. | Hold/Modify | Cavity had a (cold) leak. Full disassembly and re-test | Conforming | cold leak has been resolved
405 | Mar-21-2023 | Feb-20-2024 | L2HE-VTA-CAV-VTRF | 135 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R132 | RFTestFailure VTARFPerformance | High radiation at cavity gradient bellow 16 MV/m. Radiation induced quench events. CW RF processing for about one hour ineffective. | Hold/Modify | re-rinse & re-test (big HPR -- full disassembly) | Reject | Cavity send to Fermilab for HF rinse CTB 2024-01-19
406 | Mar-21-2023 | Mar-21-2023 | L2HE-INSP-QUAD | 6 | LCLS-II MAGNET PACKAGE ASSEMBLY | 207 | InspectionFailure | Q-factor is slightly out of tolerance .263 @100Hz. Spec is 0.3 - 0.4. | Hold/Re-measure | Retest and record the Q factor at 100 Hz for the quadrupole winding. | Use As Is | Our electrical inspection results indicate that the measured value of the Q factor at 100 Hz (0.263 (from test #1) and 0.271 (from test #2)) for the SPQA 207 magnet is marginally below the lower bound (0.3) of the anticipated value. Note that these limits were set according to the data, rounded off to one decimal place, obtained from measurements at the manufacturer site and Fermi Lab. Since the room temperature resistance of the coils and magnet assembly and the Q factor at 1 kHz fall within the suggested range, the observed discrepancy in the Q value at 100 Hz can be attributed to a statistical error in measurements. Nonetheless, it is expected to have no measurable impact on magnet performance at the actual operating conditions.
407 | Mar-21-2023 | Mar-21-2023 | L2HE-INSP-QUAD | 7 | LCLS-II MAGNET PACKAGE ASSEMBLY | 208 | InspectionFailure | Step 3.2.4, Q-factor @ 100Hz of Q+/Q- is slightly out of tolerance at 0.263. Spec is 0.3 - 0.4. | Hold/Re-measure | Retest and record the Q factor at 100 Hz for the quadrupole winding. | Use As Is | Our electrical inspection results indicate that the measured value of the Q factor at 100 Hz (0.265 (from test #1) and 0.269 (from test #2)) for the SPQA 208 magnet is marginally below the lower bound (0.3) of the anticipated value. Note that these limits were set according to the data, rounded off to one decimal place, obtained from measurements at the manufacturer site and Fermi Lab. Since the room temperature resistance of the coils and magnet assembly and the Q factor at 1 kHz fall within the suggested range, the observed discrepancy in the Q value at 100 Hz can be attributed to a statistical error in measurements. Nonetheless, it is expected to have no measurable impact on magnet performance at the actual operating conditions.
408 | Mar-27-2023 | Aug-02-2023 | L2HE-VTA-CAV-VTRF | 140 |  | HE-R096 | RFTestFailure VTARFPerformance Vacuum | RF power test in D7 aborted: cold leak. | Hold/Modify | Cavity had a cold leak We will try retorquing and re-testing the cavity | Reject | Send for HF rinse to Fermilab
409 | Mar-27-2023 | Mar-28-2023 | L2HE-VTA-CAV-VTRF | 139 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R081 | RFTestFailure | Qext1 was 7.8e9 during VTRF |  |  | Use As Is | CTB decision 2023-03-28
410 | Mar-27-2023 | Aug-02-2023 | L2HE-VTA-CAV-VTRF | 137 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R098 | RFTestFailure | FE onset at 7 MV/m during VTRF. | Hold/Modify | re-rinse & re-test (big HPR -- full disassembly) CTB 2023-05-02 re-rinse & re-test (2nd big HPR -- full disassembly) | Conforming | Test 2023-08-01 cavity qualified
411 | Mar-27-2023 | Aug-29-2023 | L2HE-VTA-CAV-VTRF | 138 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R104 | RFTestFailure | Could not lock cavity frequency during VTRF. Cables looked fine on TDR. | Hold/Modify | Cavity had a cold leak We will try retorquing and re-testing the cavity | Conforming | Cavity cleaned up after 3rd HPR for test on 2023-08-23
412 | Mar-29-2023 | Jul-13-2023 | L2HE-INV-CAV | 92 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R090 | OTHER | External Shock-watch has been triggered. See attachment DSCN0892 in traveler. |  |  | Conforming | Cavity passed RFIN inspection.
413 | Mar-31-2023 | Mar-31-2023 | L2HE-TUNE-CAV-RFIN | 75 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R090 | RFTestFailure | While Pi mode was in spec on RFIN, field flatness change is greater than 10%. |  |  | Use As Is | Additional deviation from RF scan at vendor most likely caused by handling at FNAL. Overall, spectrum is acceptable.
414 | Mar-31-2023 | Mar-31-2023 | L2HE-TUNE-CAV-RFIN | 74 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R115 | RFTestFailure | While Pi mode is within spec, field flatness change is greater than 10%. |  |  | Use As Is | Additional deviation from RF scan at vendor most likely caused by handling at FNAL. Overall, spectrum is acceptable.
415 | Mar-31-2023 | Apr-04-2023 | L2HE-VTA-CAV-VTRF | 143 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R151 | VTARFPerformance | Qext2 is out of spec slightly high: 2.9500e+12 |  |  | Use As Is | Acceptable as of CTB 2023-04-04 decision
416 | Apr-04-2023 | Apr-04-2023 | L2HE-INSP-CAV | 100 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R115 | Dings Scratches | Hom housing concave, dings present (short fill line end of cavity). Pictures attached. |  |  | Conforming | HOMs have already been tuned at FNAL, hence the marks & deformation
417 | Apr-04-2023 | Aug-11-2023 | L2HE-VTA-CAV-VTRF | 134 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R143 | VTARFPerformance | FE onset at 9 MV/m, FE loaded by 13 MV/m. | Hold/Modify | re-rinse & re-test (big HPR -- full disassembly) | Conforming | Cavity qualified on 2023-08-08
418 | Apr-05-2023 | Apr-05-2023 | L2HE-CHEM-CAV-USC | 146 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R090 | Delamination | Metal was found stuck to the backside of the bellows-end beam-line flange during cleaning. The metal looks like copper in the photos because of the lighting, it was not copper-colored. | Hold/Modify |  | Conforming
419 | Apr-06-2023 | Apr-07-2023 | L2HE-INSP-CAV | 101 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R137 | Dings | Blemish/impression on bellows. Pictures attached. |  |  | Use As Is | Acceptable defect, see reference. Will double check with Fermi that this was already present before their handling. Generally speaking, since there is no material removal, this is acceptable either way.
420 | Apr-06-2023 | Aug-02-2023 | L2HE-VTA-CAV-VTRF | 145 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R134 | VTARFPerformance | Cavity Q0 below spec at 20.8 MV/m, 2.37e10 | Hold/Re-measure | We'll try to thermally cycle the cavity and perform a retest | Use As Is | Will be used as marginal cavity in last string
421 | Apr-11-2023 | Aug-02-2023 | L2HE-VTA-CAV-VTRF | 146 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R134 | VTARFPerformance | Qo is below specification of 2.5e10. | Hold/Re-measure | Please perform a Q vs T measurement (In the next couple of days -- Justin is scheduling for that) | Use As Is | Will be used as marginal cavity in last string
422 | Apr-11-2023 | Apr-11-2023 | L2HE-INSP-BLBP | 40 | WELDMENT BELLOWS PRCM | 073 | Plating | Pitting present on copper. Pictures attached. |  |  | Return To Vendor | Return to vendor for replating
423 | Apr-12-2023 | Apr-12-2023 | L2HE-INSP-BLBP | 38 | WELDMENT BELLOWS PRCM | 070 | Plating | Lots of pitting throughout bellows plating. Pictures attached. |  |  | Return To Vendor | Return to vendor for replating
424 | Apr-12-2023 | Apr-12-2023 | L2HE-INSP-BLBP | 41 | WELDMENT BELLOWS PRCM | 076 | Dimensional | There are (3) dimensions that are OOT. See attached. |  |  | Use As Is | OOT is small enough not to effect usability of bellows
425 | Apr-13-2023 | Apr-19-2023 | L2HE-INSP-BLBP | 41 | WELDMENT BELLOWS PRCM | 076 | Plating | Shallow pitting noted on plating surfaces (ID diameters and convolutions). "Shallow pitting" as stated in Step 2. NCR generated over pitting concerns. |  |  | Use As Is | Use as is per Michelle and Katherine - pits are barely visible to the naked eye
426 | Apr-16-2023 | Apr-17-2023 | L2HE-INSP-BLBP | 39 | WELDMENT BELLOWS PRCM | 072 | Scratches | Scratches noted in NR flange seal groove. Scratches are inboard of the seal contact area. |  |  | Return To Vendor | Pitting on NCR 428 is cause for replating
427 | Apr-16-2023 | Apr-17-2023 | L2HE-INSP-BLBP | 39 | WELDMENT BELLOWS PRCM | 072 | Scratches | 1st exterior convolution from the NR flange has a scratch/cut area approximately 1/8" long, .01" wide, & .004" deep. |  |  | Return To Vendor | Pitting on NCR 428 is cause for replating
428 | Apr-16-2023 | Apr-17-2023 | L2HE-INSP-BLBP | 39 | WELDMENT BELLOWS PRCM | 072 | Plating | Concentrated areas of shallow pitting and/or blistering noted on ID plating and convolutions. |  |  | Return To Vendor | Extensive pitting is cause for replating
429 | Apr-20-2023 | Jun-06-2023 | L2HE-CLNRM-FPCC-INSP | 42 | FPC COLD PART | CPI-FPCC-HE-129 | Dings Oxidation Scratches | CPI-FPCC-HE-129 Inspection of copper plating revealed oxidations and scratches around the RF hole. |  |  | Use As Is | Oxidation spots are small and the scratches are not deep enough to affect the RF contact
429 | Apr-20-2023 | Jun-06-2023 | L2HE-CLNRM-FPCC-INSP | 42 | FPC COLD PART | CPI-FPCC-HE-129 | Dings Oxidation Scratches | CPI-FPCC-HE-129 Inspection of copper plating revealed oxidations and scratches around the RF hole. |  |  | Use As Is | Oxidation spots are small and the scratches are not deep enough to affect the RF contact
430 | Apr-20-2023 | Apr-20-2023 | L2HE-CLNRM-FPCC-INSP | 43 | FPC COLD PART | CPI-FPCC-HE-140 | Oxidation | Small spotting and dark discoloration on the copper plating. |  |  | Use As Is | Dark spots are too small to affect the function of the FPC
431 | Apr-20-2023 | Apr-20-2023 | L2HE-CLNRM-FPCC-INSP | 43 | FPC COLD PART | CPI-FPCC-HE-141 | Stains | Dark spotting on copper near RF Hole. |  |  | Return To Vendor
432 | Apr-25-2023 | Apr-25-2023 | L2HE-CHEM-CAV-USC | 158 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R095 | Scratches | Scratches in HV bellows were noticed after cavity was sent out of the clean room for ultrasonic cleaning. |  |  | Conforming | Cavity bellows are conforming. There are two minor surface finish blemishes on the top convolutions in the photo, the third convolution is unmarked as seen in the photo attached here. The mark seen on the original picture was a piece of particulate.
433 | May-01-2023 | May-04-2023 | L2HE-INSP-BLBP | 44 | WELDMENT BELLOWS PRCM | 097 | Residues | Initial visual inspection showed that the copper plating on ID surfaces has a hazy/streaky appearance. Possible film or residue from rinse. | Hold/Modify | Send to Chemistry for cleaning and return to CMM room for inspection | Use As Is | Plating is now clean | MICHELLE O - Should be conforming?
434 | May-19-2023 | May-23-2023 | L2HE-INSP-BLBP | 54 | WELDMENT BELLOWS PRCM | 100 | OTHER Oxidation Stains | Bulges/Gaps in weld may create leak path. | Hold/Re-measure | Hold for Jenord to inspect weld | Return To Vendor | Leak check failed due to bad weld
435 | May-22-2023 | Jun-27-2023 | L2HE-INSP-BLBP | 55 | WELDMENT BELLOWS PRCM | 103 | OTHER Oxidation Residues | Dust on plated convolutions | Hold/Modify | Send to chemistry for cleaning and removal of debris | Use As Is | Cleaning successful
436 | May-23-2023 | Jun-09-2023 | L2HE-INSP-BLBP | 55 | WELDMENT BELLOWS PRCM | 103 | Plating | Irregular plating at weld edge may indicate leak paths are present | Hold/Re-measure | Please send to Tom Prosper for leak check | Use As Is | Second leak test was successful
437 | May-24-2023 | May-24-2023 | L2HE-INSP-BLBP | 53 | WELDMENT BELLOWS PRCM | 101 | LeakTestFailure Plating | Non-smooth weld edge imperfections can potentially compromise hermetic requirements or cause degradation of the plated seal edge. | Hold/Re-measure | Leak test | Return To Vendor | Leak check failed due to faulty weld
438 | May-25-2023 | Apr-17-2025 | L2HE-INSP-BLBP | 56 | WELDMENT BELLOWS PRCM | 085 | Delamination Plating | 1 irregular plating ~1/10" into edge from weld area; reference photo 'Plating 2 Pt. 1' and NCR | Hold/Modify | Send to chemistry for cleaning | Conforming | NCR reopened to change from use as is to conforming. Copper plating is clean.
440 | May-26-2023 | Apr-17-2025 | L2HE-INSP-BLBP | 58 | WELDMENT BELLOWS PRCM | 079 | Stains | Lt. brown oil-like stains on flanges incl. seal surface will not clean with IPA or Acetone. | Hold/Modify | Send to chemistry for cleaning | Conforming | NCR reopened to change from use as is to conforming. Copper plating is clean, stains removed. Scratch did not affect performance.
441 | May-30-2023 | Jun-08-2023 | L2HE-INSP-BLBP | 57 | WELDMENT BELLOWS PRCM | 081 | Dings Plating | Indentation bulge defect in interior plated edge near weld measures ~1/8" W x 1/16" DP. | Hold/Re-measure | Leak test | Use As Is | Successful leak check
442 | May-30-2023 | Jun-08-2023 | L2HE-INSP-BLBP | 57 | WELDMENT BELLOWS PRCM | 081 | Stains | Lt. brown oil-like spots on rotating flange face & seal surface. | Hold/Modify | Send to chemistry for cleaning | Use As Is | Cleaning successful
443 | May-30-2023 | May-31-2023 | L2HE-CMA-UCM-DMAG | 4 | LCLS-II UPPER COLD MASS ASSEMBLY | RFCHE-UC-008 | Mechanical | One pad could not be fully demagged. (Cavity 7 area) Several attempts were made. Unit measured ~40mG | Hold/Re-measure | Please repeat the measurement to confirm the residual magnetic field magnitude | Use As Is | Though the residual field on this pad exceeded the spec of 10 mG, the effect on adjacent cavity's Q0 is minimal since the measurement is done on contact. 24 mG is expected to decay quick vs distance
444 | May-31-2023 | Jun-02-2023 | L2HE-INSP-CAV | 105 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R164 | Dimensional | Dimensional non-conforming. PDF attached. |  |  | Use As Is | Can be compensated during welding of 2pp
445 | May-31-2023 | May-31-2023 | L2HE-INSP-BLBP | 48 | WELDMENT BELLOWS PRCM | 120 | Dimensional | (1) Dimension is OOT: Bellows ID; 78.0 mm +/- 0.3 mm = 79.21 |  |  | Use As Is | OOT is small and won't effect bellows functionality
446 | Jun-01-2023 | Jul-26-2023 | L2HE-INSP-BLBP | 50 | WELDMENT BELLOWS PRCM | 127 | Stains | Oil-like stains on flange face/sealing surface (Rotating side) | Hold/Modify | Send to chemistry for cleaning | Use As Is | Cleaning was successful
447 | Jun-02-2023 | Jun-07-2023 | L2HE-INSP-BLBP | 52 | WELDMENT BELLOWS PRCM | 148 | Oxidation Residues Stains | Multiple small oxidation spots measure 1/64" to 3/16"; White residue on inner convolutions. | Hold/Modify | Send to chemistry for cleaning | Use As Is | Stains removed by chemistry
448 | Jun-07-2023 | Jun-07-2023 | L2HE-INSP-BLBP | 49 | WELDMENT BELLOWS PRCM | 082 | Dings | Small dent on interior plated surface measures ~1/8". |  |  | Use As Is | Shallow dent on flat will not effect functionality of bellows
449 | Jun-07-2023 | Jun-07-2023 | L2HE-INSP-BLBP | 47 | WELDMENT BELLOWS PRCM | 113 | Dimensional | (3) dimensions are OOT. 1) Bellows ID 78.0 mm +/- 0.3 mm = 79.20 2) Bellows OD 98.0 +/- 0.3 mm = 98.58 3) Sealing surface depth 3.1 +/- 0.1 mm = 2.96 |  |  | Use As Is | OOT values are small and won't change effectiveness of bellows
450 | Jun-07-2023 | Jun-08-2023 | L2HE-INSP-BLBP | 51 | WELDMENT BELLOWS PRCM | 084 | Stains | Large, dark stains on rotating flange sealing surface measure 1/8" and will not come clean with alcohol. | Hold/Modify | Send to chemistry for cleaning | Use As Is | Cleaning was successful
451 | Jun-07-2023 | Oct-24-2023 | L2HE-LERF-CM-PREP | 2 |  | J1.3-22 | Vacuum | On 2 June 2023 CM J22, in LERF, found to have a beamline vacuum leak in the e-6 range and localized to the area of cavity 7 on the tuner end. | Hold/Re-measure | The L2HE project permits J22 beam line leak root-cause analysis that may start with leak check at WS5 | Reject | The L2HE project decides to let JLAB fully dissemble J1.3-22 and rebuild it. See CAPA 051 (https://jlabdoc.jlab.org/docushare/dsweb/Services/Document-276859) for more details
452 | Jun-08-2023 | Jun-08-2023 | L2HE-INSP-BLBP | 50 | WELDMENT BELLOWS PRCM | 127 | Dimensional | CMM inspection performed. (3) dimensions are OOT. 1) Bellows ID 78.0 mm +/- 0.3 mm = 79.23 2) Bellows OD 98.0 +/- 0.3 mm = 98.68 3) Sealing surface depth 3.1 +/- 0.1 mm = 2.96 |  |  | Use As Is | OOT values are small and will not effect bellows performance
453 | Jun-09-2023 | Jun-09-2023 | L2HE-INSP-BLBP | 52 | WELDMENT BELLOWS PRCM | 148 | Dimensional | CMM inspection performed. (3) dimensions are OOT SN148: 1) Bellows ID 78.0 mm +/- 0.3 mm = 79.21 2) Bellows OD 98.0 +/- 0.3 mm = 98.50 3) Sealing surface depth 3.1 +/- 0.1 mm = 2.96 |  |  | Use As Is | OOT values are small and won't effect bellows performance
454 | Jun-12-2023 | Jun-12-2023 | L2HE-INSP-BLBP | 46 | WELDMENT BELLOWS PRCM | 106 | Dimensional | (3) dimensions are OOT SN106: 1) Bellows ID 78.0 mm +/- 0.3 mm = 79.21 2) Bellows OD 98.0 +/- 0.3 mm = 98.55 3) Rotating Sealing depth 3.1 +/- 0.1 mm = 2.69 |  |  | Use As Is | OOT values are small and won't effect bellows performance
455 | Jun-12-2023 | Aug-02-2023 | L2HE-VTA-CAV-VTRF | 163 |  | HE-R163 | RFTestFailure VTARFPerformance | Radiation onset at 22.8 MV/m | Hold/Modify | Proceed with small HPR | Conforming | Test on 2023-07-26 qualified
456 | Jun-21-2023 | Jun-27-2023 | L2HE-INSP-BLBS | 10 | WELDMENT BELLOWS SHORT | 019 | OTHER Plating | Possible bare spot on inner copper convolution (reference photo 'Plating 1 Mid 2.JPG' below); Category 9-12.1 Plating discoloration near weld area. | Hold/Modify | Send to chemistry to see if they can clean the spot on convolutions | Use As Is | Use as a spare, scratch is very small
457 | Jun-27-2023 | Sep-28-2023 | L2HE-CMA-VV-INSP | 7 | VESSEL LCLS-II AFTER MACHING WELDMENT | RFCHE-VV-016 | Dimensional | (3) areas on Item 1, have a crease towards the upstream side which is approximately 6" long and just above the center of Items 5-2, 5-4, & 5-5. See attached photos |  |  | Use As Is | The crease does not appear to be significant. Suggest to proceed with assembly.
458 | Jun-27-2023 | Sep-28-2023 | L2HE-CMA-VV-INSP | 7 | VESSEL LCLS-II AFTER MACHING WELDMENT | RFCHE-VV-016 | Dings Scratches | There are (8) drawing item numbers that have non-conforming issues. See attached QC Task List. All photos are on Server: M:\cmm\Pictures\Inspection Pictures\L2HE\Vacuum vessel\VV016 |  |  | Conforming | Cryomodule assembly group has performed polishing of surfaces with scratches. VV 016 is getting ready to be installed onto J1.3-25 cryomodule
459 | Jun-30-2023 | Jul-03-2023 | L2HE-INSP-BLBP | 65 | WELDMENT BELLOWS PRCM | 114 | Scratches | Mild-Moderate scratch measures 0.078" (5/64ths) caused during borescope use. |  |  | Use As Is | Use as is - hasn't gone through plating. Note scratch on accept tag | MICHELLE O - This comment is confusing, bellows looks to have Cu plating, Datamine shows nothing else was done, cleaned and used in string.
460 | Jul-05-2023 | Aug-08-2023 | L2HE-INSP-BLBP | 66 | WELDMENT BELLOWS PRCM | 115 | OTHER | Multiple pits in plating. |  |  | Return To Vendor
461 | Jul-06-2023 | Aug-08-2023 | L2HE-INSP-BLBP | 67 | WELDMENT BELLOWS PRCM | 116 | OTHER | Clusters of pitted areas. Raised metal spot on copper conv. could be a blister or plated over debris/material; can be felt with a Q-tip and visible with borescope (measures ~3/64".) MORE PHOTOS ON SERVER |  |  | Return To Vendor
462 | Jul-07-2023 | Jul-18-2023 | L2HE-INSP-BLBP | 69 | WELDMENT BELLOWS PRCM | 118 | OTHER | Several white spots/oddities on plated convolutions cannot be removed with alcohol. | Hold/Modify | Send to chemistry to see if spots can be removed | Use As Is | Use as SPARE
463 | Jul-10-2023 | Aug-08-2023 | L2HE-INSP-BLBP | 70 | WELDMENT BELLOWS PRCM | 119 | OTHER Plating | White imperfections on copper convolutions measure up to 3/8". Multiple clusters of pits in plating. |  |  | Return To Vendor
464 | Jul-12-2023 | Oct-13-2023 | L2HE-INSP-TUNC | 23 | LEVER TUNER-15.66 | 0092 | OTHER | During the removal of the (2) FC0045779 BEARINGS from the F10147500 WLDMT, TUNER SUPPORT arm for CMM inspection, The bearings had to be knocked out due to a light press fit. All other tuners inspected have had a slight slip fit during removal. After the bore diameters are corrected, the assembly will need to be check fit to the fixture. | Hold/Modify | The frame will be sent to the machine shop to correct the dimension error around the bearing. This frame will be re-inspected afterwards. Another frame from this lot will also be inspected. | Conforming | After machining, the bearing fit in the arm is acceptable.
465 | Jul-13-2023 | Jul-18-2023 | L2HE-CMA-FPCW-INSP | 30 | FPC WARM PART | CPI-FPCW-HE-132 | Plating | Dark Spots on copper plating inside DN40 Flange |  |  | Use As Is | Spots will not affect operation
466 | Jul-18-2023 | Jul-21-2023 | L2HE-CMA-FPCW-INSP | 31 | FPC WARM PART | CPI-FPCW-HE-134 | Mechanical | Push rod bellows damaged, twisted convolutions. | Hold/Re-measure | Reinstall push rod and allow to sit in a relaxed state for a day, then reinspect | Return To Vendor
467 | Jul-19-2023 | Jul-19-2023 | L2HE-CLNRM-FPCC-INSP | 47 | FPC COLD PART | CPI-FPCC-HE-085 | Oxidation Residues Stains | discoloration/oxidation possibly |  |  | Use As Is | Light oxidation will not affect operation
468 | Jul-19-2023 | Jul-19-2023 | L2HE-CMA-FPCW-INSP | 32 | FPC WARM PART | CPI-FPCW-HE-137 | N/A | Rejected because of marks on copper plating at DN40 Flange |  |  | Return To Vendor
469 | Jul-19-2023 | Jul-19-2023 | L2HE-CLNRM-FPCC-INSP | 47 | FPC COLD PART | CPI-FPCC-HE-086 | Oxidation Stains | discoloration, possibly oxidation |  |  | Return To Vendor
470 | Jul-19-2023 | Jul-19-2023 | L2HE-CLNRM-FPCC-INSP | 46 | FPC COLD PART | CPI-FPCC-HE-127 | Oxidation Stains | discoloration/oxidation |  |  | Return To Vendor
471 | Jul-20-2023 | Jul-20-2023 | L2HE-CLNRM-FPCC-INSP | 45 | FPC COLD PART | CPI-FPCC-HE-107 | Oxidation Stains |  |  |  | Use As Is | Oxidation will not affect performance
472 | Jul-20-2023 | Aug-08-2023 | L2HE-INSP-BLBP | 75 | WELDMENT BELLOWS PRCM | 035 | Dimensional Parallelism Plating | Out of box dimensional damage not caused by shipping damage. Warp in copper plating-R side. |  |  | Return To Vendor
473 | Jul-24-2023 | Jul-24-2023 | L2HE-INSP-BLBP | 76 | WELDMENT BELLOWS PRCM | 125 | Scratches | Large scratch in NR plated flat area measuring 5/16". |  |  | Use As Is | Scratch has not gone through plating
474 | Jul-24-2023 | Jul-25-2023 | L2HE-VTA-CAV-VTRF | 165 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R107 | RFTestFailure VTARFPerformance | Cavity quench limited at 18 MV/m. |  |  | Reject | CTB 2023-07-25 Send to FNAL for HF rinse
475 | Jul-24-2023 | Apr-17-2025 | L2HE-VTA-CAV-VTRF | 167 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R162 | RFTestFailure VTARFPerformance | HOMs Q with very high values (more than 1e20 on both HOMs). Suppose both HOM cables are not connected to the cavity | Hold/Re-measure | connect cables and retest | Use As Is | Cables were disconnected. Cavity met spec in test documented in traveler L2HE-VTA-CAV-VTRF 3173
476 | Jul-25-2023 | Jul-25-2023 | L2HE-CLNRM-FPCC-INSP | 49 | FPC COLD PART | CPI-FPCC-HE-121 | Oxidation Stains | unusual discoloration |  |  | Return To Vendor | Return to vendor installed on the stand with CPI-FPCC-HE-122
477 | Jul-25-2023 | Jul-25-2023 | L2HE-CLNRM-FPCC-INSP | 49 | FPC COLD PART | CPI-FPCC-HE-122 | Oxidation Stains | discoloration |  |  | Return To Vendor | Return to vendor installed on stand with CPI-FPCC-HE-121
478 | Jul-25-2023 | Apr-22-2025 | L2HE-INSP-TUNC | 24 | LEVER TUNER-15.66 | 0091 | Dimensional | F10147497; Plane to plane angles are OOT by .31°. See attached report. | Hold/Modify | Frame will be modified by machine shop if needed during installation. | Conforming | Frame was installed in J27 without need for modification from the machine shop. Correction made 22-Apr-2025 by ashleya changed J28 to J27 - powen confirmed typo in comment above, was J28, should be J27
479 | Jul-25-2023 | Sep-25-2023 | L2HE-INSP-TUNC | 24 | LEVER TUNER-15.66 | 0091 | Dimensional OTHER | During the removal of the (2) FC0045779 BEARINGS from the F10147500 WLDMT, TUNER SUPPORT arm for CMM inspection, The bearings and bearing retainers had to be knocked out due to a light press fit. All other tuners inspected have had a slight slip fit during removal. | Hold/Modify | The frame will be sent to the machine shop to correct the dimension error around the bearing. This frame will be re-inspected afterwards. Since this is the second frame of the lot to have this issue, I will have all frames from the lot (#7) to be inspected. | Conforming
480 | Jul-26-2023 | Jul-26-2023 | L2HE-INSP-BLBP | 78 | WELDMENT BELLOWS PRCM | 074 | OTHER Plating | Weld area appears compromised in a few places. | Hold/Re-measure | Leak check | Use As Is | Leak test ok
481 | Jul-28-2023 | Jul-28-2023 | L2HE-INSP-BLBP | 81 | WELDMENT BELLOWS PRCM | 177 | Dings | During Leak Check, the rotating flange fell down and damaged 4 convolution areas. Each of the 4 damaged areas are approximately .5" long, 1/8" wide. See attached photos. |  |  | Reject | Dents caused by leak check accidental damage are cause for rejection
482 | Aug-08-2023 | Aug-09-2023 | L2HE-CLNRM-FPCC-INSP | 52 | FPC COLD PART | CPI-FPCC-HE-099 | OTHER | Large hole through both outer and inner bag over the pump out port to the valve, Adam G and John V were contacted and advised techs to continue cleaning for evaluation in the clean room. |  |  | Use As Is | Use extra caution cleaning due to ripped bag.
483 | Aug-08-2023 | Mar-04-2024 | L2HE-CLNRM-FPCC-INSP | 51 | FPC COLD PART | CPI-FPCC-HE-115 | Vacuum | While doing the procedure to quantify the static vacuum of the FPC box, the pressure observed was close to atmosphere when the valve was opened. This indicates the box was not under vacuum at the time of inspection. |  |  | Return To Vendor | Return to vendor for repair
484 | Aug-08-2023 | Mar-04-2024 | L2HE-CLNRM-FPCC-INSP | 51 | FPC COLD PART | CPI-FPCC-HE-115 | Vacuum | While preforming the bagged helium leak check a leak was observed and sized at 2.57e-6 STD cc/s. |  |  | Return To Vendor | Return to vendor for repair
485 | Aug-08-2023 | Aug-11-2023 | L2HE-CLNRM-FPCC-INSP | 51 | FPC COLD PART | CPI-FPCC-HE-115 | Oxidation | During the inspection of the protection caps significant oxidation was observed on FPC 115. |  |  | Use As Is | Discoloration does not affect FPC performance.
486 | Aug-09-2023 | Aug-09-2023 | L2HE-CLNRM-FPCC-INSP | 52 | FPC COLD PART | CPI-FPCC-HE-100 | OTHER | sam issue as 99 this is the mate to it And amandaa, very large hole in both outer and inner bag, Adam G and John V were contacted and techs were advised to continue carefully with cleaning and send into clean room for their assessment of the coupler cleanliness |  |  | Use As Is | Use extra caution cleaning due to ripped bag.
487 | Aug-09-2023 | Aug-11-2023 | L2HE-CLNRM-FPCC-INSP | 52 | FPC COLD PART | CPI-FPCC-HE-099 | Oxidation Stains | oxidation on copper surface |  |  | Return To Vendor | Return to vendor for cleaning.
488 | Aug-10-2023 | May-12-2025 | L2HE-INV-FPCC | 114 |  | CPI-FPCC-HE-141 | OTHER | Hold as spare. Coupler shipped back to CPI by itself for dark spot on copper, then shipped to FNAL in error and then returned to JLab. Hold as spare due to off normal handling. |  |  | Use As Is | Hold as spare. Coupler shipped back to CPI by itself for dark spot on copper, then shipped to FNAL in error and then returned to JLab. Hold as spare due to off normal handling.
489 | Aug-14-2023 | Nov-16-2023 | L2HE-INV-CAV | 109 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R170 | OTHER | Accelerometer readings: 7.1 VM Max/3.0 Mission Max |  |  | Conforming | Passed RFIN inspection and qualified VTRF, cavity is conforming.
490 | Aug-14-2023 | Sep-19-2023 | L2HE-INV-CAV | 108 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R169 | OTHER | Accelerometer readings: 7.1 VM Max/3.0 Mission Max |  |  | Conforming | Cavity passed RFIN and qualified VTRF, conforming.
491 | Aug-14-2023 | Sep-14-2023 | L2HE-VTA-CAV-VTRF | 179 | 1.3GHZ LCLS-II RF CAVITY ASSY (BARE) | 362 | RFCMTF RFTestFailure VTARFPerformance | Cavity quench at about 23.5 MV/m. High RF power testing affected by non linearities in Qo and Q2 measurements. Edit: Qe2 out of spec (7.5e11 - 2.5e12) | Hold/Modify | Replace field probe using purge system | Use As Is | Qe2 issue, as well as calibration issues, has been resolved, see test of 2023-09-13
492 | Aug-16-2023 | Aug-16-2023 | L2HE-INSP-CAV | 107 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R169 | Scratches | Imperfections/scratches present on burst disc. Pictures attached. |  |  | Use As Is | RF incoming shows no leak, i.e. burst disc should be leak tight
493 | Aug-30-2023 | Oct-13-2023 | L2HE-INSP-TUNC | 26 | LEVER TUNER-15.66 | 0093 | Dimensional | During the removal of the (2) FC0045779 BEARINGS from the F10147500 WLDMT, TUNER SUPPORT arm for CMM inspection, The bearings and bearing retainers opposite of the tuning motor side had to be knocked out due to a press fit condition. | Hold/Modify | Frame will be sent to machine shop to have the bearing holes resized. MS #331876 | Conforming | After machining, the bearing fit in the arm is acceptable.
494 | Sep-01-2023 | Oct-13-2023 | L2HE-INSP-TUNC | 27 | LEVER TUNER-15.66 | 0094 | Dimensional | During the removal of the (2) FC0045779 BEARINGS from the F10147500 WLDMT, TUNER SUPPORT arm for CMM inspection, the bearings and bearing retainers opposite of the tuning motor side had to be knocked out due to a press fit condition. (1) bearing was destroyed during removal with a bearing puller. Bearing was replaced with new bearing. The dia. 26.0mm through holes are undersized and out of round. | Hold/Modify | Frame will be sent to machine shop to have the bearing holes resized. MS #331916 | Conforming | After machining, the bearing fit in the arm is acceptable.
495 | Sep-01-2023 | Apr-29-2025 | L2HE-INSP-TUNC | 27 | LEVER TUNER-15.66 | 0094 | Dimensional | Part# F10147497; Angle: Should Be: 5.0° ±1° = 3.78°, OOT -1.22° | Hold/Modify | Frame will be modified by machine shop if needed during installation. | Use As Is | Frame was installed on J29 cavity 6 without issue. No modifications were necessary.
496 | Sep-11-2023 | May-06-2025 | L2HE-INV-FPCC | 117 | FPC COLD PART | CPI-FPCC-HE-165 | OTHER | Per Naeem, since this FPCC was shipped as a single it needs to be designated and held as a spare. |  |  | Use As Is | This FPCC was shipped in an off-normal manner as a single unit; it needs to be designated and held as a spare.
497 | Sep-12-2023 | Nov-16-2023 | L2HE-INV-CAV | 110 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R174 | OTHER | Accelerometer: 5.46g exceeds 3.00g Max |  |  | Conforming | Passed RFIN inspection and qualified VTRF, cavity is conforming.
498 | Sep-12-2023 | Nov-16-2023 | L2HE-INV-CAV | 111 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R175 | OTHER | Accelerometer: 5.46g exceeds 3.00g Max |  |  | Conforming | Passed RFIN inspection and qualified VTRF, cavity is conforming.
499 | Sep-12-2023 | Nov-16-2023 | L2HE-INV-CAV | 112 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R177 | OTHER | Accelerometer: 4.36g exceeds 3.00g Max |  |  | Conforming | Passed RFIN inspection and qualified VTRF, cavity is conforming.
500 | Sep-12-2023 | Nov-16-2023 | L2HE-INV-CAV | 113 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R178 | OTHER | Accelerometer: 4.36g exceeds 3.00g Max |  |  | Conforming | Passed RFIN inspection, cavity is conforming.
501 | Sep-12-2023 | Nov-16-2023 | L2HE-INV-CAV | 114 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R179 | OTHER | Accelerometer: 7.32g exceeds 3.00g Max |  |  | Conforming | Passed RFIN inspection and qualified VTRF, cavity is conforming.
502 | Sep-12-2023 | Nov-16-2023 | L2HE-INV-CAV | 115 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R180 | OTHER | Accelerometer: 7.70g exceeds 3.00g Max |  |  | Conforming | Passed RFIN inspection, cavity is conforming.
503 | Sep-13-2023 | Mar-01-2024 | L2HE-LERF-CM-DMAG | 2 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-23 | OTHER | After two rounds of 65A demag and a few .65A demag, the cavity 1 fluxgate reads ~ 20 mG, which is highly above the specification of 2.5 mG. It is suspected that there is magnetic components at the Cryocan end. Preliminary survey of the interconnect bellows and its stiffening rods show high, peak at ~ 2.7 Gauss, remanent field. | Hold/Modify | Demagnetize the interconnection bellows. Monitor cavity 1 FG reading | Use As Is | Cavity 1 Q0 is measured to be 2.86e10, which is above the average Q0 acceptance criterion of 2.7e10.
503 | Sep-13-2023 | Mar-01-2024 | L2HE-LERF-CM-DMAG | 2 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-23 | OTHER | After two rounds of 65A demag and a few .65A demag, the cavity 1 fluxgate reads ~ 20 mG, which is highly above the specification of 2.5 mG. It is suspected that there is magnetic components at the Cryocan end. Preliminary survey of the interconnect bellows and its stiffening rods show high, peak at ~ 2.7 Gauss, remanent field. | Hold/Modify | Demagnetize the interconnection bellows. Monitor cavity 1 FG reading | Use As Is | Cavity 1 Q0 is measured to be 2.86e10, which is above the average Q0 acceptance criterion of 2.7e10.
504 | Sep-15-2023 | Sep-15-2023 | L2HE-INSP-CAV | 109 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R174 | Dimensional | PDF attached. Dimensional non-conforming. |  |  | Use As Is | Acceptable deviation during tuner installation
505 | Sep-18-2023 | Feb-20-2024 | L2HE-VTA-CAV-VTRF | 188 |  | HE-R132 | RFCMTF RFTestFailure VTARFPerformance | Strong multipacting over 16.5 MV/m associated with cavity gradient drop and significant Qo drop. No quench events up to 17 MV/m. | Hold/Modify | Go for another big HPR (CTB 2023-09-26) | Use As Is | Use as is = retest: see NCR 405 for final disposition.
506 | Sep-18-2023 | Jan-19-2024 | L2HE-VTA-CAV-VTRF | 189 |  | HE-R113 | RFCMTF RFTestFailure VTARFPerformance | Cavity quench limited at 21.8 MV/m | Hold/Modify | Go for another big HPR (CTB 2023-09-26) | Reject | CTB 2023-12-19 send to Fermi/Argonne for HF rinse
507 | Sep-19-2023 | Sep-19-2023 | L2HE-CLNRM-FPCC-INSP | 53 | FPC COLD PART | CPI-FPCC-HE-063 | OTHER | FPC 063 RF pin and bolt missing |  |  | Use As Is | Pin and bolt are shipped separately
508 | Sep-20-2023 | Dec-19-2023 | L2HE-LERF-CM-PREP | 3 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-23 | Vacuum | In epics warm insulating vacuum reads back 0.0017 Torr, which exceeds the required "Below 1e-4" specification. Epics reading from Pirani only, CCG not yet active. Local gauges on turbo-pumps read e-5 range. | Hold/Re-measure | We decided to check the local vacuum gauge to confirm the vacuum pressure | Use As Is | What we learned from this is that for the future L2HE CM test in LERF, we will need to make sure the insulating vacuum CCG is active
509 | Sep-22-2023 | Sep-26-2023 | L2HE-INSP-CAV | 113 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R179 | Dimensional | Parallelism of 2-phase TEE OOT. Should be 0-1.0, = 4.2. |  |  | Use As Is | acceptable for string assembly
510 | Sep-22-2023 | Sep-26-2023 | L2HE-INSP-CAV | 113 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R179 | Dings | Two small indentations noted during visual inspection: One is from the inside to outer surface and the 2nd is from outer to inner surface. |  |  | Use As Is | discussed the scratches with CTB on 2023-09-26
511 | Sep-25-2023 | Sep-26-2023 | L2HE-INSP-CAV | 114 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R180 | Scratches | Scratches present on bellows. Please see attached. |  |  | Use As Is | discussed the scratches with CTB on 2023-09-26
512 | Sep-27-2023 | Nov-07-2023 | L2HE-INSP-BLBS | 13 | WELDMENT BELLOWS SHORT | 006 | Dimensional Parallelism Plating | As rec'd. damage - bent convolutions; Warped plating; Questionable weld. Returning RTV- damage not seen in initial inspection photos. |  |  | Reject | Leak check happened prior to inspection, can't prove to vendor this damage happened prior to receipt.
513 | Oct-04-2023 | Oct-13-2023 | L2HE-CMA-UCM-INSP | 4 | LCLS-II UPPER COLD MASS ASSEMBLY | RFCHE-UC-008 | Dings | Bellows damage found. Dent/ding present. Pictures attached. |  | Let's try to measure the depth of the dent. The bellows has 3 plies, each is .3 mm or .012" thick. The inner ply holds vacuum, see drawing and email attached | Use As Is | The ~ .003" dent is deemed minor since the single-ply thickness is .012"
514 | Oct-04-2023 | Oct-05-2023 | L2HE-INSP-IMAG | 67 | MAGNETIC SHIELD CAVITY ASSEMBLY | RFCHE-MS-122 | Dimensional | Two item FC10056532 spacers are 1/32 to 1/16" short. |  |  | Use As Is | Advised the cryomodule assembly team to add additional spacers such as washers
515 | Oct-06-2023 | Feb-20-2024 | L2HE-VTA-CAV-VTRF | 194 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R179 | VTARFPerformance | Cavity performance met spec for gradient and Qo, but failed for initial FE onset at 16.3 MV/m, and final onset at 15.3 MV/m. Also, Qext HOMB is below spec at 8.5e10. | Hold/Modify | Cavity will undergo big HPR Prior to next test, HOM_B will be retuned | Conforming | Cavity passed in test on 2024-01-19
516 | Oct-09-2023 | Oct-09-2023 | L2HE-INSP-CAV | 115 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R182 | Dings | Ding on bellows. Pictures attached. |  |  | Use As Is | Acceptable as it is shallow and no sharp edges are present Discussed with SLAC (James) on 23-10-09 | Very nice comment for the disposition!
517 | Oct-10-2023 | Oct-10-2023 | L2HE-VTA-CAV-VTRF | 188 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R132 | VTARFPerformance | At 2K, an HOM survey was performed on the cavity. This measures the Ql of HOM modes, as well as the 9 passband frequencies. The passband frequencies show some deformation in the cavity with an M-parameter of 5.6 kHz when compared to this cavity's previous vertical test. This is within the elastic range of the cavity (10 kHz), but still unexpected during vertical test. For comparison, HE-113 was also tested in D5 and measured in a similar way. The M-parameter of -113 was only 0.3 kHz. |  |  | Conforming | The cavity had water/moisture in its helium vessel which most likely caused a deformation after cooldown and corresponding freezing. Subsequent measurements, see attached files, show that the deformation was only temporary i.e. elastic.
518 | Oct-11-2023 | Jan-31-2024 | L2HE-CMA-VV-INSP | 8 | VESSEL LCLS-II AFTER MACHING WELDMENT | RFCHE-VV-014 | Residues Roughness Scratches | There are (10) drawing item numbers that have non-conforming issues. See attached QC Task List. All photos are on Server: M:\cmm\Pictures\Inspection Pictures\L2HE\Vacuum vessel\VV0164 | Hold/Modify | The cryomodule assembly group will use the visual inspection findings as a guide to clean up, polish sealing surfaces | Conforming | VV 014 has been installed onto J1.3-26 cryomodule. Necessary polishing is done
519 | Oct-12-2023 | Mar-04-2024 | L2HE-LERF-CM-ACTS | 2 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-23 | CMTFRFperf Mechanical | At 4K and 2K, TE-DH-1320 (Cav3 HOM1, coupler-end) temperature is 6 and 4 degrees higher than other HOMs. Temperatures prior to cooldown were all within 0.3 K of 294.5K. High power testing showed Cav3 to quench at low gradient (8.7 MV/m.) | Hold/Re-measure |  | Conforming | The root cause analysis of abnormal temperature rises from cavity 3, HOM 1 sensor indicates an issue with the thermal connection. Subsequent thermal strap heat transfer testing led to the replacement of the thermal strap, confirming improved heat transfer. Heat transfer testing has been completed on J1.3-25, and the results have proven to be satisfactory. A similar test will be performed on J1.3-26 as well. Additionally, it has been observed that on the clamp F10048352 (refer to the attached document), holes at Datum B have been added. These holes were not present at the beginning of LCLS-II. Assembly technicians have mentioned that they do need to drill these holes on many F10048352's but not all. Furthermore, the holes at the cavity 1 HOM 1 thermal clamp have been enlarged. J1.3-24 has not undergone inspection for similar holes as it was in testing. However, J1.3-25 has undergone clamp hole inspection, and J1.3-26 will undergo inspection as well. It is believed that the pre-machined holes are added to later F10048352 clamps.
520 | Oct-13-2023 | Jan-30-2024 | L2HE-LERF-CM-ACTS | 2 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-23 | CMTFRFperf Mechanical RFCMTF RFTestFailure | Cavity 1 FPC Qext adjustment will not adjust above 2.7E7 (bandwidth measurement) or 2.6E7 (time decay measurement). This value does not meet the Minimum Acceptance Criteria of 6E7. Having to operate with a Qext this low may also constrain the maximum usable gradient. We were able to get the cavity to 20.8 MV/m but may not have enough forward power headroom to operate reliably in SELAP mode. |  |  | Use As Is | We were able to adjust the Qext to 4.3E7 after adjustments made during the mid-test warm up. However, 4.3E7 is still below the Minimum Acceptance Criteria of 6E7 and does not meet the range requirement of 1E7 - 8E7. Cavity 1's gradient reach is not limited by a Qext of 4.3E7 but coupler heating will be higher than typical.
521 | Oct-17-2023 | Mar-04-2024 | L2HE-LERF-CM-ACTS | 2 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-23 | Instrumentation | Three (3) DT-670A1-CU (bobbin-style, diode temperature sensors on HOMs) failed during insulating vacuum pumpdown and cryomodule cooldown of J23. TE-DH-1120 (Cav1 HOM1, coupler-end), TE-UH-1318 (Cav3 HOM2, tuner-end), and TE-UH-1718 (Cav7 HOM2, tuner-end) all read 295 Kelvin following hookup in LERF vault, but read 0.0K during 2K operation. | Hold/Re-measure |  | Conforming | Up to five temperature sensors were found to be malfunctioning. Replacement sensors were added, with the faulty sensors abandoned but not removed. All replacement sensors passed electrical check-out
522 | Oct-18-2023 | Dec-19-2023 | L2HE-LERF-CM-ACTS | 1 |  | J1.3-22 | Vacuum | L2HE J22 was found to have a helium process leak into the insulating vacuum space. On 2nd (And subsequent) Cooldowns a cold leak opened and insulating vacuum began to degrade. Insulating vacuum was unable to be maintained below 1e-3 Torr, cooldown was aborted to investigate. | Hold/Re-measure | Per discussion with Matt Weaks: investigation in to the He Process leak shows a cold leak from the 2K (Low Pressure) circuit into the insulating vacuum. During cooldown with active pumping, insulating vacuum degrades and rises above ~1e-3 Torr. Initial testing in LERF is unable to determine the specific source of the leak. | Conforming | The leak into insulating vacuum from the helium process line on J22 is determined not caused by the J22 cryomodule. It's rather the interface between J22 and the interconnection bellows & cryo can in LERF that leaked. Refer to CAPA 051 (https://jlabdoc.jlab.org/docushare/dsweb/View/Collection-56874) for more details. Also worth noting is that J22 heads into rework not due to the leak to insulating vacuum
523 | Oct-30-2023 | Nov-01-2023 | L2HE-INSP-CAV | 116 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R130 | Dimensional | The following are OOT: 1) DIM1, 300.0 +0.5 - 1.5= 301.0, OOT .5mm 2) DIM2, 1.0= 1.5, OOT .5 3) DIM12, 5 +/- 1.0, OOT .6° |  |  | Use As Is | Can be tolerated in string assembly
524 | Nov-01-2023 | Nov-01-2023 | L2HE-INSP-CAV | 117 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R181 | Dimensional | The following are OOT: 1) DIM2, 1.0= 1.5, OOT 1.5 2) DIM13, 272 +0.35-0.65, OOT .06 |  |  | Use As Is | Can be tolerated in string assembly
525 | Nov-07-2023 | Jan-23-2024 | L2HE-CMA-UCM-INSP | 9 | LCLS-II UPPER COLD MASS ASSEMBLY | RFCHE-UC-012 | Dings | Damage present on conflat flanges on 5k tube, 2.2k tube and 50k shield tube. Please see attached. |  |  | Use As Is | The damages appear on flanges that are for JLAB testing use only. Cryomodule assembly group is aware of the issues. They will polish the flanges if needed at the time of assembly
526 | Nov-07-2023 | Jan-23-2024 | L2HE-CMA-UCM-INSP | 9 | LCLS-II UPPER COLD MASS ASSEMBLY | RFCHE-UC-012 | OTHER | Burr/sharp edge present on one hanger. Damage found on different hanger also, it is bent slightly. Please see attached. |  |  | Use As Is | J1.3-027 cavity string has been attached and aligned to the UCM 012
527 | Nov-15-2023 | Nov-15-2023 | L2HE-CLNRM-AMGV-LEAK | 10 | ALL METAL GATE VALVE | 0041 | LeakTestFailure | Original traveler being updated to leak check bonnet. Leak check on upstream subassembly revealed bonnet leaked. | Hold/Modify | The AMGV will be sent back to vendor for repairs. | Return To Vendor | The valve will be shipping back to the vendor to undergo repair.
528 | Nov-16-2023 | Nov-16-2023 | L2HE-INSP-CAV | 118 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R183 | Dings | Damage present on pin. Please see attached. |  |  | Use As Is | This type of damage is not affecting CM assembly, especially since it is on the side of the pin.¦ In general it is concerning though to see such kind of damage. I will reach out to the vendor/SLAC to keep them informed.
529 | Nov-20-2023 | Apr-29-2025 | L2HE-INSP-CAV | 118 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R183 | Dimensional | Please see attached. |  |  | Use As Is | After discussion with CM assembly group (John F) deemed acceptable for string build
530 | Nov-20-2023 | Feb-18-2025 | L2HE-CLNRM-GV2SA-ASSY | 8 | BEAM POSITION MONITOR FEEDTHROUGH | 10 | Dings | BPM Probe number 10 has ding on surface |  |  | Reject | Rejected for use due to leak check failure. Hold as spare, can be repaired if needed.
530 | Nov-20-2023 | Feb-18-2025 | L2HE-CLNRM-GV2SA-ASSY | 8 | BEAM POSITION MONITOR FEEDTHROUGH | 10 | Dings | BPM Probe number 10 has ding on surface |  |  | Reject | Rejected for use due to leak check failure. Hold as spare, can be repaired if needed.
531 | Nov-21-2023 | Nov-27-2023 | L2HE-INSP-BLBS | 11 | WELDMENT BELLOWS SHORT | 020 | Dings | Small, shallow dent was found in convolutions during assembly in clean room. |  |  | Use As Is | Dent measures 1/16", no indication of a sharp point
532 | Nov-28-2023 | Feb-26-2024 | L2HE-CMA-CM-ASSY2 | 5 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-26 | Instrumentation | Discovered cavity #5 fluxgate cable (MAG-F sensor, ME-C5-2215), was severed just outside of magnetic shield. Attempting repair. | Hold/Modify | Please proceed with the repair | Conforming | The repair is complete, and the electrical check-out has confirmed its success
533 | Dec-12-2023 | Jan-19-2024 | L2HE-VTA-CAV-VTRF | 198 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R183 | RFTestFailure VTARFPerformance | Field Emission at 3.5 MV/m, did not process away | Hold/Modify | re-rinse & re-test big HPR | Conforming | Cavity had a clean test on 2024-01-19
534 | Dec-14-2023 | Dec-15-2023 | L2HE-INSP-CAV | 120 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R186 | Scratches | Bellows, supply line side, has (3) small areas that have scratches/cut areas. See photos. |  |  | Use As Is | Dings/scratches are acceptable (compare slides provided by SLAC)
535 | Dec-14-2023 | Dec-19-2023 | L2HE-INSP-CAV | 120 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R186 | OTHER | Partially demagnetized. The "Tee-transition joint" is OOT. Should be: .25+.40(background), .65 Gauss; Is: 1.0 Gauss; .35 Gauss OOT. Several attempts with the two hand-held demagnetizers did not work in Tee area. See photos. | Hold/Modify | Please try to degauss T with Maurer degausser in north annex. | Use As Is | Proceed with VTRF to evaluate impact on cavity quality factor.
536 | Dec-15-2023 | Jan-10-2024 | L2HE-VTA-CAV-VTRF | 200 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R132 | RFTestFailure | Could not test cavity on test stand 34-002 position 2. Low power signals showed continuity, cavity frequency could be measured with VNA and labview but during cavity coupling check signal was found to have no decay curve indicating most of the power being reflected due to a poor connection on the input cable. |  |  | Conforming | RF cable issue has been resolved prior to re-test on 2023-01-03
537 | Dec-18-2023 | Dec-18-2023 | L2HE-INSP-CAV | 119 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R185 | Scratches | Some light cuts/scratches noted on OD convolutions in two areas. See attached photos. |  |  | Use As Is | Dings/scratches are acceptable (compare slides provided by SLAC)
538 | Jan-05-2024 | Jan-08-2024 | L2HE-INSP-CAV | 121 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R086 | Dimensional | 1) Parallelism of 2-phase TEE OOT. Should be 0-1.0, = 1.1 2) Location of tuner ring (split ring) circle OOT. Should be 5° +/- 1.0°, = 6.9° OOT by 1.9° |  |  | Use As Is | 1) can be tolerated/compensated for during 2 phase pipe welding 2) should be adjustable during tuner assembly alternatively, individual adaptation of tuner arms might be required -- individual machining of these arms is common practice
539 | Jan-09-2024 | Mar-01-2024 | L2HE-CMA-FPCW-INSP | 18 | FPC WARM PART | CPI-FPCW-HE-020 | LeakTestFailure | Sapphire window leak. 2.45e-10 STD CC/sec leak rate | Hold/Modify | Replace sapphire window and retest | Conforming | Leak has been corrected with replacement window
540 | Jan-10-2024 | Jan-11-2024 | L2HE-INSP-CAV | 122 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R078 | Scratches | All (4) Support Blocks have some scratches on top & bottom sides. This cavity was apparently installed and removed. |  |  | Use As Is | Use as is, light scratches don't affect performance.
541 | Jan-10-2024 | Jan-16-2024 | L2HE-INSP-CAV | 122 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R078 | Dimensional | DIM2: // of 2-phase TEE OOT; Should be 1.0°, Is 1.8° |  |  | Use As Is | Can be compensated for during 2p pipe welding
542 | Jan-12-2024 | Mar-01-2024 | L2HE-CMA-FPCW-INSP | 23 | FPC WARM PART | CPI-FPCW-HE-021 | LeakTestFailure | Sapphire window leak on FPCW installed on Cav5. 2.81e-9 STD cc/sec leak rate | Hold/Modify | Replace sapphire window and retest | Conforming | Leak has been corrected with replacement window
543 | Jan-12-2024 | Mar-01-2024 | L2HE-CMA-FPCW-INSP | 26 | FPC WARM PART | CPI-FPCW-HE-036 | LeakTestFailure | Sapphire window leak on FPCW installed on Cav8. 3.54e-10 STD cc/sec leak rate | Hold/Modify | Replace sapphire window and retest | Conforming | Leak has been corrected with replacement window
544 | Jan-12-2024 | Jan-12-2024 | L2HE-INSP-CAV | 123 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R105 | OTHER | Loose material/fibers inside tee. Please see attached pictures. |  |  | Use As Is | SLAC has been notified of this discrepancy with the cavity. No further action is required. Does not affect performance.
545 | Jan-17-2024 | Jan-17-2024 | L2HE-INSP-CAV | 124 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R094 | Dimensional | (2) dimensions are OOT: DIM1: Should be; 300.0 +0.5 - 1.5, Is: 300.6 DIM2: Should be; 1.0°, Is: 2.1° |  |  | Use As Is | Can be tolerated/compensated for during 2p pipe welding.
546 | Jan-30-2024 | Mar-01-2024 | L2HE-LERF-CM-ACTS | 2 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-23 | CMTFRFperf Mechanical RFCMTF RFTestFailure | Cavity 2 has a measured QextFPC of 5.7E7. This is at the maximum adjustment available. Does not meet the Minimum Acceptance Criteria of 6E7. The range of adjustment is also narrower than the 1E7 - 8E7 span required by the Minimum Acceptance Criteria |  |  | Use As Is | Despite the lower measured Qext, cavity 2 tested normal. The FPC can be adjusted within limits next time when it is at 2k and the cavity is tuned.
547 | Jan-30-2024 | Feb-27-2024 | L2HE-LERF-CM-ACTS | 2 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-23 | CMTFRFperf Mechanical RFCMTF RFTestFailure | None of the FPC's on J23 have shown an adjustment range that reaches 8E7 |  |  | Use As Is | We have adjusted cavities 1 & 2 so that they are not power-limited. Others were good
548 | Jan-30-2024 | Feb-27-2024 | L2HE-LERF-CM-ACTS | 2 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-23 | CMTFRFperf RFCMTF RFTestFailure | Only one cavity (cavity 7) has a measured Qo that is equal to or higher than 2.7E10. Average Qo for the module is 1.9E10 |  |  | Use As Is | We suspect that several cavities quenched hence degraded the Qo
549 | Feb-01-2024 | Mar-18-2024 | L2HE-INSP-FPFT | 45 | PICKUP FEEDTHROUGH | 2033 | LeakTestFailure Vacuum | FPFT s/n 2033 has been verified as installed on J22 as cavity #7 (CAV-R031) and is suspected to be the source of the BL leak documented on NCR 451. Failed part to be segregated upon disassembly and investigated for root cause of failure. Also see JLab CAPA 051. | Hold/Re-measure | Upon string disassembly, it was discovered that the leak was actually on cavity #8 (s/n R012) with FPFT s/n 2014 installed. This was the result of mis-identifying the cavity number while in the vacuum vessel due to limited visibility. | Conforming | FPFT 2033 is conforming. See NCR 578 for status of bad FPFT s/n 2014.
550 | Feb-01-2024 | Feb-02-2024 | L2HE-INSP-BLBP | 90 | WELDMENT BELLOWS PRCM | 070 | Stains | Both flange seal surfaces have a discolored yellow/bronze tint, stain or film; Will not clean with IPA, Acetone or DI water. | Hold/Modify | Send to chemistry for cleaning | Use As Is | Stains have been removed
551 | Feb-06-2024 | Feb-27-2024 | L2HE-VTA-CAV-VTRF | 212 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R169 | RFTestFailure VTARFPerformance | Cavity had initial FE onset 19.98 MV/m, max radiation 0.763mR/hr. Performed 20 minutes of quench processing at 50w. Resulted in max 0.447 mR/hr at 25.85 MV. | Hold/Modify | Go for a "big HPR" i.e. full disassembly and USC | Conforming | Cavity qualified on 2024-02-21
552 | Feb-08-2024 | Feb-08-2024 | L2HE-INSP-BLBP | 91 | WELDMENT BELLOWS PRCM | 072 | Dimensional | Length, seal surface depth and bellows ID out of tolerance. Please see attached for dimensional non-conforming. |  |  | Use As Is | OOT of length and ID ok, OOT for seal surface depth less so. Keep as emergency use spare only and mark EMERGENCY USE SPARE.
553 | Feb-14-2024 | Feb-20-2024 | L2HE-VTA-CAV-VTRF | 213 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R143 | VTARFPerformance | Qext2 was 2.69e12, out of spec (7.5 e11 ~ 2.5 e12) |  |  | Use As Is | Accepted by CTB on 2024-02-20
554 | Feb-19-2024 | Sep-24-2024 | L2HE-VTA-CAV-VTRF | 215 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R094 | RFTestFailure VTARFPerformance | FE onset at 13.7 MV/m. | Hold/Modify | re-rinse & re-test (big HPR -- full disassembly) | Return To Vendor | 2024-09-24 sent to FNAL after big HPR for test in VTS
555 | Feb-19-2024 | Dec-03-2024 | L2HE-LERF-CM-ACTS | 3 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-24 | CMTFRFperf Mechanical RFCMTF RFTestFailure | Cavity 1 does not meet the tuning range criteria of +/- 20 kHz. Has an actual range of +7/-20 kHz | Hold/Re-measure | A tuner inspection plan has been developed to check if the cavity tuning range was set properly during tuner installation. The plan is to be executed to determine what correction is needed | Conforming | The G10 plunger of cavity 1 tuner is modified to increase tuning range to meet the spec of +/- 20kHz
556 | Feb-20-2024 | Feb-20-2024 | L2HE-INSP-BLBP | 94 | WELDMENT BELLOWS PRCM | 115 | OTHER | Questionable weld. Non-uniform appearance with cracks observed. |  |  | Reject | Mark as emergency use spare only pending leak check if necessary for use.
557 | Feb-20-2024 | Feb-20-2024 | L2HE-INSP-BLBP | 94 | WELDMENT BELLOWS PRCM | 115 | Scratches | Radial scratch in R Flange seal surface may have slight depth. There are no variations in Ra surface finish measurements. |  |  | Reject | Mark as emergency use spare only pending leak check if necessary for use
558 | Feb-20-2024 | Feb-20-2024 | L2HE-INSP-BLBP | 94 | WELDMENT BELLOWS PRCM | 115 | OTHER | Crack in non-critical R Flange surface appears to be covered with dark colored sealant. |  |  | Reject | Mark as emergency use spare only pending leak check if necessary for use
559 | Feb-27-2024 | Apr-23-2024 | L2HE-CMA-CM-ASSY1 | 7 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-28 | RFCMTF | Cavity 178 in position 2 FP attenuation -37.76 dB , normal attenuation -7.5 dB See NCR 507 |  | Install Pongy on coupler in flow hood and back fill with N2 | Conforming | The FP attenuation is acceptable after repair
560 | Feb-27-2024 | Feb-28-2024 | L2HE-INSP-CAV | 127 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R022 | Dimensional | Two dimensions are OOT: DIM1) Should Be: 300.0 +0.5 - 1.5; Is: 300.8 DIM2) Should Be: 0 to 1.0 ; Is: 1.7 See attached report. |  |  | Use As Is | Will be compensated for during 2 phase pipe welding
561 | Feb-28-2024 | Mar-01-2024 | L2HE-CMA-CM-SHIP | 1 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-23 | Instrumentation | Low resistance measured on J23 cav 7 FPC_top temperature sensor RTD. Sensor TE-CA-1712 on Cav_7-Flange_A02, pins E,F,G,H fails due to 4-wire measurement of 9.9 ohms (108 ohms nominal). Resistance of V+/V- and I+/I- legs are normal (~68 ohms) indicating that the lead extensions are intact and that the low resistance failure is at or near the sensor-end of the wiring harness. All other sensor measurements on the Shipping Instrumentation Checks of Step 27 are good, including the complementary cav 7 FPC_bottom temperature sensor RTD. Sensor TE-CA-1712 was last measured, and found to be good, on 23 Feb 2024. |  |  | Use As Is | The problematic sensors are inaccessible w/o disassembly of the cryomodule back to WS4
562 | Feb-28-2024 | Feb-29-2024 | L2HE-CLNRM-FPCC-INSP | 59 | FPC COLD PART | CPI-FPCC-HE-083 | Oxidation | oxidation on copper surface |  |  | Return To Vendor | Return to CPI to clean plating
563 | Feb-28-2024 | Feb-29-2024 | L2HE-CLNRM-FPCC-INSP | 59 | FPC COLD PART | CPI-FPCC-HE-084 | Oxidation | oxidation on copper surface |  |  | Return To Vendor | Return to CPI for plating cleaning
564 | Feb-28-2024 | Dec-03-2024 | L2HE-LERF-CM-PREP | 4 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-24 | Instrumentation | (4) HOM temperature sensor diodes failed at or shortly after initial instrumentation hookup. TE-UH-1318 (Cav3 HOM2, tuner-end) failed at hookup. TE-DH-1320 (Cav3 HOM1, coupler-end) read room temp, then failed 1 hour after hookup. TE-UH-1518 (Cav5 HOM2, tuner-end) failed at hookup. TE-DH-1820 (Cav8 HOM1, coupler-end) read room temp, then failed 18 hours after hookup. | Hold/Modify | Replace/repair when feasible | Use As Is | Refer to NCR# 584 for details on repair of the temperature sensors
565 | Feb-28-2024 | Feb-28-2024 | L2HE-LERF-CM-ACTS | 2 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-23 | CMTFRFperf RFCMTF RFTestFailure | Cavity 3 was performance limited to 8.9 MV/m during acceptance testing due to HOM heating which caused early quench. This issue and corrective action are documented in NCR 519, including the repair; it is signed off. Cavity 3 was not able to be retested after the HOM thermal station repair. It is expected that cavity 3 will perform better with the repaired HOM thermal strap but was not verified due to schedule limitations. This issue was discussed during the pre-warm up meeting with SLAC. J23 exceeded the required energy gain even with cav 3 at 8.9 MV/m. |  |  | Use As Is
566 | Feb-28-2024 | Feb-29-2024 | L2HE-CLNRM-FPCC-INSP | 56 | FPC COLD PART | CPI-FPCC-HE-021 | LeakTestFailure Vacuum | FPC was not under vacuum and missing washers on studs securing FPC body to coupler manifold. When opening valve to coupler the pressure on the pump cart raised to 400 Torr indicting that this set of couplers were not under vacuum. Once the couplers were under vacuum there was a leak detected that is higher than 2.69e-6 std/ccs |  |  | Return To Vendor | Return to vendor for repair
567 | Feb-28-2024 | Feb-29-2024 | L2HE-CLNRM-FPCC-INSP | 56 | FPC COLD PART | CPI-FPCC-HE-099 | LeakTestFailure Vacuum | FPC was not under vacuum. Once the coupler was under vacuum there was a leak detected that is higher than 2.69e-6 std/ccs |  |  | Return To Vendor | Return to vendor for repair
568 | Mar-01-2024 | Mar-01-2024 | L2HE-LERF-CM-ACTS | 2 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-23 | CMTFRFperf | Magnet test done according to SLAC commissioning procedure at 8.5 Amps for 1 hour rather than Magnet Acceptance Criteria which states run magnets at 18 A for 30 minutes. |  |  | Use As Is
569 | Mar-04-2024 | Mar-04-2024 | L2HE-INSP-CAV | 129 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R123 | Dimensional | One dimension is OOT: DIM2) Should Be: 0 to 1.0 ; Is: 1.3 See attached report. |  |  | Use As Is | Will be compensated for during 2 phase pipe welding
570 | Mar-04-2024 | Mar-12-2024 | L2HE-VTA-CAV-VTRF | 217 | PRODUCTION CAVITY ASSEMBLY | HE-295 | RFCMTF | Measured Q2 6.5400e+11, slightly bellow project Acceptance Criteria: 7.5 e11 |  |  | Use As Is | Accepted by CTB 2024-03-12
571 | Mar-05-2024 | Apr-10-2024 | L2HE-CMA-FPCW-INSP | 40 | FPC WARM PART | CPI-FPCW-HE-099 | ShippingDamage |  | Hold/Re-measure | Leak check | Conforming | Leak check was successful
573 | Mar-11-2024 | Dec-03-2024 | L2HE-LERF-CM-ACTS | 3 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-24 | CMTFRFperf LeakTestFailure Mechanical OTHER RFCMTF RFTestFailure Vacuum | Cav. 8 experiencing coupler vacuum/heating issues associated with anomalous drop in Qfpc from ~6e7 to ~3e7 and sagging gradient (requiring ever increasing power to maintain set point) once coupler reaches 125-130K. Initially thought to just be a vacuum issue, vacuum was processed with increasing forward power/gradient over 5 shifts until vacuum became maintainable at e-8 level. (Initially, it tripped coupler vacuum interlock at 2e-6.) Network analyzer was also used to verify Qfpc = 6.6e7 once cooled down. Hypothesized to be possible thermal strap or assembly issue. | Hold/Re-measure | Warm FPC was disassembled from the cold part. Signs of arcing were present on the cold section of coupler, and on the inside of the warm coupler. Pin on warm coupler push rod was found to be bent, indicating unusual force applied to the rod. The central screw may have been loose, causing a gap between the cold and warm coupler contact surfaces. Replace warm coupler with new unit. Clean off dark marks on cold coupler with abrasive paper. | Conforming | Original disposition 18-Jun-2024 by CHENG: The L2HE project decides to let JLAB re-test J1.3-24. Hence there will be another opportunity to verify whether the repair to the cavity 8 coupler has resolved the heating issue J1.3-24 re-test is completed. Cavity reached 22 MV/m gradient during the extended run with the coupler peak temperature of 109.436 K. The coupler heating issue is resolved
574 | Mar-11-2024 | Apr-23-2024 | L2HE-INSP-CAV | 130 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R191 | OTHER | BellowsBrace1; slot interference caused the 5/16-18 Hex-screw to damage the threads during installation of the screw. It appears that the slot was re-machined but did not clear the threads. | Hold/Modify | Please check out two brace arms from Mike Dickey (Marked with an "X") and compare if their slotted holes are at the wider diameter of the partially widen up holes on these brace arms. If that is the case, please get back to me so we can plan further work/repair on that cavity John V | Conforming | brace arms and screws replaced, holes re-tapped
575 | Mar-12-2024 | Mar-27-2024 | L2HE-VTA-CAV-VTRF | 219 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R123 | VTARFPerformance | During decay measurements, HOMB had pretty high value: 1.73e16. L2HE project specifications: QextHOMb Acceptance Criteria: >/= 2.7 e11. | Hold/Re-measure | Cavity will be tested in VTA again | Conforming | Broken power meter on HOM B line identified as error Retest on 2024-03-21 has correct power measurement of HOM B
576 | Mar-13-2024 | May-01-2024 | L2HE-CLNRM-CAV-TSTD | 51 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R186 | LeakTestFailure | Upon opening of cavity for vacuum check pressure rose to 1.7e-02 mBar |  |  | Conforming | As no leak could be found, the bottom flange assembly was replaced after an HPR of the cavity. The vacuum problem has been resolved.
577 | Mar-15-2024 | Mar-15-2024 | L2HE-INSP-BLBP | 95 | WELDMENT BELLOWS PRCM | 116 | Plating | Imperfections in copper plating have the appearance of plated over debris rather than blistering. |  |  | Use As Is | Use as spare
578 | Mar-15-2024 | Mar-15-2024 | L2HE-CMA-FPFT-LEAK | 25 | PICKUP FEEDTHROUGH | 2014 | LeakTestFailure | FPFT s/n 2014 has been verified as installed on J22 as cavity #8 (CAV-R012) and was found to be the source of the BL leak documented on NCR 451. Failed part to be segregated upon disassembly and investigated for root cause of failure. Also see JLab CAPA 051. | Hold/Re-measure | bench check feedthrough to verify leak. | Reject | FPFT s/n 2014 from J22 cavity 8 is confirmed as leaking and is rejected for use. Further investigation of leak will be documented on CAPA 051
579 | Mar-15-2024 | Mar-15-2024 | L2HE-INSP-BLBP | 96 | WELDMENT BELLOWS PRCM | 119 | Plating | Multiple dents in copper plating measure <= 1/8" and are caused by a deep, center pits- NR side. More photos at: M:\cmm\Pictures\Inspection Pictures\L2HE\Bellows production (BLBP-PrCM) F10041075\BLBP_119 2024.Feb.23\Dents-Pits |  |  | Reject | Dents on collar are deep
580 | Mar-21-2024 | May-29-2024 | L2HE-VTA-CAV-VTRF | 223 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R035 | RFTestFailure VTARFPerformance | Cavity has FE onset at 22.7 MV/m, quench at 22.8 MV/m. | Hold/Modify | re-rinse & re-test (small HPR -- slow bleed-up) CTB 2024-03-26 | Conforming | Cavity cleaned up in test on 2024-05-25
581 | Mar-21-2024 | Mar-27-2024 | L2HE-VTA-CAV-VTRF | 224 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R012 | RFTestFailure VTARFPerformance | Cavity quenched limited to 20.3 MV/m, no FE |  |  | Use As Is | CTB 2024-03-26 decision, set aside with chance to put into last string if no better cavities are available
582 | Mar-25-2024 | May-07-2024 | L2HE-CMA-CM-ASSY2 | 7 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-28 | Mechanical | It was noticed the UCM is missing aluminum rivets at one of the connection areas of the top 50k shield. | Hold/Modify | Please add rivets. See drawing attached, detail B shows rivets. Note 1.2 indicates that the center panels, i.e. items 4 and 5, shall be riveted only, no welding | Conforming | Missed rivets are added
583 | Mar-28-2024 | Jul-17-2024 | L2HE-VTA-CAV-VTRF | 226 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R030 | VTARFPerformance | Emax was 21.6 MV/m, less than spec (23 MV/m) |  |  | Conforming | Quench limit within spec during next test on 2024-05-14
584 | Apr-02-2024 | Dec-09-2024 | L2HE-LERF-CM-ACTS | 3 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-24 | Instrumentation | Four (4) DT-670A1-CU (bobbin-style, diode temperature sensors on HOMs) failed during insulating vacuum pumpdown and cryomodule cooldown of J24. TE-UH-1318 (Cav3 HOM2, tuner-end) TE-DH-1320 (Cav3 HOM1, coupler-end), TE-UH-1518 (Cav5 HOM2, tuner-end), and TE-DH-1820 (Cav8 HOM1, coupler-end) all read 295 Kelvin prior to and shortly after hookup in LERF vault, but then failed prior to cooldown. All worked properly when cold, but all four failed again during warmup. All four are presently failed at room temperature. | Hold/Re-measure | Measure resistance of individual wire pairs leading to each sensor terminal and measure fwd and reverse bias voltage of each diode sensor. | Use As Is | The L2HE project decides to let JLAB re-test J1.3-24. Hence there will be another opportunity to verify the fix to the secondary sensors installed. The three sensors replaced before retest all worked and are conforming. The fourth sensor replaced after retest, passes all warm checks but will not be cold tested, thus a use as is disposition.
585 | Apr-11-2024 | Apr-11-2024 | L2HE-CLNRM-CAV-TSTD5 | 57 | LCLS-II PRODUCTION CAVITY ASSEMBLY | 0351 | LeakTestFailure | HE 351 removed from TS due to leak at burst disk for burst disk replacement | Hold/Modify | Replace burst disc using clean assembly techniques. | Conforming | Cavity is leak tight with new burst disc. For the record, this was a deviation from standard practices (usually the cavity would have undergone a 2p HPR). The decision for a "simple" replacement of the BD was driven by time constraints and the fact that the risk of contaminating the cavity during the replacement off of the TS is very low. The CTB was informed and approved the decision on 2024-04-09.
586 | Apr-11-2024 | May-01-2024 | L2HE-CLNRM-CAV-TSTD5 | 57 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R135 | LeakTestFailure | RAV leaks by. Removed from TS to remove bottom flange rerinse and replace Valve | Hold/Modify | remove bottom flange, rerinse (2P) and replace RA valve during subsequent reassembly | Conforming | Vacuum leak has been resolved by replacing bottom flange assembly.
587 | Apr-17-2024 | Oct-10-2024 | L2HE-INSP-CAV | 132 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R036 | Dings Scratches | Scratches/defects on bellows in multiple areas (total of 8 areas). Please see attached pictures. |  |  | Return To Vendor | CTB 2024-05-28 return to vendor for warranty repair
589 | Apr-18-2024 | Apr-23-2024 | L2HE-CMA-FPCW-INSP | 35 | FPC WARM PART | CPI-FPCW-HE-140 | RFCMTF RFTestFailure VTARFPerformance | FPCW removed from J24, cavity 8, to investigate coupler heating during acceptance testing. |  |  | Reject | Reject due to arcing
590 | Apr-19-2024 | Jun-25-2024 | L2HE-CHEM-COMP-DEGR | 256 | HOM ANTENNA | 0843 | InspectionFailure OTHER | HOM Antenna is bent. Was told by CAM to put an NCR on the part and reach out to John V. or Liang Z. to find a replacement. | Hold/Modify | Replace with new feedthrough with SN 303 provided by inventory group | Conforming | HOM got replaced
591 | Apr-22-2024 | Oct-24-2024 | L2HE-INV-CAV | 135 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R014 | OTHER | Accelerometer blinking RED. See traveler for file. |  |  | Conforming | HE-CAVR-014 passed RFIN inspection 23 April 2024, cavity is conforming.
592 | Apr-23-2024 | May-29-2024 | L2HE-VTA-CAV-VTRF | 233 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R018 | VTARFPerformance | Low quench gradient 19.5 MV/m. No field emission or signs of multipacting |  |  | Conforming | gradient improved in test on 2024-05-18 after HEP rinse
593 | Apr-23-2024 | Sep-24-2024 | L2HE-VTA-CAV-VTRF | 232 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R135 | VTARFPerformance | Field Emission Onset at 14 MV/m, Neutron Alarm at 28 MV/m. | Hold/Modify | HPR, retest | Return To Vendor | 2024-09-24 sent to FNAL after big HPR for test in VTS
594 | Apr-23-2024 | Jul-01-2024 | L2HE-VTA-CAV-VTRF | 231 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R023 | RFTestFailure VTARFPerformance | Cold Leak developed after quench event. Unable to test due to plasma breakdown. |  |  | Conforming | cold leak has been resolved in next test (2024-05-25)
595 | Apr-25-2024 | Apr-25-2024 | L2HE-INSP-CAV | 134 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R014 | OTHER | Magnetic hygiene check of the helium vessel failed in one area of the Short fill line-transition joint. Gauss measured 9.6 before degaussing. Should Be: .65 Gauss max (.40 Gauss background), Is: 1.5 Gauss. See attached photos. |  |  | Use As Is | We will go forward and check on the performance of the cavity in the vertical test.
596 | Apr-25-2024 | Apr-25-2024 | L2HE-INSP-CAV | 134 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R014 | Dimensional | Location of tuner ring is OOT. Should Be: 5° +/- 1.0°, Is: 3.8°, OOT by -1.2° |  |  | Use As Is | Will be accounted for when tuner arms are fitted
597 | Apr-25-2024 | May-28-2024 | L2HE-CLNRM-FPCC-INSP | 63 | FPC COLD PART | CPI-FPCC-HE-127 | OTHER | During inspection of ceramic piece and copper area technician found hole #4 (from 12 o clock position) of protective cap area to have bad threads | Hold/Modify | Remove the FPC stand from the cleanroom and repair the thread while the FPC is still installed on the stand. The cap should remain in place while the threaded hole is being repaired. After repair, return the FPC pair to the cleanroom following the relevant cleaning steps from L2HE-PR-INSP-FPCC | Conforming | Bolt hole threads are repaired
598 | Apr-26-2024 | Oct-24-2024 | L2HE-CLNRM-CAV-TSTD5 | 61 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R031 | Dings LeakTestFailure | FPC test fixture was damaged causing leak in cavity which also damaged the FPC cavity flange | Hold/Modify | Please correct damage on FPC port (file/mechanical work) and replace test FPC (high Q antenna) with used one from string cavity (storage) | Conforming | Damage was corrected and cavity was built into CM J31, cavity string was leak tight.
599 | Apr-26-2024 | May-01-2024 | L2HE-CLNRM-FPCC-INSP | 65 | FPC COLD PART | CPI-FPCC-HE-188 | InspectionFailure | scratch found around RF probe hole |  |  | Use As Is | Scratch won't affect performance
600 | Apr-26-2024 | May-24-2024 | L2HE-CLNRM-FPCC-INSP | 65 | FPC COLD PART | CPI-FPCC-HE-189 | InspectionFailure | scratch found around rf probe hole |  |  | Use As Is | marks are covered by the Warm Antenna, not an rf surface
601 | Apr-29-2024 | May-01-2024 | L2HE-VTA-CAV-VTRF | 234 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R036 | VTARFPerformance | Qext2 was 3.2e12, out of spec (7.5e11~2.5e12). |  |  | Use As Is | Accepted by CTB 2024-04-30
602 | Apr-29-2024 | May-29-2024 | L2HE-VTA-CAV-VTRF | 235 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R186 | VTARFPerformance | FE onset at 8.7 MV/m, Q0 rolled off at 14 MV/m. | Hold/Modify | go for full disassembly and big HPR | Conforming | cavity cleaned up in test on 2024-05-17
603 | Apr-29-2024 | May-01-2024 | L2HE-CLNRM-FPCC-INSP | 64 | FPC COLD PART | CPI-FPCC-HE-190 | InspectionFailure | Flange plate has scratches at 2 and 6 o clock positions as well as pit looking spot at 12 and 6 o clock positions |  |  | Use As Is | Scratches won't affect performance
604 | Apr-29-2024 | May-24-2024 | L2HE-CLNRM-FPCC-INSP | 64 | FPC COLD PART | CPI-FPCC-HE-190 | InspectionFailure | scratches located on rf contact surface |  |  | Use As Is | marks are covered by the Warm Antenna, not an rf surface
605 | Apr-29-2024 | May-01-2024 | L2HE-CLNRM-FPCC-INSP | 64 | FPC COLD PART | CPI-FPCC-HE-190 | InspectionFailure | scratches located on sloped surface where rf pin is supposed to be installed |  |  | Use As Is | Scratch won't affect performance
606 | Apr-29-2024 | May-01-2024 | L2HE-CLNRM-FPCC-INSP | 64 | FPC COLD PART | CPI-FPCC-HE-191 | InspectionFailure | flange plate has pit looking spot at 5 o clock position |  |  | Use As Is | Scratch won't affect performance
607 | Apr-29-2024 | May-24-2024 | L2HE-CLNRM-FPCC-INSP | 64 | FPC COLD PART | CPI-FPCC-HE-191 | InspectionFailure | rf conn surface has 2 different sized circle indentations as well as scratches on surface |  |  | Use As Is | marks are covered by the Warm Antenna, not an rf surface
608 | Apr-29-2024 | May-01-2024 | L2HE-CLNRM-FPCC-INSP | 64 | FPC COLD PART | CPI-FPCC-HE-191 | InspectionFailure | rf hole has scratches and large dark spot that could be copper flaked off or pitting |  |  | Use As Is | Dark spot won't affect performance
609 | May-01-2024 | Jul-17-2024 | L2HE-VTA-CAV-VTRF | 237 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R012 | RFTestFailure VTARFPerformance | HE-R012 quench limited at 18 MV/m. No FE onset up to Emax of 18 MV/m. |  |  | Use As Is | Quench limit overcome in test on 2024-06-04. Cavity still has FE (NCR 626)
610 | May-03-2024 | Sep-13-2024 | L2HE-VTA-CAV-VTRF | 239 |  | HE-R191 | VTARFPerformance | Radiation onset first run at cavity gradient 13.5 MV/m, Second mode at 17 MV/m, MP at about 18 MV/m and RF processing on 18.5 MV/m for one hour ineffective in decreasing radiation level. | Hold/Modify | Go for a full disassembly and "big" HPR | Return To Vendor | 2024-09-11 sent to FNAL after small HPR for test in VTS
612 | May-08-2024 | Jul-01-2024 | L2HE-VTA-CAV-VTRF | 240 |  | HE-R031 | VTARFPerformance | Cavity quench at 16.7 MV/m. No Fe onset up to Emax. | Hold/Modify | Cavity will go to HEP rinse | Use As Is | CTB accepted cavity with marginal performance after next test on 2024-05-25
613 | May-09-2024 | Jul-01-2024 | L2HE-VTA-CAV-VTRF | 242 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R135 | VTARFPerformance | Lambda leak identified while attempting High RF power qualification test. | Hold/Modify | When cavity comes out of dewar, please check state of RAV. After removal from test stand, please perform leak check in cleanroom/VAA | Conforming | Leak resolved in test on 2024-05-30
614 | May-09-2024 | May-09-2024 | L2HE-VTA-CAV-VTRF | 238 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R014 | VTARFPerformance | Qext2=1.68e+11 smaller than L2HE project specifications (Qext2 Acceptance Criteria: 7.5 e11 |  |  | Use As Is | Is withing acceptable limit at installation. Accepted by CTB 2024-05-07
615 | May-13-2024 | May-17-2024 | L2HE-CMA-FPCW-INSP | 49 | FPC WARM PART | CPI-FPCW-HE-101 | ShippingDamage | Two tears in the outer bag |  |  | Use As Is | Cleanliness will be verified by particulate count at a later point in the inspection process
616 | May-13-2024 | Oct-10-2024 | L2HE-CMA-VV-INSP | 9 | VESSEL LCLS-II AFTER MACHINING WELDMENT | RFCHE-VV-012 | Dings OTHER Scratches | See attached spreadsheet for NCR conditions (scratches, dings, and paint overspray). | Hold/Modify | Clean, polish, especially the defects affecting sealing | Conforming | VV006 is installed on J1.3-28 which has completed final assembly. All necessary VV surface cleaning and polishing are done. The CM has passed final leak check
617 | May-14-2024 | May-24-2024 | L2HE-CMA-UCM-INSP | 10 | LCLS-II UPPER COLD MASS ASSEMBLY | RFCHE-UC-010 | OTHER | Line D vertical supports twisted/bent. Please see attached. |  |  | Use As Is | Similar distortion in rods has been identified previously. Fermi's recommendation is to use them as is. Please refer to the attached email for details.
619 | May-20-2024 | Jun-04-2024 | L2HE-VTA-CAV-VTRF | 246 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R018 | RFTestFailure | Cavity frequency below spec, measured 1300.095 MHz at 2K, previous test before HEP rinse was 1300.103 MHz. |  |  | Conforming | Set as conforming according to spec (1300.1MHz) by CTB 2024-06-04
620 | May-22-2024 | Sep-13-2024 | L2HE-VTA-CAV-VTRF | 243 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R030 | VTARFPerformance | FE onset 20 MV/m. | Hold/Modify | Go for a big HPR | Return To Vendor | 2024-09-11 sent to FNAL after small HPR for test in VTS
621 | May-22-2024 | Sep-24-2024 | L2HE-VTA-CAV-VTRF | 247 | LCLS-II PRODUCTION CAVITY ASSEMBLY | 0351 | VTARFPerformance | Final Emax was 20.3 MV/m, and radiation was above FE threshold beyond 20 MV/m. | Hold/Modify | try to clean up cavity with a small HPR | Return To Vendor | 2024-09-24 sent to FNAL after big HPR for test in VTS
622 | May-28-2024 | Jun-04-2024 | L2HE-VTA-CAV-VTRF | 248 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R031 | VTARFPerformance | HE-R031 (Pos1) quench limited at 21.6 MV/m. |  |  | Use As Is | Set aside as marginal CTB 2024-06-04
623 | May-28-2024 | Sep-13-2024 | L2HE-VTA-CAV-VTRF | 249 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R023 | VTARFPerformance | Radiation onset at 19.7 Mv/m. Qo drop below specifications over 23 MV/m. | Hold/Modify | Go for another big HPR | Return To Vendor | 2024-09-11 sent to FNAL after small HPR for test in VTS
625 | May-29-2024 | Sep-19-2024 | L2HE-VTA-CAV-VTRF | 245 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R040 | RFTestFailure VTARFPerformance | FE onset at 23.4 MV/m, radiation at max gradient 0.020 mR/hr. | Hold/Modify | Go for another small HPR | Return To Vendor | 2024-09-12 sent to FNAL after big HPR for test in VTS
626 | May-29-2024 | Sep-19-2024 | L2HE-VTA-CAV-VTRF | 252 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R012 | RFTestFailure VTARFPerformance | 21.15 MV/m quench, rad 0.464 mR/hr, onset 16.38 MV/m. | Hold/Modify | Go for a big HPR | Return To Vendor | 2024-09-12 sent to FNAL after big HPR for test in VTS
627 | May-30-2024 | May-30-2024 | L2HE-INSP-IMAG | 82 | MAGNETIC SHIELD CAVITY ASSEMBLY | RFCHE-MS-158 | Scratches | Blemish/imperfections present on a couple of the mag shields. Please see attached. |  |  | Use As Is | Imperfections are localized on a small scale and are not expected to degrade the overall shielding performance significantly
628 | May-31-2024 | Jun-03-2024 | L2HE-CLNRM-CAV-ASSY1 | 91 | HOM FEEDTHROUGH | 2046 | OTHER | Cavity HE-R023 was tested in VTA on May25. After VTA testing, this HOM feedthrough HEH2046 was disassembled from cavity to prepare for a "big" HPR. Prior cleaning for re-use, chemistry room staff noticed that it was bent. The probe is tilting slightly. Two dragging marks are visible, one near the probe tip, the other near the base of the tip. | Hold/Modify | Replace with new HOM feedthrough | Reject | HOM feedthrough replaced with new feedthrough (HEH2318).
629 | Jun-03-2024 | Jul-01-2024 | L2HE-VTA-CAV-VTRF | 256 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R040 | RFTestFailure VTARFPerformance | Qext2 measured high at 1.5100e+13. tested on TS 28-010. Previous test done on TS 34-002. |  |  | Use As Is | CTB accepted as use-as-is
630 | Jun-03-2024 | Jun-04-2024 | L2HE-CLNRM-CAV-ASSY1 | 87 | HOM FEEDTHRU | 0181 | OTHER | Cavity HE-R012 was tested in VTA on May18, first test post HEP rinse. After VTA testing, this HOM feedthrough 0181 was disassembled from cavity to prepare for a "big" HPR. Prior cleaning for re-use, chemistry room staff noticed that it was bent. The probe is tilting slightly. Two dragging marks are visible, one near the probe tip, the other near the base of the probe. |  |  | Reject | Both HOM feedthroughs on HE-CAVR012 were replaced with the new feedthroughs HEH2316 and HEH2317. (The second feedthrough was found to be intact, however, not serialized, hence, it was replaced too.)
631 | Jun-04-2024 | Dec-03-2024 | L2HE-LERF-CM-ACTS | 3 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-24 | CMTFRFperf RFCMTF | Measured Qo at 20.8 MV/m is 2.0E10 and does not meet the minimum acceptance criteria of 2.7E10. |  |  | Conforming | The L2HE project decides to let JLAB re-test J1.3-24. Hence there will be another opportunity to measure cavity Q0's 2024-12-03 J1.3-24 re-test results an improved Q0 of 3.0e10, which meets the acceptance criteria
632 | Jun-14-2024 | Apr-23-2025 | L2HE-CMA-UCM-INSP | 10 | LCLS-II UPPER COLD MASS ASSEMBLY | RFCHE-UC-010 | Dimensional | Dimensional inspection report shows multiple Out-of-Tolerance (OOT)'s. The following are deemed as outstanding: Line E upstream dx OOT is -10.5 mm Line D upstream dy OOT is -9.7 mm Line A downstream dx OOT is -7.3 mm Line F downstream dy OOT is 4.9 mm | Hold/Modify | Having discussed with CMA, we will make necessary adjustments and try to re-install the vendor-supplied pipe location template | Use As Is | According to the OTD dimensions, the cryo pipe locations have shifted during cryomodule assembly. The downstream end of Line A now shows a -13.34 mm deviation along the X direction. Line A is anchored using brackets and G-10 blocks, as shown in the pictures. However, the supports are not firm enough to hold Line A in position, so any adjustments made may not be maintained during transportation.
633 | Jun-21-2024 | Jun-21-2024 | L2HE-CMA-CM-ASSYF | 5 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-26 | OTHER | During RT of the process piping in PV piping assemblies, an area of concern was noticed. |  |  | Use As Is | On the top assembly drawing JL10127855 (see attached), sheet 2, Detail H, welds are required to be done per notes 5, 6 & 7. Then Note 5 states " All Process Piping Shall Be in Accordance to ASME B31.3 Process Piping Code Normal Fluid Service." The RT contractor uses a form ((see attached) that invokes ASME B31.3 Acceptance Criteria. Since the contractor says that he would fail the reference weld #4, that means weld#4 does not meet the B31.3 acceptance criteria. Hence there is a nonconformance. Weld # 10 is required to be RT'ed since it is a closure weld that can not be pressurized to 110% of its design pressure due to the risk of damaging the niobium cavity. However, the weld #4 and its companion welds can be high pressure tested. This is the Line A high pressure test that are required to be carried out on all L2HE cryomodules. Completed PS-7 form for the J1.3-26 Line A high pressure test is attached. Weld #4 as part of Line A has passed the pressure test prior to RT of the weld #10. The nonconformance on weld #4 RT (not required) is hence dismissed.
634 | Jun-25-2024 | Sep-19-2024 | L2HE-VTA-CAV-VTRF | 263 | PRODUCTION CAVITY ASSEMBLY | HE-040 | VTARFPerformance | FE onset at 21 MV/m after field emitter blew up at about 25 MV/m |  |  | Return To Vendor | 2024-09-12 sent to FNAL after big HPR for test in VTS
635 | Jul-01-2024 | Sep-24-2024 | L2HE-VTA-CAV-VTRF | 262 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R106 | RFTestFailure | Unable to test cavity R106 on 26 June 2024 due to suspected liquid helium leak. Pi mode could be measured with VNA but could not be locked with high power RF. Subsequent check with VNA showed a signature of helium in the cavity when the Pi mode frequency was viewed with 1 kHz span. |  |  | Return To Vendor | 2024-09-24 sent to FNAL after big HPR for test in VTS
636 | Jul-02-2024 | Jul-17-2024 | L2HE-LERF-CM-ACTS | 4 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-25 | CMTFRFperf RFCMTF RFTestFailure | The Maximum Usable Gradient for Cavity 1 is 15.6 NV/m which is lower than the Minimum Acceptance Criteria of 16 MV/m |  |  | Use As Is | John Hogan speaking for SLAC has approved a "Use As Is" disposition for this NCR.
637 | Jul-02-2024 | Jul-17-2024 | L2HE-LERF-CM-ACTS | 4 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-25 | CMTFRFperf RFCMTF RFTestFailure | The Average Qo for J25 is 2.6E10 which is lower than the Minimum Acceptance Criteria of 2.7E10. This also means the Dynamic Heat load is higher than 137 W at 173 MV |  |  | Use As Is | John Hogan speaking for SLAC has approved a "Use As Is" disposition for this NCR.
638 | Jul-03-2024 | Jul-03-2024 | L2HE-INSP-BLBP | 104 | WELDMENT BELLOWS PRCM | 006 | Plating | Bellows removed from cleanroom assembly. Please see pictures attached for copper flaking (rotatable flange end). | Hold/Modify | Send to chemistry to polish/clean flaking copper. | Reject | Flaky plating persists, for emergency use only
639 | Jul-09-2024 | Jul-31-2024 | L2HE-INV-CAV | 137 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R102 | OTHER | Accelerometer "Mission Duration" = 4h 48m, data full on day 1 of shipment before midnight. |  |  | Conforming | Cavity passed RFIN inspection.
640 | Jul-09-2024 | Jul-31-2024 | L2HE-INV-CAV | 136 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R050 | OTHER | Accelerometer "Mission Duration" = 4h 48m, data full on day 1 of shipment before midnight. |  |  | Conforming | Cavity passed RFIN inspection.
641 | Jul-09-2024 | Aug-01-2024 | L2HE-CLNRM-FPCC-INSP | 67 | FPC COLD PART | CPI-FPCC-HE-055 | LeakTestFailure Vacuum | On the FPC coupler 055 and 052 static vacuum and leak check a discrepancy was observed. The pump cart pressure was 3e-8mbar and when the right angle valve to the coupler was opened the pressure went to 2e-2mbar. Following this observation the subsequent bagged leak check was preformed which revealed an approximately 1e-7 leak. Inspection continued and no visual discrepancies were observed during the inspection. | Hold/Re-measure | Individual vacuum check | Conforming | Successful leak check
642 | Jul-10-2024 | Jul-29-2024 | L2HE-CLNRM-FPCC-INSP | 67 | FPC COLD PART | CPI-FPCC-HE-052 | Dings Oxidation Scratches | During the inspection of FPC052 oxidation, dings or dents, and scratches were observed on the copper ring on the bellow. Pictures were taken to document and attached here. |  |  | Use As Is | Dents and oxidation will not effect performance of heat sink
643 | Jul-10-2024 | Jul-17-2024 | L2HE-CLNRM-FPCC-INSP | 68 | FPC COLD PART | CPI-FPCC-HE-084 | Oxidation | FPC084 Discoloration on the copper around the RF pin. |  |  | Use As Is | Light oxidation won't affect performance
644 | Jul-10-2024 | Jul-17-2024 | L2HE-CLNRM-FPCC-INSP | 68 | FPC COLD PART | CPI-FPCC-HE-083 | Oxidation | FPC 083 Discoloration on the copper near the RF Pin. |  |  | Use As Is | Light oxidation won't affect performance
645 | Jul-11-2024 | Jul-12-2024 | L2HE-INSP-CAV | 135 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R102 | Dimensional | Burst disk orientation deviates from the 180 degrees from valve. |  |  | Use As Is | After consultation with string assembly group, the conclusion is that this issue will require a deviation in the alignment procedure during assemble, however, should not affect the precision of the alignment of the flange/cavity within the string itself.
646 | Jul-12-2024 | Jul-12-2024 | L2HE-INSP-CAV | 135 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R102 | Dings | Blemish on bellows. Please see attached. |  |  | Use As Is | Defects should not impair the functionality of the bellows or threaten to cause a leak. Will follow up with Fermi if defect had been identified before.
647 | Jul-12-2024 | Aug-08-2024 | L2HE-CLNRM-FPCC-INSP | 70 | FPC COLD PART | CPI-FPCC-HE-119 | Oxidation Stains | Both couplers 119's and 120's copper plating around the cold ceramics are severely discolored coupled with apparent staining. Copper plating around antennas for both were in a good state. | Hold/Modify | Send to chemistry for cleaning | Conforming | Copper surface free of stains
648 | Jul-15-2024 | Jul-15-2024 | L2HE-INSP-CAV | 136 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R050 | Dimensional | Right Angle Valve flange not 180° to burst disk, per F10023864. |  |  | Use As Is | After consultation with string assembly group, the conclusion is that this issue will require a deviation in the alignment procedure during assemble, however, should not affect the precision of the alignment of the flange/cavity within the string itself. See NCR 645
649 | Jul-16-2024 | Jul-17-2024 | L2HE-CLNRM-FPCC-INSP | 71 | FPC COLD PART | CPI-FPCC-HE-051 | Oxidation Stains | Light discoloration on FPCC051 copper ring above the bellows. |  |  | Use As Is | Discoloration won't affect performance
650 | Jul-19-2024 | Aug-01-2024 | L2HE-CLNRM-FPCC-INSP | 67 | FPC COLD PART | CPI-FPCC-HE-052 | LeakTestFailure | On the FPC coupler 055 and 052 static vacuum and leak check a discrepancy was observed. The pump cart pressure was 3e-8mbar and when the right angle valve to the coupler was opened the pressure went to 2e-2mbar. Following this observation the subsequent bagged leak check was preformed which revealed an approximately 1e-7 leak. Inspection continued and no visual discrepancies were observed during the inspection. | Hold/Re-measure | Vacuum check separately from SN 55 | Conforming | Successful leak check
651 | Jul-19-2024 | Aug-08-2024 | L2HE-CLNRM-FPCC-INSP | 70 | FPC COLD PART | CPI-FPCC-HE-120 | Oxidation Stains | Both couplers 119's and 120's copper plating around the cold ceramics are severely discolored coupled with apparent staining. Copper plating around antennas for both were in a good state. | Hold/Modify | Send to chemistry for cleaning | Conforming | Copper surface free of stains
652 | Jul-19-2024 | Apr-23-2025 | L2HE-CMA-VV-INSP | 10 | VESSEL LCLS-II AFTER MACHINING WELDMENT | RFCHE-VV-010 | Dimensional Dings Residues Scratches Stains | Visual inspection completed on 18-Jul-2024. All photos are on Server: M:\cmm\Pictures\Inspection Pictures\L2HE\Vacuum vessel\VV010 Scratches/pitted areas and paint over-spray were observed. See attached photos in the traveler or on server. | Hold/Modify | Use the attached VV010 QC task list as a guide, Clean, polish, especially the defects affecting sealing | Conforming | RFCHE-VV-010 is installed on J1.3-29 cryomodule (see picture). The sealing surfaces have been polished at WS5. The cryomodule has passed final leak check. See traveler L2HE-CMA-CM-ASSYF for J1.3-29 for details
653 | Aug-05-2024 | Aug-09-2024 | L2HE-CMA-CM-ASSY1 | 9 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-30 | OTHER | During installation of SPQA209 on the J30 string at WS2, it was noted that the outer-most cooling strap (F10037116_rev.F_CORE COOLING STRAP, LH) on the LH side of SPQA209 (F10009383_rev.W_ASSEMBLY, HALF-LH) was not properly adhered to the magnet. This could adversely affect the heat transfer out of the SC magnet core to the 2-phase pipe. The JLab receipt inspection of SPQA209 was completed on 11-Aug-2023. The magnet met all visual, electrical and mechanical inspection criteria, with the Inspector's comment, "Article 209 appearance was in good order." No defects were noted during a partial visual inspection of the magnet on 24-Jul-2024, following its removal from inventory and being mounted on the magnet transfer cart for assembly (IMG_3025). SPQA209 was installed on cavity string J30 on 30-Jul-2024. After installation, the two lower sections of the LH Core Cooling Strap were noted as having partially detached from the magnet body (IMG_3052). Inspection of the upper-most section showed a visible gap between the Core Cooling Strap and the magnet body (IMG_3047). A "tap test" was performed on the accessible areas of both the LH and RH cooling straps. The lower area of the RH strap sounds well attached. The upper and lower area of the LH strap makes a hollow, tinny sound when tapped. SPQA209 was removed from string J30 on 1-Aug-2024 and the entire Core Cooling Strap, LH came off of the magnet assembly (IMG_3089). The newly exposed blue Stycast between the strap and the magnet has an unusual, oily residue, suggesting that it may not be completely cured. Large voids up to 3mm in depth are visible over much of the mating surface (IMG_3076). |  |  | Return To Vendor
654 | Aug-06-2024 | Oct-15-2024 | L2HE-CMA-CM-ASSY1 | 9 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-30 | Dimensional Mechanical | Measurements on cavity 1 features during string fiducialization identified dimensional anomalies that need to be understood. The helium vessel appears to be shifted towards the cold coupler ~7mm.Current and past RI CMM data uploaded here for reference. See Tab labeled "Sheet 2" for cavity 1 lug location to the Cold coupler datum. | Hold/Re-measure | See attached. We looked up the history of Zannon cavities having similar non-conformances as well as the history of CAV295. A mock-up model is built to check the warm and cold positions of the supporting blocks w.r.t. the UCM cavity hangar pads. We decide to proceed with assembly. Will need to 1) check how the entire magnetic shield assembly fits with the non-conforming cavity; 2)re-measure distances between the cavity supporting blocks and roller bearings, at the U/S and D/S locations (4 places), once the UCM is engaged with the cavity string. | Use As Is | We have verified the roller bearing to HV lug distances at all 4 locations on cavity 1 of J30, no risk of rolling off is identified. The magnetic shield is fully installed w/o any need to modify the magnetic shield
655 | Oct-01-2024 | Dec-03-2024 | L2HE-LERF-CM-ACTS | 5 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-24 | CMTFRFperf RFCMTF | Cavity 8 HOM1 emits too much power at 20.8 MV/m. Measured power is 1.85 W (32.6 dBm), greater than the 1.7W spec. Qext still meets spec, 2.5e11 | Hold/Modify | The HOM coupler can be re-tuned before shipping. It can be accessed, with some difficulty, from the end of the cryomodule. | Use As Is | The 2K performance of this HOM cannot be verified without a full 2K test, hence the disposition can only be 'Use As Is'. However, we have confidence that the HOM tuning will fall within spec when the CM is tested again. The measured notch frequency is near the desired frequency when measured warm and is in-line with other HOMs.
656 | Nov-07-2024 | Nov-21-2024 | L2HE-INSP-CAV | 137 | LCLS-II PRODUCTION CAVITY ASSEMBLY | 0317 | Dings | Ding/damage present on bellows. Pictures attached. | Hold/Re-measure | Cavity under hold as of CTB 2024-11-08 Special case of two-ply bellows under investigation | Use As Is | This damage is accepted as "use as is" CTB 2024-11-08
657 | Nov-07-2024 | Feb-13-2025 | L2HE-INSP-CAV | 137 | LCLS-II PRODUCTION CAVITY ASSEMBLY | 0317 | Dings | Ding/sharp edge (#2) present on bellows. Please see attached. | Hold/Modify | Cavity under hold as of CTB 2024-11-08 Special case of two-ply bellows under | Return To Vendor | This cavity had initially been accepted by the CTB as "use as is" (2024-11-08), further investigation, however, let us revise this decision. The cavity is "on hold" for further use. Independent of that, it has been selected to be send to FNAL for potential use in one of their strings if no better cavities are qualified.
658 | Nov-07-2024 | Nov-13-2024 | L2HE-INSP-CAV | 137 | LCLS-II PRODUCTION CAVITY ASSEMBLY | 0317 | OTHER | Burst disk location deviates from 180 degrees from valve. Please see attached. |  |  | Use As Is | Issue noted, assembly tooling in cleanroom will have to be moved/adapted.
659 | Nov-08-2024 | Nov-19-2024 | L2HE-INSP-CAV | 138 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R202 | Roughness | 1) All (4) support blocks (rolling pads) exceed 0.8 micrometer surface finish (mm). Looks/feels about 2.0 surface finish. | Hold/Re-measure | Lugs have been honed, remeasure surface finish and recheck dimensions on CMM (John Vennekate) & repeat RF in measurement | Conforming | Cavity is conforming after repair as verified by CMM measurement.
660 | Nov-11-2024 | Nov-13-2024 | L2HE-INSP-CAV | 139 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R074 | OTHER | Valve/burst disk not properly rotated. |  |  | Use As Is | Issue noted, assembly tooling in cleanroom will have to be moved/adapted.
661 | Nov-12-2024 | Nov-13-2024 | L2HE-INSP-CAV | 141 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R111 | Dimensional | Rotation of Right-Angle Valve flange is 90° from Burst disk, not 180°. |  |  | Use As Is | Issue noted, assembly tooling in cleanroom will have to be moved/adapted.
662 | Nov-13-2024 | Nov-13-2024 | L2HE-INSP-CAV | 140 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R199 | OTHER | Burst disk is not rotated properly. |  |  | Use As Is | Issue noted, assembly tooling in cleanroom will have to be moved/adapted.
663 | Nov-14-2024 |  | L2HE-INSP-CAV | 137 | LCLS-II PRODUCTION CAVITY ASSEMBLY | 0317 | OTHER | Additional hardware was installed. "One spacer between split ring halves (usually on the left side, F10008700)"; There are two spacers installed. | Hold/Modify | Assembly discrepancy can be corrected at WS3, hold until corrected. | Use As Is | One spacer will be removed during tuner installation.
664 | Nov-14-2024 |  | L2HE-INSP-CAV | 141 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R111 | OTHER | Additional hardware was installed. "One spacer between split ring halves (usually on the left side, F10008700)"; There are two spacers installed. | Hold/Modify | Assembly discrepancy can be corrected at WS3, hold until corrected. | Use As Is | One spacer will be removed during tuner installation
665 | Nov-14-2024 | Nov-14-2024 | L2HE-INSP-CAV | 142 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R173 | OTHER | Burst disk is in incorrect location relative to valve. |  |  | Use As Is | Issue noted, assembly tooling in cleanroom will have to be moved/adapted.
666 | Nov-14-2024 | Nov-19-2024 | L2HE-INSP-CAV | 142 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R173 | Scratches | Tooling marks/scratches present on bellows between convolutions. Please see pictures attached below. |  |  | Reject | Rejected by CTB 2024-11-19
667 | Nov-14-2024 | Nov-14-2024 | L2HE-INSP-CAV | 140 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R199 | OTHER | Nothing engraved on field probe installed on cavity, it is missing a serial number. |  |  | Use As Is | According to FNAL data the FP is within spec, i.e. use as is.
668 | Nov-14-2024 | Feb-07-2025 | L2HE-LERF-CM-PREP | 7 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-26 | Instrumentation | J26 FPC 2_Bottom temperature sensor faulty. TE-CB-1213 (ACCL:L1B:0320:CPLRTEMP2) was reading 310K at room temperature, just prior to cooldown. The RTD is intermittently off-scale high. | Hold/Re-measure |  | Use As Is | The faulty temperature sensor is not accessible and can be repaired or replaced when the cryomodule is disassembled at the next available opportunity
669 | Nov-15-2024 | Nov-15-2024 | L2HE-INSP-CAV | 142 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R173 | OTHER | HOM Probe A - short has no engraving present. Missing serial number. |  |  | Use As Is | According to vertical test at FNAL the HOM is within spec & leak tight, i.e. use as is.
670 | Nov-15-2024 | Nov-15-2024 | L2HE-INSP-CAV | 143 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R013 | Dimensional | Rotation of Right-Angle Valve flange is 90° from Burst disk, not 180°. |  |  | Use As Is | Issue noted, assembly tooling in cleanroom will have to be moved/adapted.
671 | Nov-17-2024 | Nov-18-2024 | L2HE-INSP-CAV | 144 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R201 | Dimensional | Rotation of Right-Angle Valve flange is 90° from Burst disk, not 180°. |  |  | Use As Is | Issue noted, assembly tooling in cleanroom will have to be moved/adapted.
672 | Nov-17-2024 | Nov-18-2024 | L2HE-INSP-CAV | 144 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R201 | Dings | (1) small ding located on the outer sidewall, on the edge of the 1st convolution. See attached photos. |  |  | Conforming | As per SLAC definition of rejectable defects (scratch & material removal) this is an acceptable deviation.
673 | Nov-18-2024 | Nov-18-2024 | L2HE-INSP-CAV | 145 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R198 | Dimensional | Rotation of Right-Angle Valve flange is 90° from Burst disk, not 180°. |  |  | Use As Is | Issue noted, assembly tooling in cleanroom will have to be moved/adapted.
674 | Nov-25-2024 | Jan-23-2025 | L2HE-CLNRM-CST-ASSY | 13 | LCLS-II PRODUCTION CAVITY ASSEMBLY | HE-R198 | OTHER | During vacuum verification pressure before cavity opened at 3.7e-05 mBar upon opening vacuum rose to 2.3e+01 mBar | Hold/Re-measure | Perform big HPR. | Return To Vendor | Cavity to be shipped to FNAL for HPR and test CTB 2025-01-21.
675 | Dec-04-2024 | Dec-04-2024 | L2HE-INV-CAV | 149 | LCLS-II PRODUCTION CAVITY ASSEMBLY | 0313 | ShippingDamage | Accelerometer light is blinking RED |  |  | Conforming | Cavity passed RFIN inspection, conforming.
676 | Dec-05-2024 | Dec-12-2024 | L2HE-INSP-CAV | 146 | LCLS-II PRODUCTION CAVITY ASSEMBLY | 0313 | OTHER | Split ring has two spacers, traveler states to have only one installed. |  |  | Use As Is | One spacer will be removed during tuner installation.
677 | Dec-05-2024 | Dec-12-2024 | L2HE-INSP-CAV | 146 | LCLS-II PRODUCTION CAVITY ASSEMBLY | 0313 | Dimensional | Burst disk location is at wrong angle. |  |  | Use As Is | Issue noted, assembly tooling in cleanroom will have to be moved/adapted.
678 | Dec-05-2024 | Dec-12-2024 | L2HE-INSP-CAV | 146 | LCLS-II PRODUCTION CAVITY ASSEMBLY | 0313 | Dimensional | Distance inside of bellow bracket slot to other bracket slot distance across cavity out of tolerance. Also, lug to lug distance side to side across cavity out of tol. PDF report attached. |  |  | Use As Is | Can be compensated for during tuner installation.
679 | Dec-05-2024 | Dec-06-2024 | L2HE-CLNRM-CST-ASSY | 4 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-24 | Dimensional OTHER | The "Short" Bellows, F10023437-H-005 was installed between cavities 2/3 instead of between cavity 8 and the BPM. It is 2.6mm shorter than F10041075-H-003 which was installed downstream between cavity 8 and magnet instead of between cavities 2/3. Traveler ID: L2HE-D3 Traveler Seq Number: 67 was initiated 21-Nov-2022 to document this discrepancy. The difference between these two SNs is 2.6 mm according to the drawings. Both are copper plated per TR, Michelle Oast. |  |  | Use As Is | Both bellows F10041075, F11023435 are copper plated and have a free length difference of 2.6mm. CM has been cooled and tested twice with no BL vacuum issues.
680 | Dec-09-2024 | Mar-05-2025 | L2HE-LERF-CM-ACTS | 6 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-26 | CMTFRFperf | Cavity 2 Limit Switch: During extended range test of the cavity tuner, cavity 2 low limit switch was activated when the cavity frequency was 1299.553 MHz. The frequency range requirement is 1299.535 MHz, 18 kHz lower than what was measured. | Hold/Modify | When the module is warm, the tuner range will be re-measured. If possible, the low limit plunger will be adjusted for an extra 20 kHz range. | Use As Is | The lower limit of tuning range has been adjusted as planned
681 | Dec-13-2024 | Feb-10-2025 | L2HE-LERF-CM-ACTS | 6 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-26 | CMTFRFperf RFCMTF RFTestFailure | J26-CAV4 Fundamental Power Coupler exhibited excessive heating during cavity qualification. Over a 5-hour period, the FPC temperature rose to 170K and the Loaded-Q dropped from 5e7 to 2.5e7. This made the cavity limited by SSA power. Operating the FPC also increased the pressure in the FPC vacuum line, up to 1e-6 torr. During the unit test, Cavity 4 was held at 16 MV/m to avoid instabilities from FPC temperature and vacuum. After 10 hours at the reduced gradient, FPC temperature was 120K and the coupler vacuum pressure was stable. The cavity would not have been able to maintain 20.8 MV/m for 10 hours due to power limitation, excess coupler vacuum, or FPC temperatures exceeding 200K. The exact cause of the heating is unknown, but these symptoms are similar to (though less severe than) J24-Cav8 FPC, which had arcing between the warm and cold coupler. That FPC was cleaned and replaced in the Test Lab, and performed nominally when retested. | Hold/Modify | See the attached document for the initial investigation and proposed path forward. The cavity 4 warm coupler and waveguide need to be replaced. Root cause analysis of the heating issue is ongoing. J26 will be re-tested to verify the repair. | Conforming
682 | Jan-09-2025 | [5 empty] | N/A
683 | Jan-14-2025 | [5 empty] | N/A
684 | Jan-14-2025 | [5 empty] | N/A
685 | Jan-14-2025 | [5 empty] | N/A
686 | Jan-27-2025 |  | L2HE-INSP-FPCWG | 60 | ALUMINUM WAVEGUIDE | 044 | N/A | Spot on mating surface has no plating. Also, this is serial number 040 not 044 as the waveguide box was marked incorrectly.
687 | Jan-27-2025 |  | L2HE-INSP-FPCWG | 61 | ALUMINUM WAVEGUIDE | 041 | N/A | Spot on mating flange with no plating. Serial Number on box was 042 not 041
688 | Jan-28-2025 |  | L2HE-INSP-FPCWG | 62 | ALUMINUM WAVEGUIDE | 042 | N/A | There is a spot on the mating surface with no plating.
689 | Jan-28-2025 |  | L2HE-INSP-FPCWG | 63 | ALUMINUM WAVEGUIDE | 043 | N/A | There is a spot on the mating surface with no plating.
690 | Feb-07-2025 | Mar-06-2025 | L2HE-LERF-CM-ACTS | 8 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-26 | Instrumentation | J26 FPC 5 bottom RTD intermittent when warm. TE-CB-1513 (ACCL:L1B:0350:CPLRTEMP2) was reading normally during both periods of 2K operation, but was intermittently off-scale low while warm, from ~ 10DEC2024 to 14JAN2025. |  |  | Use As Is | The sensor showed intermittent off-scale readings; however, it functioned properly during the 2K tests. The sensor is deemed usable. This NCR serves as a future reference in case the sensor develops a failure or error in operation.
691 | Feb-27-2025 | Feb-28-2025 | L2HE-INSP-BLA | 2 | BEAM LINE HOM ABSORBER (STD LENGTH) | 316 | N/A | Deterioration of copper plating and/or discoloration present on multiple convolutions. Please see attached pictures. |  |  | Conforming | Cleaning successfully removed oxidation
693 | Mar-20-2025 | Mar-24-2025 | L2HE-LERF-CM-ACTS | 9 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-27 | Mechanical | J27 Cav7 tuner failed to meet the low frequency range during the extended range test. The cavity needs to reach 1299.535 MHz, or -465 kHz from 1.3 GHz. The cavity 7 frequency stopped changing at 1299.564 MHz, 30 kHz above the spec. No limit switches were tripped during the range test. |  |  | Use As Is | Use-as-is requested by customer representative, Sebastian Aderhold, during warm-up meeting 21-Mar-2025
694 | Apr-07-2025 |  | L2HE-CMA-CM-ASSY2 | 10 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-22R | N/A | The UCM002 used on J22R was found to have a non-conforming dimension on its 2-phase tee. Please see the attached slides for details.
695 | Apr-11-2025 | Apr-17-2025 | L2HE-LERF-CM-ACTS | 9 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-27 | N/A | Expected Heat load at gradients established during the unit test is 146.5 W. This is higher than the Acceptance requirement of 137 W and due to several Qo's that do not meet the Acceptance criteria. |  |  | Use As Is | Reopened NCR to add note. Use-as-is requested by customer representative, Sebastian Aderhold, during warm-up meeting 21-Mar-2025. Low Qo likely caused by poor fast cool down rate of cavity due to leaking JT valve. Investigated on D3 227. Performance expected to improve after JT valve cleaning. Not retested at JLab.
696 | Apr-11-2025 | May-01-2025 | L2HE-LERF-CM-ACTS | 10 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-28 | N/A | J28Cav7 Coupler: During the cold tuning of the J28 couplers, Cavity 7 WFPC tuning handle malfunctioned. Before the incident, it was turned 3 times CW, with recorded frequency and QL. Before completing the next rotation, the technician heard and felt a "clunk" sound from the coupler. At that point the handle had no effect on the tuning and felt free of tension. There was very little torque on the handle prior to it malfunctioning, ~8in/lb. Upon second inspection, the technician noted a "hard stop" where the handle couldn't turn any further They also noticed that the handle "wobbled" when held at the 12 or 6 positions. This is different compared to the known-good coupler on cavity 6. There was still no effect on tuning the QL of the cavity, it is stuck at 1.5e7. The required tuning range is 1e7 to 8e7, with operational coupling at 6e7. |  |  | Conforming | The repair to the cavity 7 FPC tuning system is successful. This FPC behaves normal during the subsequent cold test.
697 | Apr-17-2025 | Apr-17-2025 | L2HE-CMA-CM-SHIP | 6 |  | J1.3-27 | OTHER | Steps 21 and 22, measuring Q_FPC and Q_FieldProbe were not done before the module was shipped. A mathematical analysis of the measurement shows that the measurement cannot be completed accurately. The error bars are several times larger than the measured value. We will evaluate the necessity of this measurement before future shipments. |  |  | Use As Is | SLAC Cryomodule System Manager verbally approved to Use As Is on Apr. 17, 2025
698 | Apr-24-2025 |  | L2HE-CMA-VV-INSP | 11 | VESSEL LCLS-II AFTER MACHING WELDMENT | RFCG-VV-010 | Dings | Item 9; CL Port has dings on knife edge.
699 | Apr-30-2025 |  | L2HE-LERF-CM-ACTS | 11 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-28 | N/A | Cavity 1 is limited to a maximum useable gradient of 18.2 MV/m. This is below the requirement for a nominal gradient of 20.8 MV/m. This due to the radiation dose rate reaching 50 mR/hr at 18.2 MV/m |  |  |  |  | Need final disposition
700 | Apr-30-2025 |  | L2HE-LERF-CM-ACTS | 11 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-28 | N/A | Cavity 5 is limited to a maximum useable gradient of 15.5 MV/m. This is below the minimum acceptable gradient of 16.0 MV/m. This due to the radiation dose rate reaching 50 mR/hr at 15.5 MV/m |  |  |  |  | Need final disposition
701 | May-09-2025 |  | L2HE-LERF-CM-ACTS | 10 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-28 | Instrumentation | J28 Cav1 Limit switches - During cooldown of module J28 in the LERF, both switches on cavity 1 intermittently measured as open. The low-frequency limit switch stayed open most of the time, but not constantly. During a period when both switches were closed, the tuning range of cavity 1 was measured to meet the minimum +/- 20 kHz around 1.3 GHz. Escalation of D3 #231.
702 | May-09-2025 |  | L2HE-LERF-CM-ACTS | 11 | 1.3GHz HE CM LCLS-II ASSEMBLY | J1.3-28 | OTHER | J28 Microphonics Data - Microphonic detuning of cavities is usually recorded for 1 hour during the unit test. During the unit test, it was discovered that the computer which performs the data recording was no longer functional. There was not enough time to resolve the computer issue before the warmup review with SLAC and module warmup. Qualitatively, the microphonics seemed nominal compared to other module tests in the LERF as they did not present any issues during testing, but this cannot be verified quantitatively. |  |  |  |  | Need final disposition
703 | May-09-2025 |  | L2HE-LERF-CM-ACTS | 11 |  | J1.3-28 | N/A | J28 Cav1 QextFP - During cryomodule testing of J28, the Field Probe Q-external coupling factor was measured to be 1.2*10^14. The specification range for this value is 0.75-2.5*10^12. The high Qext (by 100x) also led to a weak signal for the RF control system. A 10dB attenuator was removed to get sufficient signal strength at the control system, the calibration factors were updated accordingly. The cavity and control system functioned normally for the remainder of the test. |  |  |  |  | Need final disposition