| VTA Test Component – Pressure Relief Calculation |  |  |  |
| --- | --- | --- | --- |
| Document Number: | <SRF-##-XX-###> | Effective Date: | DD Mmm YYYY |
| Revision Number: | <1, 2, 3> | Periodic Review Date: | DD Mmm YYYY |
| Document Owner: | <First Name Last Name> | Department Owner: | SRF Operations |

*** Form Instructions. To be deleted when creating your procedure. ***

- USE this template as a guide to create SRF Ops Reference Documents, SRF Programmatic Control and Master Documents, Project Documents and other general documents.

- DO NOT use this template for Procedures or QMS Process Documents.

- DO NOT use this template for Travelers.

- In the footer, replace this word document template name with the document number, title, and revision of your written document.

- Example text and section headings are included below to aid in font and format selection but may be changed.

- Periodic Review Date:

  - SRF Documents must be reviewed no later than three years from the previous revision.

  - A project document does not require periodic review except:

    - Document is being moved/used in another project

    - Changes to facilities, machinery, or infrastructure

    - Document order or process changes

Consult the Jefferson Lab Editorial Standards documents for guidance on good writing practice at https://www.jlab.org/sites/default/files/documents/comms/JLabEditorial%20Standards.pdf

*** Form Instructions. To be deleted when creating your procedure. ***

# Background

Components in the VTA typically are sealed and demonstrated as helium leak tight prior to immersion in a liquid helium environment.

# Problem

From time to time, the seals fail allowing liquid helium into the internal cavity space.

# Solution

A burst disk is installed as an engineered control to mitigate the potential for pressurization of the interior volume due to trapped helium during warm up.

Images of ASME Burst Disk – required on test components loaded into dewars for VTA testing

Analysis conducted determines the maximum mass flow rate as a function of temperature for the relief valve. The maximum flow rate is driven by the inner diameter of the tube in the burst disk sub-assembly. Considering that helium is a compressible gas, the flow is actually choked in the tube because sonic conditions for flow are achieved when the pressure ratio is greater than 0.49 for helium gas.

Reference: Compressible Fluid Flow by White

Calculation below is for three different tube sizes over a range of temperatures to account for a variety of cooldown, testing and warm up conditions.

At the lowest flow rates, cavity pressures will be relieved in less than 1 hour assuming cavity vacuum volumes less than 120 liters which is typical. This is sufficient when compared with dewar warm up times of 12 to 24 hours typically.

Conclusion: Burst disk with tube sizes ranging from 1/2“ to 1” diameter are acceptable to relieve cavity vacuum space in the event of helium ingress during warm up.

# Release and Revision History

| Rev # | Major Changes | Revision Date: |
| --- | --- | --- |
| 1 | Initial version (Utilizing SRF-07-FM-003 Word Document Template, R3) | DD Mmm YYY |
| <2> | <brief description of major changes>(Utilizing SRF-07-FM-003 Word Document Template, R3) | DD Mmm YYY |

# Approvals

| Approved by: | Name: |
| --- | --- |
| Document Owner | <First Last Name> |
| Document Reviewer 1 JLab SME or SRFOPS WCL | <First Last Name> |
| Document Approver 2 SRFOPS WCL or GL | <First Last Name> |

Document Processor Instructions:

- Put valid dates everywhere DD is found and verify they are accurate

- Attach DocuShare Approval Picture here